PS3.14
DICOM PS3.14 2016a - Grayscale Standard Display Function
PS3.14: DICOM PS3.14 2016a - Grayscale Standard Display Function
Copyright © 2016 NEMA
Table of Contents
Notice and Disclaimer	0
Foreword	0
1. Scope and Field of Application	0
2. Normative References	0
3. Definitions	0
4. Symbols and Abbreviations	0
5. Conventions	0
6. Overview	0
7. The Grayscale Standard Display Function	0
7.1. General Formulas	0
7.2. Transmissive Hardcopy Printers	0
7.3. Reflective Hardcopy Printers	0
8. References	0
A. Derivation of the Grayscale Standard Display Function (Informative)	0
A.1. Rationale For Selecting the Grayscale Standard Display Function	0
A.2. Details of the Barten Model	0
A.3. References	0
B. Table of the Grayscale Standard Display Function (Informative)	0
C. Measuring the Accuracy With Which a Display System Matches the Grayscale Standard Display Function (Informative)	0
C.1. General Considerations Regarding Conformance and Metrics	0
C.2. Methodology	0
C.3. References	0
D. Illustrations for Achieving Conformance with the Grayscale Standard Display Function (Informative)	0
D.1. Emissive Display Systems	0
D.1.1. Measuring the System Characteristic Curve	0
D.1.2. Application of the Standard Formula	0
D.1.3. Implementation of the Standard	0
D.1.4. Measures of Conformance	0
D.2. Transparent Hardcopy Devices	0
D.2.1. Measuring the System Characteristic Curve	0
D.2.2. Application of the Grayscale Standard Display Function	0
D.2.3. Implementation of the Grayscale Standard Display Function	0
D.2.4. Measures of Conformance	0
D.3. Reflective Display Systems	0
D.3.1. Measuring the System Characteristic Curve	0
D.3.2. Application of the Grayscale Standard Display Function	0
D.3.3. Implementation of the Grayscale Standard Display Function	0
D.3.4. Measures of Conformance	0
E. Realizable JND Range of a Display Under Ambient Light (Informative)	0
List of Figures
6-1. The Grayscale Standard Display Function is an element of the image presentation after several modifications to the image have been completed by other elements of the image acquisition and presentation chain.	0
6-2. The conceptual model of a Standardized Display System maps P-Values to Luminance via an intermediate transformation to Digital Driving Levels of an unstandardized Display System.	0
7-1. The Grayscale Standard Display Function presented as logarithm-of-Luminance versus JND-Index	0
A-1. Illustration for determining the transform that changes the Characteristic Curve of a Display System to a Display Function that approximates the Grayscale Standard Display Function	0
C-1. Illustration for the LUM and FIT conformance measures	0
D.1-1. The test pattern will be a variable intensity square in the center of a low Luminance background area.	0
D.1-2. Measured Characteristic Curve with Ambient Light of an emissive Display System	0
D.1-3. Measured and interpolated Characteristic Curve, Grayscale Standard Display Function and transformed Display Function of an emissive Display System. The transformed Display Function for this Display System matches the Grayscale Standard Display Function and the two curves are superimposed and indistinguishable.	0
D.1-4. LUM and FIT measures of conformance for a the transformed Display Function of an emissive Display System	0
D.2-1. Layout of a Test Pattern for Transparent Hardcopy Media	0
D.2-3. Plot of OD vs P-Value for an 8-Bit Printer	0
D.3-1. Measured and interpolated Characteristic Curve and Grayscale Standard Display Function for a printer producing reflective hard-copies	0
D.3-2. Transformation for modifying the Characteristic Curve of the printer to a Display Function that approximates the Grayscale Standard Display Function	0
D.3-3. Transformed Display Function and superimposed Grayscale Standard Display Function for a reflective hard-copy Display System. The transformed Display Function for this Display System matches the Grayscale Standard Display Function and the two curves are superimposed and indistinguishable.	0
D.3-4. Measures of conformance for a reflective hard-copy Display System with equal input and output digitization resolution of 8 bits	0
List of Tables
B-1. Grayscale Standard Display Function: Luminance versus JND Index	0
D.1-1. Measured Characteristic Curve plus Ambient Light	0
D.1-2. Look-Up Table for Calibrating Display System	0
D.2-1. Optical Densities for Each P-Value for an 8-Bit Printer	0
D.3-1. Look-Up Table for Calibrating Reflection Hardcopy System	0
Notice and Disclaimer
The information in this publication was considered technically sound by the consensus of persons engaged in the development and approval of the document at the time it was developed. Consensus does not necessarily mean that there is unanimous agreement among every person participating in the development of this document.
NEMA standards and guideline publications, of which the document contained herein is one, are developed through a voluntary consensus standards development process. This process brings together volunteers and/or seeks out the views of persons who have an interest in the topic covered by this publication. While NEMA administers the process and establishes rules to promote fairness in the development of consensus, it does not write the document and it does not independently test, evaluate, or verify the accuracy or completeness of any information or the soundness of any judgments contained in its standards and guideline publications.
NEMA disclaims liability for any personal injury, property, or other damages of any nature whatsoever, whether special, indirect, consequential, or compensatory, directly or indirectly resulting from the publication, use of, application, or reliance on this document. NEMA disclaims and makes no guaranty or warranty, expressed or implied, as to the accuracy or completeness of any information published herein, and disclaims and makes no warranty that the information in this document will fulfill any of your particular purposes or needs. NEMA does not undertake to guarantee the performance of any individual manufacturer or seller's products or services by virtue of this standard or guide.
In publishing and making this document available, NEMA is not undertaking to render professional or other services for or on behalf of any person or entity, nor is NEMA undertaking to perform any duty owed by any person or entity to someone else. Anyone using this document should rely on his or her own independent judgment or, as appropriate, seek the advice of a competent professional in determining the exercise of reasonable care in any given circumstances. Information and other standards on the topic covered by this publication may be available from other sources, which the user may wish to consult for additional views or information not covered by this publication.
NEMA has no power, nor does it undertake to police or enforce compliance with the contents of this document. NEMA does not certify, test, or inspect products, designs, or installations for safety or health purposes. Any certification or other statement of compliance with any health or safety-related information in this document shall not be attributable to NEMA and is solely the responsibility of the certifier or maker of the statement.
Foreword
This DICOM Standard was developed according to the procedures of the DICOM Standards Committee.
While other parts of the DICOM Standard specify how digital image data can be moved from system to system, it does not specify how the pixel values should be interpreted or displayed. PS3.14 specifies a function that relates pixel values to displayed Luminance levels.
A digital signal from an image can be measured, characterized, transmitted, and reproduced objectively and accurately. However, the visual interpretation of that signal is dependent on the varied characteristics of the systems displaying that image. Currently, images produced by the same signal may have completely different visual appearance, information, and characteristics on different display devices.
In medical imaging, it is important that there be a visual consistency in how a given digital image appears, whether viewed, for example, on the display monitor of a workstation or as a film on a light-box. In the absence of any standard that regulates how these images are to be visually presented on any device, a digital image that has good diagnostic value when viewed on one device could look very different and have greatly reduced diagnostic value when viewed on another device. Accordingly, PS3.14 was developed to provide an objective, quantitative mechanism for mapping digital image values into a given range of Luminance. An application that knows this relationship between digital values and display Luminance can produce better visual consistency in how that image appears on diverse display devices. The relationship that PS3.14 defines between digital image values and displayed Luminance is based upon measurements and models of human perception over a wide range of Luminance, not upon the characteristics of any one image presentation device or of any one imaging modality. It is also not dependent upon user preferences, which can be more properly handled by other constructs such as the DICOM Presentation Lookup Table.
The DICOM Standard is structured as a multi-part document using the guidelines established in [ISO/IEC Directives, Part 3].
DICOM® is the registered trademark of the National Electrical Manufacturers Association for its standards publications relating to digital communications of medical information, all rights reserved.
HL7® and CDA® are the registered trademarks of Health Level Seven International, all rights reserved.
SNOMED®, SNOMED Clinical Terms®, SNOMED CT® are the registered trademarks of the International Health Terminology Standards Development Organisation (IHTSDO), all rights reserved.
LOINC® is the registered trademark of Regenstrief Institute, Inc, all rights reserved.
1 Scope and Field of Application
PS3.14 specifies a standardized Display Function for display of grayscale images. It provides examples of methods for measuring the Characteristic Curve of a particular Display System for the purpose of either altering the Display System to match the Grayscale Standard Display Function, or for measuring the conformance of a Display System to the Grayscale Standard Display Function. Display Systems include such things as monitors with their associated driving electronics and printers producing films that are placed on light-boxes or alternators.
PS3.14 is neither a performance nor an image display standard. PS3.14 does not define which Luminance and/or Luminance Range or optical density range an image presentation device must provide. PS3.14 does not define how the particular picture element values in a specific imaging modality are to be presented.
PS3.14 does not specify functions for display of color images, as the specified function is limited to the display of grayscale images. Color Display Systems may be calibrated to the Grayscale Standard Display Function for the purpose of displaying grayscale images. Color images, whether associated with an ICC Profile or not, may be displayed on standardized grayscale displays, but there are no normative requirements for the display of the luminance information in a color image using the GSDF.
2 Normative References
The following standards contain provisions, which, through reference in this text, constitute provisions of this Standard. At the time of publication, the editions indicated were valid. All standards are subject to revision, and parties to agreements based on this Standard are encouraged to investigate the possibilities of applying the most recent editions of the standards indicated below.
[ISO/IEC Directives, Part 3] ISO/IEC. 1989. Drafting and presentation of International Standards.
3 Definitions
For the purposes of this Standard the following definitions apply.
3.1 Display Definitions
Characteristic Curve	The inherent Display Function of a Display System including the effects of ambient light. The Characteristic Curve describes Luminance versus DDL of an emissive display device, such as a CRT/display controller system, or Luminance of light reflected from a print medium, or Luminance derived from the measured optical density versus DDL of a hard-copy medium and the given Luminance of a light-box. The Characteristic Curve depends on operating parameters of the Display System.
Note
The Luminance generated by an emissive display may be measured with a photometer. Diffuse optical density of a hard-copy may be measured with a densitometer.
Contrast Sensitivity	characterizes the sensitivity of the average human observer to Luminance changes of the Standard Target. Contrast Sensitivity is inversely proportional to Threshold Modulation.
Contrast Threshold	A function that plots the Just-Noticeable Difference divided by the Luminance across the Luminance Range.
Digital Driving Level (DDL)	A digital value that given as input to a Display System produces a Luminance. The set of DDLs of a Display System is all the possible discrete values that can produce Luminance values on that Display System. The mapping of DDLs to Luminance values for a Display System produces the Characteristic Curve of that Display System. The actual output for a given DDL is specific to the Display System and is not corrected for the Grayscale Standard Display Function.
Display Function	A function that describes a defined grayscale rendition of a Display System, the mapping of the DDLs in a defined space to Luminance, including the effects of ambient light at a given state of adjustment of the Display System. Distinguished from Characteristic Curve, which is the inherent Display Function of a Display System.
Display System	A device or devices that accept DDLs to produce corresponding Luminance values. This includes emissive displays, transmissive hardcopy viewed on light boxes, and reflective hardcopy.
Illuminance	Light from the environment surrounding the Display System that illuminates the display medium. It contributes to the Luminance that is received by an observer from the image display. Ambient Light reduces the contrast in the image.
Just-Noticeable Difference (JND)	The Luminance difference of a given target under given viewing conditions that the average human observer can just perceive.
JND Index	The input value to the Grayscale Standard Display Function, such that one step in JND Index results in a Luminance difference that is a Just-Noticeable Difference.
Luminance	is the luminous intensity per unit area projected in a given direction. The International System unit (used in PS3.14) is candela per square meter (cd/m2), which is sometimes called nit. Another unit often used is footlambert (fL). 1 fL = 3.426 cd/m2.
Luminance Range	The span of Luminance values of a Display System from a minimum Luminance to a maximum Luminance.
P-Value	A device independent value defined in a perceptually linear grayscale space. The output of the DICOM Presentation LUT is P-Values, i.e., the pixel value after all DICOM defined grayscale transformations have been applied. P-Values are the input to a Standardized Display System.
Grayscale Standard Display Function	The mathematically defined mapping of an input JND index to Luminance values defined in PS3.14.
Standardized Display System	A device or devices that produce Luminance values that are related to input P-Values by the Grayscale Standard Display Function. How this is performed is not defined, though it may be achieved by transformation of P-Values into DDLs accepted by a Display System.
Standard Luminance Level	Any one of the Standard Luminance levels in Table B-1.
Standard Target	A 2-deg x 2-deg square filled with a horizontal or vertical grating with sinusoidal modulation of 4 cycles per degree. The square is placed in a uniform background of a Luminance equal to the mean Luminance of the Target.
Note
The Standard Target is defined in terms of the subtended viewing angle, not in terms of the distance from the viewer to the target.
Threshold Modulation	The minimum Luminance modulation required by the average human observer to detect the Standard Target at a given mean Luminance level. The Threshold Modulation corresponds to the Just-Noticeable Difference in Luminance of the Standard Target.
4 Symbols and Abbreviations
The following symbols and abbreviations are used in PS3.14.
ACR	American College of Radiology
ANSI	American National Standards Institute
CEN TC251	Comite' Europeen de Normalisation - Technical Committee 251 - Medical Informatics
DICOM	Digital Imaging and Communications in Medicine
HL7	Health Level 7
IEEE	Institute of Electrical and Electronics Engineers
ISO	International Standards Organization
JIRA	Japan Medical Imaging and Radiological Systems Industries Association
NEMA	National Electrical Manufacturers Association
5 Conventions
The following conventions are used in PS3.14:
The terminology defined in Section 3 above is capitalized throughout PS3.14.
6 Overview
PS3.14 defines, mathematically, the Grayscale Standard Display Function of Standardized Display Systems. These systems may be printers producing hard-copies viewed on light-boxes or electronic Display Systems for soft-copies.
Hard-copies may consist of transmissive films as well as reflective prints. The image in these prints is represented by optical density variations in transmission or diffuse reflection. To an observer, every element of the image appears with a certain Luminance depending on the Illuminance and the optical density of the image element.
Soft-copies may be produced by emissive Display Systems (such as CRT monitors) or electronic light valves (such as light sources and liquid crystal displays).
For the purpose of PS3.14, Display Systems take a Digital Driving Level and produce Luminance or optical density variations that represent the image. Predictable application of image transformations, such as the modality, value-of-interest, and presentation look-up tables specified in the DICOM standard, requires knowledge of the Characteristic Curve of the Display System. Standardizing the response function expected of the Display System simplifies the application of such image transformations across several different Display Systems such as encountered in a network environment.
PS3.14 does not define when conformance with the Grayscale Standard Display Function is achieved or how to characterize the degree of conformance reached.
Note
A definition of conformance would require thorough evaluations of human visual system sensitivity to deviations of Display Functions from the Grayscale Standard Display Function for medical images.
Figure 6-1 and Figure 6-2 show the context for the Grayscale Standard Display Function. The Grayscale Standard Display Function is part of the image presentation. There will be a number of other modifications to the image before the Grayscale Standard Display Function is applied. The image acquisition device will adjust the image as it is formed. Other elements may perform a "window and level" to select a part of the dynamic range of the image to be presented. Yet other elements can adjust the selected dynamic range in preparation for display. The Presentation LUT outputs P-Values (presentation values). These P-Values become the Digital Driving Levels for Standardized Display Systems. The Grayscale Standard Display Function maps P-Values to the log-luminance output of the Standardized Display System. How a Standardized Display System performs this mapping is implementation dependent.
The boundary between the DICOM model of the image acquisition and presentation chain, and the Standardized Display System, expressed in P-Values, is intended to be both device independent and conceptually (if not actually) perceptually linear. In other words, regardless of the capabilities of the Standardized Display System, the same range of P-Values will be presented ìsimilarlyî.

Figure 6-1. The Grayscale Standard Display Function is an element of the image presentation after several modifications to the image have been completed by other elements of the image acquisition and presentation chain.

Figure 6-2. The conceptual model of a Standardized Display System maps P-Values to Luminance via an intermediate transformation to Digital Driving Levels of an unstandardized Display System.
The main objective of PS3.14 is to define mathematically an appropriate Grayscale Standard Display Function for all image presentation systems. The purpose of defining this Grayscale Standard Display Function is to allow applications to know a priorihow P-Values are transformed to viewed Luminance values by a Standardized Display System. In essence, defining the Grayscale Standard Display Function fixes the "units" for the P-Values output from the Presentation LUT and used as Digital Driving Levels to Standardized Display Systems.
A second objective of PS3.14 is to select a Display Function that provides some level of similarity in grayscale perception or basic appearance for a given image between Display Systems of different Luminance and that facilitates good use of the available Digital Driving Levels of a Display System. While many different functions could serve the primary objective, this Grayscale Standard Display Function was chosen to meet the second objective. With such a function, P-Values are approximately linearly related to human perceptual response. Similarity does not guarantee equal information content. Display Systems with a wider Luminance Range and/or higher Luminance will be capable of presenting more just-noticeable Luminance differences to an observer. Similarity also does not imply strict perceptual linearity, since perception is dependent on image content and on the viewer. In order to achieve strict perceptual linearity, applications would need to adjust the presentation of images to match user expectations through the other constructs defined in the DICOM Standard (e.g., VOI and Presentation LUT). Without a defined Display Function, such adjustments on the wide variety of Display Systems encountered on a network would be difficult.
The choice of the function is based on several ideas that are discussed further in Annex A.
Annex B contains the Grayscale Standard Display Function in tabular form.
Informative Annex C provides an example procedure for comparing mathematically the shape of the actual Display Function with the Grayscale Standard Display Function and for quantifying how well the actual discrete Luminance intervals match those of the Grayscale Standard Display Function.
Display Systems often will have Characteristic Curves different from the Grayscale Standard Display Function. These devices may contain means for incorporating externally defined transformations that make the devices conform with the Grayscale Standard Display Function. PS3.14 provides examples of test patterns for Display Systems with which their behavior can be measured and the approximation to the Grayscale Standard Display Function evaluated (see Informative Section D.1, Section D.2 and Section D.3).
7 The Grayscale Standard Display Function
As explained in greater detail in Annex A, the Grayscale Standard Display Function is based on human Contrast Sensitivity. Human Contrast Sensitivity is distinctly non-linear within the Luminance Range of the Grayscale Standard Display Function . The human eye is relatively less sensitive in the dark areas of an image than it is in the bright areas of an image. This variation in sensitivity makes it much easier to see small relative changes in Luminance in the bright areas of the image than in the dark areas of the image. A Display Function that adjusts the brightness such that equal changes in P-Values will result in the same level of perceptibility at all driving levels is "perceptually linearized". The Grayscale Standard Display Function incorporates the notion of perceptual linearization without making it an explicit objective of PS3.14.
The employed data for Contrast Sensitivity are derived from Barten's model of the human visual system (Ref. 1, 2 and Annex B). Specifically, the Grayscale Standard Display Function refers to Contrast Sensitivity for the Standard Target consisting of a 2-deg x 2-deg square filled with a horizontal or vertical grating with sinusoidal modulation of 4 cycles per degree. The square is placed in a uniform background of Luminance equal to the mean Luminance L of the Target. The Contrast Sensitivity is defined by the Threshold Modulation at which the grating becomes just visible to the average human observer. The Luminance modulation represents the Just-Noticeable Difference (JND) for the Target at the Luminance L.
Note
The academic nature of the Standard Target is recognized. With the simple target, the essential objectives of PS3.14 appear to be realizable. Only spurious results with more realistic targets in complex surroundings were known at the time of writing PS3.14 and these were not assessed.
The Grayscale Standard Display Function is defined for the Luminance Range from 0.05 to 4000 cd/m2. The minimum Luminance corresponds to the lowest practically useful Luminance of cathode-ray-tube (CRT) monitors and the maximum exceeds the unattenuated Luminance of very bright light-boxes used for interpreting X-Ray mammography. The Grayscale Standard Display Function explicitly includes the effects of the diffused ambient Illuminance.
Within the Luminance Range happen to fall 1023 JNDs (see Annex A).
7.1 General Formulas
The Grayscale Standard Display Function is defined by a mathematical interpolation of the 1023 Luminance levels derived from Barten's model. The Grayscale Standard Display Function allows us to calculate luminance, L, in candelas per square meter, as a function of the Just-Noticeable Difference (JND) Index, j:

(7-1)

with:
Ln referring to the natural logarithm
j the index (1 to 1023) of the Luminance levels Lj of the JNDs
a = -1.3011877
b = -2.5840191E-2
c = 8.0242636E-2
d = -1.0320229E-1
e = 1.3646699E-1
f = 2.8745620E-2
g = -2.5468404E-2
h = -3.1978977E-3
k = 1.2992634E-4
m = 1.3635334E-3
The logarithms to the base 10 of the Luminance Lj are very well interpolated by this function over the entire Luminance Range. The relative deviation of any log(Luminance) -value from the function is at most 0.3%, and the root-mean-square-error is 0.0003. The continuous representation of the Grayscale Standard Display Function permits a user to compute discrete JNDs for arbitrary start levels and over any desired Luminance Range.
Note
To apply Equation 7-1 to a device with a specific range of L values, it is convenient to also have the inverse of this relationship, which is given by:
(7-2)

where:
Log10 represents logarithm to the base 10
A = 71.498068
B = 94.593053
C = 41.912053
D = 9.8247004
E = 0.28175407
F = -1.1878455
G = -0.18014349
H = 0.14710899
I = - 0.017046845
1. When incorporating the formulas for L(j) and j(L) into a computer program, the use of double precision is recommended.
2. Alternative methods may be used to calculate the JND Index values. One method is use a numerical algorithm such as the Van Vijngaarden-Dekker-Brent method described in Numerical Recipes in C(Cambridge University press, 1991). The value j may be calculated from L iteratively given the Grayscale Standard Display Function's formula for L(j). Another method would be to use the Grayscale Standard Display Function's tabulated values of j and L to calculate the j corresponding to an arbitrary L by linearly interpolating between the two nearest tabulated L,j pairs.
3. No specification is intended as to how these formulas are implemented. These could be implemented dynamically, by executing the equation directly, or through discrete values, such as a LUT, etc.
Annex B lists the Luminance levels computed with this equation for the 1023 integer JND Indices and Figure 7-1 shows a plot of the Grayscale Standard Display Function. The exact value of the Luminance levels, of course, depends on the start level of 0.05 cd/m 2.
The Characteristic Curve of a Display System represents the Luminance produced by a Display System as a function of DDL and the effect of ambient Illuminance. The Characteristic Curve is measured with Standard Test Patterns (see Annex D). In general, the Display Function describes, for example,
1. the Luminance (including ambient Illuminance) measured as a function of DDL for emissive displays such as a CRT-monitor/digital display controller system,
2. the Luminance (including ambient Illuminance) as a function of DDL measured for a transmissive medium hung in front of a light-box after a printer produced an optical density, depending on DDL, on the medium,
3. the Luminance (including ambient light) as a function of DDL measured for a diffusely reflective medium illuminated by a office lights after a printer produced a reflective density, depending on DDL, on the medium.
By internal or external means, the system may have been configured (or calibrated) such that the Characteristic Curve is consistent with the Grayscale Standard Display Function.
Some Display Systems adapt themselves to ambient light conditions. Such a system may conform to the Grayscale Standard Display Function for one level of ambient Illuminance only, unless it had the capability of adjusting its Display Function without user-intervention so that it remains in conformance with the Grayscale Standard Display Function.
7.2 Transmissive Hardcopy Printers
For transmissive hardcopy printing, the relationship between luminance, L, and the printed optical density, D, is:

(7-3)

where:
L0 is the luminance of the light box with no film present
La is the luminance contribution due to ambient illuminance reflected off the film
If film is to be printed with a density ranging from Dmin to Dmax, the final luminance will range between and and the j values will correspondingly range from jmin = j(Lmin) to jmax = j(Lmax).
If this span of j values is represented by an N-bit P-Value, ranging from 0 for jmin to 2N-1 for jmax, the j values will correspond to P-Values as follows:

(7-4)

and the corresponding L values will be L(j(p)).
Finally, converting the L(j(p)) values to densities results in:

(7-5)

Note
Typical values for the parameters used in transmissive hardcopy printing are L0 = 2000 cd/m2, La = 10 cd/m2.
7.3 Reflective Hardcopy Printers
For reflective hardcopy printing, the relationship between luminance, L, and the printed optical density, D, is:

(7-6)

where:
L0 is the maximum luminance obtainable from diffuse reflection of the illumination that is present.
If film is to be printed with a density ranging from Dmin to Dmax, the final luminance will range between and and the j values will correspondingly range from jmin = j(Lmin) to jmax = j(Lmax).
If this span of j values is represented by an N-bit P-Value, ranging from 0 for jmin to 2N-1 for jmax, the j values will correspond to P-Values as follows:

(7-7)

and the corresponding L values will be L(j(p)).
Finally, converting the L(j(p)) values to densities results in

(7-8)

Note
Typical values for the parameters used in reflective hardcopy printing are L0 = 150 cd/m2.
8 References
1) Barten, P.G.J., Physical model for the Contrast Sensitivity of the human eye. Proc. SPIE 1666, 57-72 (1992)
2) Barten, P.G.J., Spatio-temporal model for the Contrast Sensitivity of the human eye and its temporal aspects. Proc. SPIE 1913-01 (1993)

Figure 7-1. The Grayscale Standard Display Function presented as logarithm-of-Luminance versus JND-Index
A Derivation of the Grayscale Standard Display Function (Informative)
A.1 Rationale For Selecting the Grayscale Standard Display Function
In choosing the Grayscale Standard Display Function, it was considered mandatory to have only one continuous, monotonically behaving mathematical function for the entire Luminance Range of interest. Correspondingly, for simplicity of implementing the Grayscale Standard Display Function, it was felt to be useful to define it by only one table of data pairs. As a secondary objective, it was considered desirable that the Grayscale Standard Display Function provide similarity in grayscale rendition on Display Systems of different Luminance Range and that good use of the available DDLs of a Display System was facilitated.
Perceptual linearization was thought to be a useful concept for arriving at a Grayscale Standard Display Function for meeting the above secondary objectives; however, it is not considered an objective by itself. Apart from the fact that is probably an elusive goal to perceptually linearize all types of medical images under various viewing conditions by one mathematical function, medical images are mostly presented by application-specific Display Functions that assign contrast non-uniformly according to clinical needs.
Intuitively, one would assume that perceptually linearized images on different Display Systems will be judged to be similar. To achieve perceptual linearization, a model of the human visual system response was required and the Barten model [A1] was chosen.
Early experiments showed that an appealing degree of contrast equalization and similarity could be obtained with a Display Function derived from Barten's model of human visual system response. The employed images were square patterns, the SMPTE pattern, and the Briggs' pattern [A2].
It was wished to relate DDLs of a Display System to some perceptually linear scale, primarily, to gain efficient utilization of the available input levels. If digitization levels lead to luminance or optical density levels that are perceptually indistinguishable, they are wasted. If they are too far apart, the observer may see contours. Hence, the concept of perceptual linearization was retained, not as a goal for the Grayscale Standard Display Function, but to obtain a concept for a measure of how well these objectives have been met.
Perceptual linearization is realizable, in a strict sense, only for rather simple images like square patterns or gratings in a uniform surrounding. Nevertheless, the concept of a perceptually linearized Display Function derived from experiments with simple test patterns has been successfully applied to complex images as described in the literature [A3-A8]. While it was clearly recognized that perceptual linearization can never be achieved for all details or spatial frequencies and object sizes at once, perceptual linearization for frequencies and object sizes near the peak of human Contrast Sensitivity seemed to do a ìreasonable jobî also in complex images.
Limited (unpublished) experiments have indicated that perceptual linearization for a particular detail in a complex image with a wide Luminance Range and heterogeneous surround required Display Functions that are rather strongly bent in the dark regions of the image and that such Display Functions for a low-luminance and a high-luminance display system would not be part of a continuous, monotonic function. This experience may underly the considerations of the CIELab curve [A9] proposed by other standards groups.
Other experiments and observations with computed radiographs seemed to suggest that similarity could also be obtained between grayscale renditions on Display Systems of different Luminance when the same application-specific function is combined with log-linear Characteristic Curves of the Display Systems. Thus similarity, if not contrast equalization, could be gained by a straight, luminance-independent shape for the Display Function.
While it might have been equally sensible to choose the rather simple log-linear Display Function as a standard, this was not done for the following reason, among others.
For high-resolution Display Systems with high intrinsic video bandwidth, digitization resolution is limited to 8 or 10 bits because of technology and other constraints. The more a Grayscale Standard Display Function deviates from the Characteristic Curve of a Display System, the poorer the utilization of DDLs typically is from a perception point of view. The Characteristic Curve of CRT Display Systems has a convex curvature with respect to a log-linear straight line. It differs much less from Display Functions derived from human vision models and the concept of perceptual linearization than from a log-linear Display Function.
When using application-specific display processes that cause the resultant Display Function to deviate strongly from the Grayscale Standard Display Function, the function conceivably does not provide good similarity. In this case, other functions may yield better similarity.
In summary, a Display Function was derived from Barten's model of the human visual system to gain a single continuous mathematical function which in its curvature falls between a log-linear response and a Display Function that may yield perceptual linearization in complex scenery with a wide luminance range within the image. Other models of human contrast sensitivity may potentially provide a better function, but were not evaluated. The notion of perceptual linearization was chosen to meet the secondary objectives of the Grayscale Standard Display Function, but not as an explicit goal of the Grayscale Standard Display Function itself. It is recognized that better functions may exist to meet these objectives. It is believed that almost any single mathematically defined Standard Function will greatly improve image presentations on Display Systems in communication networks.
A.2 Details of the Barten Model
Barten's model considers neural noise, lateral inhibition, photon noise, external noise, limited integration capability, the optical modulation transfer function, orientation, and temporal filtering. Neuron noise represents the upper limit of Contrast Sensitivity at high spatial frequencies. Low spatial frequencies appear to be attenuated by lateral inhibition in the ganglion cells that seems to be caused by the subtraction of a spatially low-pass filtered signal from the original. Photon noise is defined by the fluctuations of the photon flux h, the pupil diameter d, and quantum detection efficiency η of the eye. At low light levels, the Contrast Sensitivity is proportional to the square-root of Luminance according to the de Vries-Rose law. The temporal integration capability in the model used here is simply represented by a time constant of T = 0.1 sec. Temporal filtering effects are not included. Next to the temporal integration capability, the eye also has limited spatial integration capability: There is a maximum angular size XE x YE as well as a maximum number of cycles NE over which the eye can integrate information in the presence of various noise sources. The optical modulation transfer function

(A-1)

(u, spatial frequency in c/deg) is derived from a Gaussian point-spread function including the optical properties of the eye-lens, stray light from the optical media, diffusion in the retina, and the discrete nature of the receptor elements as well as from the spherical aberration, Csph, which is the main pupil-diameter-dependent component. σ0 is the value of σ at small pupil sizes. External noise may stem from Display System noise and image noise. Contrast sensitivity varies approximately sinusoidally with the orientation of the test pattern with equal maximum sensitivity at 0 and 90 deg and minimal sensitivity at 45 de.g., The difference in Contrast Sensitivity is only present at high spatial frequencies. The effect is modeled by a variation in integration capability.
The combination of these effects yields the equation for contrast as a function of spatial frequency:

(A-2)

The effect of noise appears in the first parenthesis within the square-root as a noise contrast related to the variances of photon (first term), filtered neuron (second term), and external noise. The Illuminance, IL = π/4 d2L, of the eye is expressed in trolands [td], d is the pupil diameter in mm, and L the Luminance of the Target in cd/m2. The pupil diameter is determined by the formula of de Groot and Gebhard:

d = 4.6 - 2.8 . tanh(0.4 . Log10(0.625 . L))	(A-3)

The term (1 - F(u))2 = 1 - exp(-u2/u0 2) describes the low frequency attenuation of neuron noise due to lateral inhibition (u0 = 8 c/deg). Equation A-2 represents the simplified case of square targets, X0 = Y0 [deg]. Φext is the contrast variance corresponding to external noise. k = 3.3, η = 0.025, h = 357.3600 photons/td sec deg2; the contrast variance corresponding to the neuron noise Φ0 = 3.10-8 sec deg2, XE = 12 deg, NE = 15 cycles (at 0 and 90 deg and NE = 7.5 cycles at 45 deg for frequencies above 2 c/deg), σ0 = 0.0133 deg, Csph = 0.0001 deg/mm3 [A1]. Equation A-2 provides a good fit of experimental data for 10-4 ≤ L ≤ 103 cd/m2, 0.5 ≤ X0 ≤ 60 deg, 0.2 ≤ u ≤ 50 c/deg.
After inserting all constants, Equation A-2 reduces to

(A-4)

with q1 = 0.1183034375, q2 = 3.962774805 . 10-5, and q3 = 1.356243499 . 10-7.
When viewed from 250 mm distance, the Standard Target has a size of about 8.7 mm x 8.7 mm and the spatial frequency of the grid equals about 0.92 line pairs per millimeter.
The Grayscale Standard Display Function is obtained by computing the Threshold Modulation Sj as a function of mean grating Luminance and then stacking these values on top of each other. The mean Luminance of the next higher level is calculated by adding the peak-to-peak modulation to the mean Luminance Lj of the previous level:

(A-5)

Thus, in PS3.14, the peak-to-peak Threshold Modulation is called a just-noticeable Luminance difference.
When a Display System conforms with the Grayscale Standard Display Function, it is perceptually linearized when observing the Standard Target: If a Display System had infinitely fine digitization resolution, equal increments in P-Value would produce equally perceivable contrast steps and, under certain conditions, just-noticeable Luminance differences (displayed one at a time) for the Standard Target (the grating with sinusoidal modulation of 4 c/degree over a 2 degree x 2 degree area, embedded in a uniform background with a Luminance equal to the mean target Luminance).
The display of the Standard Target at different Luminance levels one at a time is an academic display situation. An image containing different Luminance levels with different targets and Luminance distributions at the same time is in general not perceptually linearized. It is once more emphasized that the concept of perceptual linearization of Display Systems for the Standard Target served as a logical means for deriving a continuous mathematical function and for meeting the secondary goals of the Grayscale Standard Display Function. The function may represent a compromise between perceptual linearization of complex images by strongly-bent Display Functions and gaining similarity of grayscale perception within an image on Display Systems of different Luminance by a log-linear Display Function.
The Characteristic Curve of the Display System is measured and represented by {Luminance, DDL}-pairs Lm = F(Dm). A discrete transformation may be performed that maps the previously used DDLs, Dinput, to Doutput according to Equations (A6) and (A7) such that the available ensemble of discrete Luminance levels is used to approximate the Grayscale Standard Display Function L = G(j). The transformation is illustrated in Fig. A1. By such an operation, conformance with the Grayscale Standard Display Function may be reached.

Doutput = s . F-1[G(j)]	(A-6)

s is a scale factor for accommodating different input and output digitization resolutions.
The index j (which in general will be a non-integer number) of the Standard Luminance Levels is determined from the starting index j0 of the Standard Luminance level at the minimum Luminance of the Display System (including ambient light), the number of Standard JNDs, NJND, over the Luminance Range of the Display System, the digitization resolution DR, and the DDLs, Dinput, of the Display System:

I = I0 + NJND / DR . Dinput	(A-7)

A detailed example for executing such a transformation is given in Annex D.
A.3 References
[A1] P.G.J. Barten: Physical model for the Contrast Sensitivity of the human eye. Proc. SPIE 1666, 57-72 (1992) and Spatio-temporal model for the Contrast Sensitivity of the human eye and its temporal aspects. Proc. SPIE 1913-01 (1993)
[A2] S.J. Briggs: Digital test target for display evaluation .Proc. SPIE 253, 237-246 (1980)
[A3] S.J. Briggs: Photometric technique for deriving a "best gamma" for displays .Proc. SPIE 199, Paper 26 (1979) and Opt. Eng. 20,651-657 (1981)
[A4] S.M. Pizer: Intensity mappings: linearization, image-based, user-controlled .Proc. SPIE 271, 21-27 (1981)
[A5] S.M. Pizer: Intensity mappings to linearize display devices .Comp. Graph. Image. Proc. 17, 262-268 (1981)
[A6] R.E. Johnston, J.B. Zimmerman, D.C. Rogers, and S.M. Pizer: Perceptual standardization .Proc. SPIE 536, 44-49 (1985)
[A7] R.C. Cromartie, R.E. Johnston, S.M. Pizer, D.C. Rogers: Standardization of electronic display devices based on human perception .University of North Carolina at Chapel Hill, Technical Report 88-002, Dec. 1987
[A8] B. M. Hemminger, R.E. Johnston, J.P. Rolland, K.E. Muller: Perceptual linearization of video display monitors for medical image presentation .Proc. SPIE 2164, 222-241 (1994)
[A9] CIE 1976

Figure A-1. Illustration for determining the transform that changes the Characteristic Curve of a Display System to a Display Function that approximates the Grayscale Standard Display Function
B Table of the Grayscale Standard Display Function (Informative)
The Grayscale Standard Display Function based on the Barten model was introduced in Section 7 and details are presented in Annex A above. This annex presents the Grayscale Standard Display Function as a table of values for Luminance as a function of the Just-Noticeable Difference Index for integer values of the Just-Noticeable Difference Index.
Table B-1. Grayscale Standard Display Function: Luminance versus JND Index
JND	L[cd/m 2]		JND	L[cd/m 2]		JND	L[cd/m 2]		JND	L[cd/m 2]
1	0.0500		2	0.0547		3	0.0594		4	0.0643
5	0.0696		6	0.0750		7	0.0807		8	0.0866
9	0.0927		10	0.0991		11	0.1056		12	0.1124
13	0.1194		14	0.1267		15	0.1342		16	0.1419
17	0.1498		18	0.1580		19	0.1664		20	0.1750
21	0.1839		22	0.1931		23	0.2025		24	0.2121
25	0.2220		26	0.2321		27	0.2425		28	0.2532
29	0.2641		30	0.2752		31	0.2867		32	0.2984
33	0.3104		34	0.3226		35	0.3351		36	0.3479
37	0.3610		38	0.3744		39	0.3880		40	0.4019
41	0.4161		42	0.4306		43	0.4454		44	0.4605
45	0.4759		46	0.4916		47	0.5076		48	0.5239
49	0.5405		50	0.5574		51	0.5746		52	0.5921
53	0.6100		54	0.6281		55	0.6466		56	0.6654
57	0.6846		58	0.7040		59	0.7238		60	0.7440
61	0.7644		62	0.7852		63	0.8064		64	0.8278
65	0.8497		66	0.8718		67	0.8944		68	0.9172
69	0.9405		70	0.9640		71	0.9880		72	1.0123
73	1.0370		74	1.0620		75	1.0874		76	1.1132
77	1.1394		78	1.1659		79	1.1928		80	1.2201
81	1.2478		82	1.2759		83	1.3044		84	1.3332
85	1.3625		86	1.3921		87	1.4222		88	1.4527
89	1.4835		90	1.5148		91	1.5465		92	1.5786
93	1.6111		94	1.6441		95	1.6775		96	1.7113
97	1.7455		98	1.7802		99	1.8153		100	1.8508
101	1.8868		102	1.9233		103	1.9601		104	1.9975
105	2.0352		106	2.0735		107	2.1122		108	2.1514
109	2.1910		110	2.2311		111	2.2717		112	2.3127
113	2.3543		114	2.3963		115	2.4388		116	2.4817
117	2.5252		118	2.5692		119	2.6137		120	2.6587
121	2.7041		122	2.7501		123	2.7966		124	2.8436
125	2.8912		126	2.9392		127	2.9878		128	3.0369
129	3.0866		130	3.1367		131	3.1875		132	3.2387
133	3.2905		134	3.3429		135	3.3958		136	3.4493
137	3.5033		138	3.5579		139	3.6131		140	3.6688
141	3.7252		142	3.7820		143	3.8395		144	3.8976
145	3.9563		146	4.0155		147	4.0754		148	4.1358
149	4.1969		150	4.2586		151	4.3209		152	4.3838
153	4.4473		154	4.5115		155	4.5763		156	4.6417
157	4.7078		158	4.7745		159	4.8419		160	4.9099
161	4.9785		162	5.0479		163	5.1179		164	5.1886
165	5.2599		166	5.3319		167	5.4046		168	5.4780
169	5.5521		170	5.6269		171	5.7024		172	5.7786
173	5.8555		174	5.9331		175	6.0114		176	6.0905
177	6.1702		178	6.2508		179	6.3320		180	6.4140
181	6.4968		182	6.5803		183	6.6645		184	6.7496
185	6.8354		186	6.9219		187	7.0093		188	7.0974
189	7.1863		190	7.2760		191	7.3665		192	7.4578
193	7.5500		194	7.6429		195	7.7366		196	7.8312
197	7.9266		198	8.0229		199	8.1199		200	8.2179
201	8.3167		202	8.4163		203	8.5168		204	8.6182
205	8.7204		206	8.8235		207	8.9275		208	9.0324
209	9.1382		210	9.2449		211	9.3525		212	9.4611
213	9.5705		214	9.6809		215	9.7922		216	9.9044
217	10.0176		218	10.1318		219	10.2469		220	10.3629
221	10.4800		222	10.5980		223	10.7169		224	10.8369
225	10.9579		226	11.0799		227	11.2028		228	11.3268
229	11.4518		230	11.5779		231	11.7050		232	11.8331
233	11.9622		234	12.0925		235	12.2237		236	12.3561
237	12.4895		238	12.6240		239	12.7596		240	12.8963
241	13.0341		242	13.1730		243	13.3130		244	13.4542
245	13.5965		246	13.7399		247	13.8844		248	14.0302
249	14.1770		250	14.3251		251	14.4743		252	14.6247
253	14.7763		254	14.9291		255	15.0831		256	15.2384
257	15.3948		258	15.5525		259	15.7114		260	15.8716
261	16.0330		262	16.1957		263	16.3596		264	16.5249
265	16.6914		266	16.8592		267	17.0283		268	17.1987
269	17.3705		270	17.5436		271	17.7180		272	17.8938
273	18.0709		274	18.2494		275	18.4293		276	18.6105
277	18.7931		278	18.9772		279	19.1626		280	19.3495
281	19.5378		282	19.7275		283	19.9187		284	20.1113
285	20.3054		286	20.5009		287	20.6980		288	20.8965
289	21.0966		290	21.2981		291	21.5012		292	21.7058
293	21.9120		294	22.1197		295	22.3289		296	22.5398
297	22.7522		298	22.9662		299	23.1818		300	23.3990
301	23.6179		302	23.8383		303	24.0605		304	24.2842
305	24.5097		306	24.7368		307	24.9656		308	25.1961
309	25.4283		310	25.6622		311	25.8979		312	26.1353
313	26.3744		314	26.6153		315	26.8580		316	27.1025
317	27.3488		318	27.5969		319	27.8468		320	28.0985
321	28.3521		322	28.6075		323	28.8648		324	29.1240
325	29.3851		326	29.6481		327	29.9130		328	30.1798
329	30.4486		330	30.7193		331	30.9920		332	31.2667
333	31.5434		334	31.8220		335	32.1027		336	32.3854
337	32.6702		338	32.9570		339	33.2459		340	33.5369
341	33.8300		342	34.1251		343	34.4224		344	34.7219
345	35.0235		346	35.3272		347	35.6332		348	35.9413
349	36.2516		350	36.5642		351	36.8790		352	37.1960
353	37.5153		354	37.8369		355	38.1608		356	38.4870
357	38.8155		358	39.1463		359	39.4795		360	39.8151
361	40.1530		362	40.4933		363	40.8361		364	41.1813
365	41.5289		366	41.8790		367	42.2316		368	42.5866
369	42.9442		370	43.3043		371	43.6669		372	44.0321
373	44.3998		374	44.7702		375	45.1431		376	45.5187
377	45.8969		378	46.2778		379	46.6613		380	47.0475
381	47.4365		382	47.8281		383	48.2225		384	48.6197
385	49.0196		386	49.4224		387	49.8279		388	50.2363
389	50.6475		390	51.0616		391	51.4786		392	51.8985
393	52.3213		394	52.7470		395	53.1757		396	53.6074
397	54.0421		398	54.4798		399	54.9205		400	55.3643
401	55.8112		402	56.2611		403	56.7142		404	57.1704
405	57.6298		406	58.0923		407	58.5581		408	59.0270
409	59.4992		410	59.9747		411	60.4534		412	60.9354
413	61.4208		414	61.9094		415	62.4015		416	62.8969
417	63.3958		418	63.8980		419	64.4037		420	64.9129
421	65.4256		422	65.9418		423	66.4615		424	66.9848
425	67.5117		426	68.0422		427	68.5763		428	69.1140
429	69.6555		430	70.2006		431	70.7495		432	71.3021
433	71.8585		434	72.4187		435	72.9827		436	73.5505
437	74.1222		438	74.6978		439	75.2773		440	75.8608
441	76.4482		442	77.0396		443	77.6351		444	78.2346
445	78.8381		446	79.4458		447	80.0576		448	80.6735
449	81.2936		450	81.9179		451	82.5464		452	83.1792
453	83.8163		454	84.4577		455	85.1034		456	85.7535
457	86.4079		458	87.0668		459	87.7302		460	88.3980
461	89.0703		462	89.7472		463	90.4286		464	91.1147
465	91.8053		466	92.5006		467	93.2006		468	93.9053
469	94.6147		470	95.3289		471	96.0480		472	96.7718
473	97.5005		474	98.2341		475	98.9726		476	99.7161
477	100.4646		478	101.2181		479	101.9767		480	102.7403
481	103.5091		482	104.2830		483	105.0621		484	105.8464
485	106.6359		486	107.4308		487	108.2309		488	109.0364
489	109.8473		490	110.6637		491	111.4854		492	112.3127
493	113.1455		494	113.9838		495	114.8278		496	115.6773
497	116.5326		498	117.3935		499	118.2602		500	119.1326
501	120.0109		502	120.8950		503	121.7850		504	122.6809
505	123.5828		506	124.4907		507	125.4047		508	126.3247
509	127.2508		510	128.1831		511	129.1215		512	130.0662
513	131.0172		514	131.9745		515	132.9381		516	133.9082
517	134.8847		518	135.8676		519	136.8571		520	137.8531
521	138.8557		522	139.8650		523	140.8810		524	141.9037
525	142.9331		526	143.9694		527	145.0125		528	146.0625
529	147.1195		530	148.1835		531	149.2545		532	150.3326
533	151.4178		534	152.5101		535	153.6097		536	154.7166
537	155.8307		538	156.9523		539	158.0812		540	159.2175
541	160.3614		542	161.5128		543	162.6718		544	163.8384
545	165.0128		546	166.1948		547	167.3847		548	168.5824
549	169.7880		550	171.0015		551	172.2230		552	173.4526
553	174.6902		554	175.9360		555	177.1900		556	178.4522
557	179.7227		558	181.0016		559	182.2889		560	183.5846
561	184.8889		562	186.2017		563	187.5232		564	188.8533
565	190.1921		566	191.5398		567	192.8963		568	194.2617
569	195.6360		570	197.0194		571	198.4119		572	199.8134
573	201.2242		574	202.6442		575	204.0735		576	205.5122
577	206.9603		578	208.4179		579	209.8851		580	211.3618
581	212.8482		582	214.3444		583	215.8503		584	217.3661
585	218.8919		586	220.4276		587	221.9733		588	223.5292
589	225.0952		590	226.6715		591	228.2581		592	229.8550
593	231.4624		594	233.0803		595	234.7088		596	236.3479
597	237.9977		598	239.6583		599	241.3297		600	243.0120
601	244.7054		602	246.4097		603	248.1252		604	249.8519
605	251.5899		606	253.3392		607	255.0999		608	256.8721
609	258.6559		610	260.4512		611	262.2583		612	264.0772
613	265.9079		614	267.7506		615	269.6052		616	271.4720
617	273.3509		618	275.2420		619	277.1455		620	279.0614
621	280.9897		622	282.9306		623	284.8841		624	286.8504
625	288.8294		626	290.8213		627	292.8262		628	294.8442
629	296.8752		630	298.9195		631	300.9770		632	303.0480
633	305.1324		634	307.2304		635	309.3420		636	311.4673
637	313.6065		638	315.7595		639	317.9266		640	320.1077
641	322.3030		642	324.5126		643	326.7365		644	328.9749
645	331.2278		646	333.4953		647	335.7776		648	338.0747
649	340.3867		650	342.7137		651	345.0558		652	347.4131
653	349.7858		654	352.1738		655	354.5773		656	356.9964
657	359.4312		658	361.8818		659	364.3483		660	366.8308
661	369.3294		662	371.8442		663	374.3754		664	376.9229
665	379.4869		666	382.0676		667	384.6650		668	387.2793
669	389.9105		670	392.5587		671	395.2241		672	397.9068
673	400.6069		674	403.3245		675	406.0596		676	408.8125
677	411.5833		678	414.3719		679	417.1787		680	420.0036
681	422.8468		682	425.7085		683	428.5886		684	431.4875
685	434.4051		686	437.3415		687	440.2970		688	443.2717
689	446.2655		690	449.2788		691	452.3116		692	455.3640
693	458.4361		694	461.5282		695	464.6402		696	467.7724
697	470.9249		698	474.0977		699	477.2911		700	480.5052
701	483.7400		702	486.9958		703	490.2726		704	493.5706
705	496.8900		706	500.2308		707	503.5932		708	506.9774
709	510.3835		710	513.8116		711	517.2619		712	520.7344
713	524.2294		714	527.7471		715	531.2874		716	534.8507
717	538.4370		718	542.0465		719	545.6793		720	549.3356
721	553.0155		722	556.7192		723	560.4469		724	564.1986
725	567.9746		726	571.7750		727	575.6000		728	579.4497
729	583.3242		730	587.2238		731	591.1486		732	595.0988
733	599.0744		734	603.0758		735	607.1030		736	611.1563
737	615.2357		738	619.3415		739	623.4738		740	627.6328
741	631.8187		742	636.0316		743	640.2717		744	644.5392
745	648.8343		746	653.1571		747	657.5079		748	661.8867
749	666.2939		750	670.7295		751	675.1937		752	679.6868
753	684.2089		754	688.7602		755	693.3409		756	697.9512
757	702.5913		758	707.2613		759	711.9615		760	716.6921
761	721.4531		762	726.2450		763	731.0678		764	735.9217
765	740.8070		766	745.7238		767	750.6723		768	755.6529
769	760.6655		770	765.7106		771	770.7882		772	775.8986
773	781.0420		774	786.2187		775	791.4287		776	796.6724
777	801.9500		778	807.2616		779	812.6075		780	817.9880
781	823.4031		782	828.8533		783	834.3386		784	839.8594
785	845.4158		786	851.0081		787	856.6365		788	862.3012
789	868.0025		790	873.7407		791	879.5158		792	885.3283
793	891.1783		794	897.0661		795	902.9919		796	908.9559
797	914.9585		798	920.9998		799	927.0801		800	933.1997
801	939.3588		802	945.5577		803	951.7966		804	958.0758
805	964.3956		806	970.7561		807	977.1578		808	983.6008
809	990.0853		810	996.6118		811	1003.1800		812	1009.7910
813	1016.4450		814	1023.1420		815	1029.8820		816	1036.6650
817	1043.4930		818	1050.3640		819	1057.2800		820	1064.2400
821	1071.2460		822	1078.2960		823	1085.3920		824	1092.5340
825	1099.7220		826	1106.9570		827	1114.2380		828	1121.5670
829	1128.9420		830	1136.3660		831	1143.8370		832	1151.3570
833	1158.9250		834	1166.5420		835	1174.2080		836	1181.9240
837	1189.6890		838	1197.5050		839	1205.3710		840	1213.2890
841	1221.2570		842	1229.2770		843	1237.3480		844	1245.4720
845	1253.6480		846	1261.8770		847	1270.1600		848	1278.4950
849	1286.8850		850	1295.3290		851	1303.8270		852	1312.3810
853	1320.9900		854	1329.6540		855	1338.3740		856	1347.1510
857	1355.9840		858	1364.8750		859	1373.8230		860	1382.8290
861	1391.8930		862	1401.0160		863	1410.1970		864	1419.4380
865	1428.7390		866	1438.1000		867	1447.5220		868	1457.0040
869	1466.5480		870	1476.1530		871	1485.8210		872	1495.5510
873	1505.3440		874	1515.2010		875	1525.1210		876	1535.1050
877	1545.1540		878	1555.2680		879	1565.4470		880	1575.6930
881	1586.0040		882	1596.3820		883	1606.8280		884	1617.3410
885	1627.9220		886	1638.5710		887	1649.2900		888	1660.0780
889	1670.9350		890	1681.8630		891	1692.8620		892	1703.9310
893	1715.0730		894	1726.2860		895	1737.5730		896	1748.9320
897	1760.3650		898	1771.8720		899	1783.4530		900	1795.1090
901	1806.8410		902	1818.6490		903	1830.5330		904	1842.4940
905	1854.5330		906	1866.6500		907	1878.8450		908	1891.1190
909	1903.4730		910	1915.9060		911	1928.4200		912	1941.0160
913	1953.6930		914	1966.4520		915	1979.2940		916	1992.2190
917	2005.2270		918	2018.3200		919	2031.4980		920	2044.7620
921	2058.1110		922	2071.5470		923	2085.0700		924	2098.6800
925	2112.3790		926	2126.1670		927	2140.0440		928	2154.0110
929	2168.0690		930	2182.2170		931	2196.4580		932	2210.7910
933	2225.2170		934	2239.7360		935	2254.3500		936	2269.0580
937	2283.8620		938	2298.7620		939	2313.7590		940	2328.8530
941	2344.0450		942	2359.3350		943	2374.7250		944	2390.2140
945	2405.8040		946	2421.4960		947	2437.2890		948	2453.1850
949	2469.1840		950	2485.2860		951	2501.4940		952	2517.8060
953	2534.2250		954	2550.7500		955	2567.3820		956	2584.1230
957	2600.9720		958	2617.9310		959	2634.9990		960	2652.1790
961	2669.4710		962	2686.8740		963	2704.3910		964	2722.0220
965	2739.7670		966	2757.6270		967	2775.6040		968	2793.6970
969	2811.9080		970	2830.2380		971	2848.6870		972	2867.2550
973	2885.9440		974	2904.7550		975	2923.6880		976	2942.7450
977	2961.9250		978	2981.2300		979	3000.6600		980	3020.2170
981	3039.9020		982	3059.7140		983	3079.6550		984	3099.7260
985	3119.9270		986	3140.2600		987	3160.7260		988	3181.3240
989	3202.0570		990	3222.9240		991	3243.9280		992	3265.0680
993	3286.3460		994	3307.7620		995	3329.3180		996	3351.0140
997	3372.8520		998	3394.8310		999	3416.9540		1000	3439.2210
1001	3461.6330		1002	3484.1910		1003	3506.8970		1004	3529.7500
1005	3552.7520		1006	3575.9030		1007	3599.2060		1008	3622.6610
1009	3646.2680		1010	3670.0300		1011	3693.9460		1012	3718.0180
1013	3742.2480		1014	3766.6350		1015	3791.1810		1016	3815.8880
1017	3840.7550		1018	3865.7850		1019	3890.9780		1020	3916.3350
1021	3941.8580		1022	3967.5470		1023	3993.4040

C Measuring the Accuracy With Which a Display System Matches the Grayscale Standard Display Function (Informative)
C.1 General Considerations Regarding Conformance and Metrics
To demonstrate conformance with the Grayscale Standard Display Function is a much more complex task than, for example, validating the responses of a totally digital system to DICOM messages.
Display systems ultimately produce analog output, either directly as Luminances or indirectly as optical densities. For some Display Systems, this analog output can be affected by various imperfections in addition to whatever imperfections exist in the Display System's Display Function that is to be validated. For example, there may be spatial non-uniformities in the final presented image (e.g., arising from film, printing, or processing non-uniformities in the case of a hardcopy printer) that are measurable but are at low spatial frequencies that do not ordinarily pose an image quality problem in diagnostic radiology.
It is worth noting that CRTs and light-boxes also introduce their own spatial non-uniformities. These non-uniformities are outside the scope of the Grayscale Standard Display Function and the measurement procedures described here. But because of them, even a test image that is perfectly presented in terms of the Grayscale Standard Display Function will be less than perfectly perceived on a real CRT or a real light-box.
Furthermore, the question "How close (to the Grayscale Standard Display Function) is close enough?" is currently unanswered, since the answer depends on psychophysical studies not yet done to determine what difference in Display Function is "just noticeable" when two nearly identical image presentations (e.g., two nearly identical films placed on equivalent side-by-side light-boxes) are presented to an observer.
Furthermore, the evaluation of a given Display System could be based either on visual tests (e.g., assessing the perceived contrast of many low-contrast targets in one or more test images) or by quantitative analysis based on measured data obtained from instruments (e.g., photometers or densitometers).
Even the quantitative approach could be addressed in different ways. One could, for example, simply superimpose plots of measured and theoretical analog output (i.e., Luminance or optical density) vs. P-Value, perhaps along with "error bars" indicating the expected uncertainty (non-repeatable variations) in the measured output. As a mathematically more elegant alternative, all the measured data points could be used as input to a statistical mathematical analysis that could attempt to determine the underlying Display Function of the Display System, yielding one or more quantitative values (metrics) that define how well the Display System conforms with the Grayscale Standard Display Function.
In what follows in this and the following annexes, an example of the latter type of metric analysis is used, in which measured data is analyzed using a "FIT" test that is intended to validate the shape of the Characteristic Curve and a "LUM" test that is intended to show the degree of scatter from the ideal Grayscale Standard Display Function. This approach has been applied, for example, to quantitatively demonstrate how improvements were successfully made to the Display Function of certain Display Systems.
Before proceeding with the description of the methodology of this specific metric approach, it should be noted that it is offered as one possible approach, not necessarily as the most appropriate approach for evaluating all Display Systems. In particular, the following notes should be considered before selecting or interpreting results from any particular metric approach.
1. There may be practical issues that limit the number of P-Values that can be meaningfully used in the analysis. For example, it may be practical to measure all 256 possible Luminances from a fixed position on the screen of an 8-bit video monitor, but it may be impractical to meaningfully measure all 4096 densities theoretically printable by a 12-bit film printer. One reason for the impracticality is the limited accuracy of densitometers (or even film digitizers). A second reason is that the film density measurements, unlike the CRT photometer measurements, are obtained from different locations on the display area, so any spatial non-uniformity that is present in the film affects the hardcopy measurement. Current hardcopy printers and densitometers both have absolute optical density accuracy limitations that are significantly worse than the change that would be caused by a change in just the least significant bit of a 12-bit P-Value. In general, selecting a larger number of P-Values allows, in principle, more localized aberrations from the Grayscale Standard Display Function to be "caught", but the signal-to-noise ratio (or significance) of each of these will be decreased.
2. If the measurement data for a particular Display System has significant "noise" (as indicated by limited repeatability in the data when multiple sets of measurements are taken), it may be desirable to apply a statistical analysis technique that goes beyond the "FIT" and "LUM" metric by explicitly utilizing the known standard deviations in the input data set, along with the data itself, to prevent the fitting technique from over-reacting to noise. See, for example, the section "General Linear Least Squares" in Reference C1 and the chapter "Least-Squares Fit to a Polynomial" in Reference C2. If measurement noise is not explicitly taken into account in the analysis, the metric's returned root-mean-square error of the data points relative to the fit could be misleadingly high, since it would include the combined effect of errors due to incorrectness in the Display Function and errors due to measurement noise.
3. If possible, the sensitivity and specificity of the metric being considered should be checked against visual tests. For example, a digital test pattern with many low-contrast steps at many ambient Luminances could be printed on a "laboratory standard" Grayscale Standard Display Function printer and also printed on a printer being evaluated. The resultant films could then be placed side-by-side on light-boxes for comparison by a human observer. A good metric technique should detect as sensitively and repeatably as the human observer the existence of deviations (of any shape) from the Grayscale Standard Display Function. For example, if a Display System has a Characteristic Curve that, for even a very short interval of DDL values, is too contrasty, too flat, or (worse yet) non-monotonic, the metric should be able to detect and respond to that anomaly as strongly as the human observer does.
4. Finally, in addition to the experimentally encountered non-repeatabilities in the data from a Display System, there may be reason to consider additional possible causes of variations. For example, varying the ordering of P-Values in a test pattern (temporally for CRTs, spatially for printers) might affect the results. For printers, switching to different media might affect the results. A higher confidence can be placed in the results obtained from any metric if the results are stable in the presence of any or all such changes.
C.2 Methodology
Step (1)
The Characteristic Curve of the test Display System should be determined with as many measurements as practical (see Section D.1, Section D.2, and Section D.3). Using the Grayscale Standard Display Function, the fractional number of JNDs are calculated for each Luminance interval between equally spaced P-Value steps. The JNDs/Luminance interval may be calculated directly, or iteratively. For example, if only a few JNDs belong to every Luminance interval, a linear interpolation may be performed. After transformation of the grayscale response of the Display System, the Luminance Levels for every P-Value are Li and the corresponding Standard Luminance Levels are Lj; dj specifies the JNDs /Luminance-Interval on the Grayscale Standard Display Function for the given number of P-Values. Then, the JNDs/Luminance interval for the transformed Display Function are

r = dj(Li+1 - Li)(Lj+1 + Lj) / ((Li+1 + Li)(Lj+1 - Lj))	(C-1)

Additionally, an iterative method can be used to calculate the number of JNDs per Luminance interval, requiring only the Grayscale Standard Display Function that defines a JND step in Luminance given a Luminance value. This is done by simply counting the number of complete JND steps in the Luminance interval, and then the remaining fractional step. Start at the Luminance low end of the interval, and calculate from the Grayscale Standard Display Function the Luminance step required for one JND step. Then continue stepping from the low Luminance value to the high Luminance value in single JND steps, until the Luminance value of the upper end of the Luminance Range is passed. Calculate the fraction portion of one JND that this last step represents. the total number of completed integer JND steps plus the fractional portion of the last uncompleted step is the fractional number of JND steps in the Luminance interval.
Plot the number of JNDs per Luminance interval (vertical axis) versus the index of the Luminance interval (horizontal axis). This curve is referred to as the Luminance intervals vs JNDscurve. An example of a plot of Luminance intervals vs JNDs is shown in figure C-1. The plot is matched very well by a horizontal line when a linear regression is applied.

Figure C-1. Illustration for the LUM and FIT conformance measures
The JNDs/Luminance interval data are evaluated by two statistical measures [C4]. The first assesses the global match of the test Display Function with the Grayscale Standard Display Function. The second measure locally analyses the approximation of the Grayscale Standard Display Function to the test Display Function.
Step (2)
Two related measures of a regression analysis are applied after normal multiple linear regression assumptions are verified for the data [C3].The first measure, named the FITtest, attempts to match the Luminance-Intervals-vs-JNDs curve of the test Luminance distribution with different order polynomial fits. The Grayscale Standard Display Function is characterized by exactly one JND per Luminance interval over the entire Luminance Range. Therefore, ideally, the data of JNDs/Luminance intervals vs index of the Luminance interval are best fit by a horizontal line of a constant number of JNDs/Luminance interval, indicating that both the local and global means of JNDs/Luminance interval are constant over the given Luminance Range. If the curve is better matched by a higher-order curve, the distribution is not closely approximating the Grayscale Standard Display Function. The regression analysis should test comparisons through third-order curves.
The second measure, the Luminance uniformity metric (LUM), analyzes whether the size of Luminance steps are uniform in perceptual size (i.e., JNDs) across the Luminance Range. This is measured by the Root Mean Square Error (RMSE) of the curve fit by a horizontal line of the JNDs/Luminance interval. The smaller the RMSE of the JNDs/Luminance interval, the more closely the test Display Function approximates the Grayscale Standard Display Function on a microscopic scale.
Both the FIT and LUM measures can be conveniently calculated on standard statistical packages.
Assuming the test Luminance distribution passes the FIT test, then the measure of quality of the distribution is determined by the single quantitative measurement (LUM) of the standard deviation of the JNDs/Luminance interval from their mean. Clinical practice is expected to determine the tolerances for the FIT and LUM values.
An important factor in reaching a close approximation of a test Display Function to the Grayscale Standard Display Function is the number of discrete output levels of the Display System. For instance, the LUM measure can be improved by using only a subset of the available DDLs while maintaining the full available output digitization resolution at the cost of decreasing contrast resolution.
While the LUM is influenced by the choice of the number of discrete output gray levels in the Grayscale Standard Display Function, the appropriate number of output levels is determined by the clinical application, including possible gray scale image processing that may occur independently of the Grayscale Standard Display Function standardization. Thus, PS3.14 does not prescribe a certain number of gray levels of output. However, in general, the larger the number of distinguishable gray levels available, the higher the possible image quality because the contrast resolution is increased. It is recommended that the number of necessary output driving levels for the transformed Display Function be determined prior to standardization of the Display System (based on clinical applications of the Display System), so that this information can be used when calculating the transformation in order to avoid using gray scale distributions with fewer output levels than needed.
C.3 References
[C1] Press, William H, et al., Numerical Recipes in C, Cambridge University Press, 1988, Section "General Linear Least Squares"
[C2] Bevington, Phillip R., Data Reduction and Error Analysis for the Physical Sciences, McGraw-Hill, 1969, the chapter "Least-Squares Fit to a Polynomial" .
[C3] Kleinbaum DG, Kupper LL, Muller KE, Applied Regression Analysis and Other Multivariable Methods, Duxbury Press, 2nd Edition, pp 45-49, 1987.
[C4] Hemminger, B., Muller, K., "Performance Metric for evaluating conformance of medical image displays with the ACR/NEMA display function standard", SPIE Medical Imaging 1997, editor Yongmin Kim, vol 3031-25, 1997.
D Illustrations for Achieving Conformance with the Grayscale Standard Display Function (Informative)
The following sections illustrate how conformance with the Grayscale Standard Display Function may be achieved for emissive (soft-copy) Display Systems as well as systems producing image presentations (hard-copies) on transmissive and reflective media. Each section contains four sub-sections on 1) a procedure for measuring the system Characteristic Curve, 2) the application of the Grayscale Standard Display Function to the Luminance Range of the Display System, 3) the implementation of the Grayscale Standard Display Function, and 4) the application of the conformance metrics as proposed in Annex C.
It is emphasized that there are different ways to configure a Display System or to change its performance so that it conforms to the Grayscale Standard Display Function. In fact, conceivably, a Display System may calibrate itself automatically to maintain conformance with the Standard. Hence, the following three illustrations are truly only examples.
Luminance of any Display System, hard-copy or softcopy, may be measured with a photometer. The photometer should have the following characteristics:
- be accurate to within 3% or less of the absolute Luminance level across its full range of operation;
- have a relative accuracy of at least two times the least significant digit at any Luminance level in its range of operation;
- maintain this accuracy at Luminance levels that are one-tenth of the minimum measured Luminance of the Display System;
- have an acceptance angle that is small enough to incorporate only the measurement field without overlapping the surrounding background.
Note
The photometer may be of the type that attaches directly to the display face (with a suction cup) or of the type that is held away from the display face. If of the latter type, the photometer should be well baffled to exclude extraneous light sources, including light from the background area of the test pattern.
For a film Display System the photometer may be appropriately used to measure the background Illuminance and the Luminance of the light-box on which the film will be displayed. The Luminance characteristics of the film Display System may be measured directly with the photometer or indirectly using measured optical density of the film and the values for the measured background Illuminance and the light-box Luminance.
D.1 Emissive Display Systems
D.1.1 Measuring the System Characteristic Curve
Before the characteristic Luminance response of the emissive Display System is measured, it is allowed to warm up as recommended by the manufacturer and is adjusted such that it conforms to the manufacturer's performance specifications. In particular, adjustment procedures for setting the black and white levels of the display should be obtained from the Display System manufacturer. The goal is to maximize the dynamic Luminance Range of the display without introducing artifacts, resulting in the highest possible number of Just-Noticeable Differences (JNDs).
Note
A simple test that the system is set up properly can be performed by viewing the 5% and 95% squares in the SMPTE pattern. The perceived contrast between the 5% square and its 0% surrounding should be equal to the perceived contrast between the 95% square and a white square.
Measurement of the Characteristic Curve of the Display System may be accomplished using a test pattern (Figure D.1-1) consisting of:
- a square measurement field comprising 10% of the total number of pixels displayed by the system positioned in the center of the display;
- a full-screen uniform background of 20% of maximum Luminance surrounding the target.
Note
With a measurement field of 10% of the total number of displayed pixels and a surrounding set to 20% of maximum Luminance, internal light scatter in the monitor causes the Luminance Range to be typically comparable to that found in radiographs, such as a thorax radiograph, when displayed on the CRT monitor.

Figure D.1-1. The test pattern will be a variable intensity square in the center of a low Luminance background area.
Note
1. For example, on a 5-megapixel Display System with a matrix of 2048 by 2560 pixels, the target would be a square with 724 pixels on each side.
2. Ideally, the test pattern should fill the entire screen. Under certain windowed operating environments, it may be difficult to eliminate certain user-interface objects from the display, in particular, menu bars at the top of the screen. In this case, the background should fill as much of the screen as possible.
The Characteristic Curve of the Display System may be determined by
- turning off all ambient lighting (necessary only when a suction cup photometer is used or when a handheld photometer casts a shadow on the display screen);
- displaying the above test pattern;
- setting the DDL for the measurement field to a sequence of different values, starting with 0 and increasing at each step until the maximum DDL is reached;
- using a photometer to measure and record the Luminance of the measurement field at each command value.
As discussed in Annex C, the number and distribution of DDLs at which measurements are taken must be sufficient to accurately model the Characteristic Curve of the Display System over the entire Luminance Range.
Note
1. If a handheld photometer is used, it should be placed at a distance from the display screen so that Luminance is measured in the center of the measurement field, without overlapping the surrounding background. This distance can be calculated using the acceptance angle specification provided by the photometer manufacturer.
2. The exact number and distribution of DDLs should be based both on the characteristics of the Display System and on the mathematical technique used to interpolate the Characteristic Curve of the system. It is recommended that at least 64 different command values be used in the procedure.
3. Successive Luminance measurements should be spaced in time such that the Display System always reaches a steady state. It may be particularly important to allow the system to settle before taking the initial measurement at DDL 0.
As stated in the normative section, the effect of ambient light on the apparent Characteristic Curve must always be included when configuring a Display System to conform with the Grayscale Standard Display Function.
If a handheld photometer that does not cast a shadow on the display screen is used to measure the Characteristic Curve, then the Luminance produced by the display plus the effect of ambient light may be measured simultaneously.
When a suction cup photometer is used to take the Luminance measurements or when a handheld photometer casts a shadow on the display screen, all ambient lighting should be turned off while measuring the Characteristic Curve. The effect of ambient light is determined separately: The Display System is turned off, the ambient light is turned on, and the Luminance produced by scattering of ambient light at the display screen is measured by placing the photometer at a distance from the display screen so that its acceptance angle includes a major portion of the screen and that the measurement is not affected by direct illumination from areas outside the display screen. The Luminance related to ambient light is added to the previously measured Luminance levels produced by the Display System to determine the effective Characteristic Curve of the system.
Note
Changes in ambient lighting conditions may require recalibration of the display subsystem in order to maintain conformance to this standard.
In the following, an example for measurements and transformation of a Display Function is presented. The Display System for this example is a CRT monitor with display controller. It is assumed that the display controller allows a transformation of the DDLs with 8-bit input precision and 10-bit output precision.
The Luminance is measured with a photometer with a narrow (1°) acceptance angle. The ambient light level was adjusted as low as possible. No localized highlights were visible.
1. The maximum Luminance was measured when setting the DDL for the measurement field to the value that yielded the highest Luminance and the DDL of the surrounding to the middle DDL range. From this measurement, the Luminance - 20% of the maximum Luminance - for the surrounding of the measurement field was calculated.
2. The ambient light was turned off. With the photometer centered on the measurement field of the test pattern of Figure D.1-1, the Luminance was measured when varying the input level Dm in increments of 1 from 0 to 255. The transformation operator of the hypothetical display controller linearly mapped 8 bits on the input to 10 bits on the output. The measured data represent the Characteristic Curve L = F(Dm) for the given operating conditions and this test pattern.
3. Next, the CRT was turned off and the ambient light turned on. The photometer was placed on the center axis of the CRT sufficiently far away so that it did not cast a shadow on the CRT face and its aperture intercepted light scattered from a major portion of the CRT face. The measured Luminance of 0.3 cd/m2produced by the ambient light on the CRT face was added to the measured Luminance values of the Characteristic Curve without ambient light. The result is listed in Table D.1-1 and plotted in Figure D.1-2.

Figure D.1-2. Measured Characteristic Curve with Ambient Light of an emissive Display System
Table D.1-1. Measured Characteristic Curve plus Ambient Light
DDL	Luminance		DDL	Luminance		DDL	Luminance		DDL	Luminance
0	0.305		1	0.305		2	0.305		3	0.305
4	0.305		5	0.305		6	0.305		7	0.305
8	0.305		9	0.305		10	0.305		11	0.307
12	0.307		13	0.307		14	0.307		15	0.307
16	0.307		17	0.307		18	0.307		19	0.307
20	0.307		21	0.307		22	0.310		23	0.310
24	0.310		25	0.310		26	0.310		27	0.320
28	0.320		29	0.320		30	0.330		31	0.330
32	0.340		33	0.350		34	0.360		35	0.370
36	0.380		37	0.392		38	0.410		39	0.424
40	0.442		41	0.464		42	0.486		43	0.512
44	0.534		45	0.562		46	0.594		47	0.626
48	0.674		49	0.710		50	0.750		51	0.796
52	0.842		53	0.888		54	0.938		55	0.994
56	1.048		57	1.108		58	1.168		59	1.232
60	1.294		61	1.366		62	1.438		63	1.512
64	1.620		65	1.702		66	1.788		67	1.876
68	1.960		69	2.056		70	2.154		71	2.248
72	2.350		73	2.456		74	2.564		75	2.670
76	2.790		77	2.908		78	3.022		79	3.146
80	3.328		81	3.460		82	3.584		83	3.732
84	3.870		85	4.006		86	4.156		87	4.310
88	4.456		89	4.608		90	4.766		91	4.944
92	5.104		93	5.268		94	5.444		95	5.630
96	5.864		97	6.050		98	6.238		99	6.438
100	6.610		101	6.820		102	7.024		103	7.224
104	7.428		105	7.644		106	7.872		107	8.066
108	8.298		109	8.528		110	8.752		111	8.982
112	9.330		113	9.574		114	9.796		115	10.060
116	10.314		117	10.560		118	10.820		119	11.080
120	11.340		121	11.620		122	11.880		123	12.180
124	12.460		125	12.700		126	13.020		127	13.300
128	13.720		129	14.020		130	14.360		131	14.640
132	14.940		133	15.300		134	15.600		135	15.900
136	16.240		137	16.560		138	16.920		139	17.220
140	17.600		141	17.940		142	18.240		143	18.640
144	19.120		145	19.460		146	19.800		147	20.260
148	20.560		149	20.920		150	21.360		151	21.760
152	22.060		153	22.520		154	22.960		155	23.300
156	23.700		157	24.080		158	24.600		159	24.980
160	25.520		161	26.040		162	26.480		163	26.700
164	27.380		165	27.620		166	28.040		167	28.580
168	28.980		169	29.400		170	29.840		171	30.540
172	30.800		173	31.380		174	31.880		175	32.400
176	33.060		177	33.400		178	34.040		179	34.400
180	34.840		181	35.360		182	35.900		183	36.400
184	37.060		185	37.400		186	38.300		187	38.420
188	39.160		189	39.760		190	39.980		191	40.840
192	41.540		193	41.900		194	42.800		195	43.060
196	43.620		197	44.520		198	44.620		199	45.500
200	46.100		201	46.380		202	47.400		203	47.600
204	48.320		205	49.060		206	49.380		207	50.320
208	50.920		209	51.600		210	52.420		211	52.680
212	53.520		213	54.220		214	54.620		215	55.420
216	56.100		217	56.600		218	57.400		219	57.820
220	58.660		221	59.320		222	59.800		223	60.720
224	61.520		225	62.240		226	63.040		227	63.480
228	64.460		229	65.020		230	65.500		231	66.500
232	66.960		233	67.840		234	68.600		235	68.980
236	70.040		237	70.520		238	71.420		239	72.180
240	72.900		241	73.980		242	74.580		243	75.320
244	76.200		245	76.540		246	77.720		247	78.220
248	79.200		249	79.880		250	80.420		251	81.560
252	81.960		253	83.140		254	83.720		255	84.340

D.1.2 Application of the Standard Formula
The section of the Grayscale Standard Display Function for the Luminance Range of the CRT monitor Display System is shown in Figure D.1-3. Minimum and maximum Luminance levels correspond to JND indices of JNDmin = 32.54 and JNDmax = 453.85, respectively. Thus, there are theoretically about 420 just-noticeable Luminance differences for the Standard Target (see Normative Section 6). Obviously, with 8-bit input digitization resolution, at best 256 noticeable Luminance increments can be realized.
D.1.3 Implementation of the Standard
The measured Characteristic Curve is interpolated for the available output levels Doutput, in this case, yielding 1024 Luminance levels LI,m. The Grayscale Standard Display Function is also interpolated between JNDmin and JNDmax ((JND= [JNDmax - JNDmin]/1023 = [453.85 - 32.54]/1023) yielding 1024 Standard Luminance levels LI,STD. Interpolations can be performed by a variety of techniques. Here, a cubic spline technique was employed.
For every LI,STD, the closest LJ,m is determined. The data pair I,J defines the transformation between Dinput and Doutput (Table D.1-2) by which the Luminance response of the Display System is made to approximate the Grayscale Standard Display Function.
Table D.1-2. Look-Up Table for Calibrating Display System
Input	Output		Input	Output		Input	Output		Input	Output
0	0		1	118		2	131		3	140
4	148		5	153		6	160		7	164
8	169		9	173		10	178		11	182
12	185		13	189		14	191		15	194
16	198		17	201		18	204		19	207
20	210		21	214		22	217		23	219
24	222		25	225		26	228		27	231
28	234		29	237		30	240		31	243
32	245		33	248		34	251		35	253
36	255		37	257		38	260		39	263
40	265		41	268		42	271		43	274
44	276		45	279		46	282		47	284
48	287		49	290		50	292		51	295
52	298		53	301		54	303		55	306
56	308		57	311		58	314		59	317
60	319		61	320		62	323		63	326
64	329		65	331		66	334		67	336
68	339		69	342		70	345		71	347
72	350		73	353		74	356		75	359
76	361		77	364		78	367		79	370
80	372		81	375		82	378		83	381
84	383		85	385		86	388		87	391
88	393		89	396		90	399		91	402
92	405		93	407		94	410		95	413
96	416		97	419		98	422		99	425
100	428		101	431		102	434		103	437
104	440		105	443		106	445		107	448
108	450		109	452		110	456		111	459
112	462		113	465		114	468		115	471
116	474		117	477		118	480		119	483
120	486		121	490		122	492		123	495
124	499		125	502		126	505		127	509
128	511		129	513		130	516		131	519
132	522		133	526		134	529		135	532
136	535		137	539		138	542		139	545
140	549		141	552		142	555		143	559
144	562		145	565		146	569		147	572
148	575		149	578		150	581		151	585
152	588		153	591		154	595		155	599
156	602		157	605		158	609		159	613
160	616		161	619		162	623		163	627
164	631		165	633		166	637		167	640
168	643		169	646		170	650		171	655
172	657		173	663		174	666		175	669
176	674		177	678		178	682		179	684
180	688		181	693		182	696		183	700
184	703		185	706		186	711		187	714
188	719		189	723		190	727		191	731
192	735		193	738		194	743		195	745
196	752		197	754		198	758		199	764
200	766		201	769		202	775		203	777
204	783		205	787		206	789		207	796
208	799		209	805		210	808		211	811
212	818		213	821		214	827		215	830
216	834		217	838		218	841		219	848
220	851		221	856		222	861		223	864
224	870		225	874		226	880		227	883
228	889		229	893		230	897		231	901
232	905		233	911		234	915		235	922
236	925		237	931		238	935		239	941
240	945		241	951		242	955		243	960
244	964		245	969		246	975		247	979
248	985		249	991		250	995		251	1002
252	1006		253	1012		254	1016		255	1023

D.1.4 Measures of Conformance
The FIT and the LUM metrics proposed in Annex C are applied to determine the macroscopic and microscopic approximation of the L J,mto the L I,STD. Figure D.1-3 shows the perceptually linearized Display Function superimposed on the Grayscale Standard Display Function and Figure D.1-4 summarizes the results of the two metrics. A good global fit was achieved as demonstrated by the nearly horizontal-line fit as best fit obtained with the FIT metric. The RMSE is acceptable. All 255 P-Value intervals lead to JNDs on the transformed Display Function for the Standard Target.

Figure D.1-3. Measured and interpolated Characteristic Curve, Grayscale Standard Display Function and transformed Display Function of an emissive Display System. The transformed Display Function for this Display System matches the Grayscale Standard Display Function and the two curves are superimposed and indistinguishable.

Figure D.1-4. LUM and FIT measures of conformance for a the transformed Display Function of an emissive Display System
D.2 Transparent Hardcopy Devices
D.2.1 Measuring the System Characteristic Curve
A transparent hardcopy device is exemplified by a laser printer (including processor) that prints (exposes and processes) one or more images on a sheet of transparent film (typically a 14" x 17" film). This film is eventually placed over a high Luminance light-box in a darkened room for viewing.
The Characteristic Curve for such a transparent hardcopy device is obtained by printing a test image consisting of a pattern of n bars, each bar having a specific numeric value (DDL). The optical density of each printed bar is then measured, using a transmission densitometer, for each of the printed bars.
To accurately define a printer's Characteristic Curve, it is desirable that n be as large as possible (to capture as many points as possible on the Characteristic Curve). However, the limitations on absolute quantitative repeatability imposed by the printer, processor, or media technologies may dictate that a much smaller value of n be used (to prevent a conformance metric that is sensitive to differences from becoming unstable and meaningless, as the density differences between adjacent bars become "in the noise" as the number of bars becomes large).
One example of a test image is a pattern of 32 approximately equal-height bars, spanning the usable printable region of the film, having 32 approximately equi-spaced DDLs as follows:

Figure D.2-1. Layout of a Test Pattern for Transparent Hardcopy Media
To define a test pattern with n DDLs for a printer with an N-bit input, the DDL of step # i can be set to

DDLi = (2N-1)i/(n-1)	(D.2-1)

rounded to the nearest integer.
The tabulated values of DDLi and the corresponding measured optical densities ODi constitute a Characteristic Curve of the printer.
D.2.2 Application of the Grayscale Standard Display Function
The films that are produced by transparent hardcopy printers are often brought to a variety of locations, where they may be viewed on different light-boxes and under a variety of viewing conditions. Accordingly, the approach of PS3.14 is to define, for hardcopy transparent printers, what densities (rather than Luminances) should be produced, and to provide here a method of applying the Grayscale Standard Display Function to the transparent hardcopy case, based on parameters that are typical of the expected range of light-box Luminances and other viewing parameters.
The specific parameters that are used in the following example are as follows:
L0 (Luminance of light-box with no film present): 2000 cd/m2
La (ambient room light reflected by film): 10 cd/m2
Dmin (minimum optical density obtainable on film): 0.20
Dmax (maximum optical density desirable on film): 3.00.
The process of constructing a table of desired OD values from the Grayscale Standard Display Function begins with defining the Luminance Range and the corresponding range of the Just-Noticeable Difference Index, j. The minimum and maximum Luminance values are given respectively by

Lmin= La + L010-Dmax = 12.0 cd/m2	(D.2-2)

Lmax= La + L010-Dmin = 1271.9 cd/m2	(D.2-3)

Next, calculate the corresponding Just-Noticeable Difference Index values, jmin and jmax. For the current example, we obtain

jmin = 233.32	(D.2-4)

jmax = 848.75	(D.2-5)

This gives us the range of j-values that the printer should cover. The printer should map its minimum input (P-Value = 0) to jmin and the corresponding Lmin. It should map its maximum input (P-Value = 2N-1 where N is the number of input bits) to jmax and the corresponding Lmax. At any intermediate input it should map its input proportionately:

j(PV) = jmin + (jmax-jmin)	(D.2-6)

and target values for the Luminance given by the Standard's formula: L(j(P-Value)). This "targeting" consists of producing an optical density OD for this P-Value that will give the desired Luminance L(j(P-Value)) under the conditions of L0 and La previously defined. The required density can thus be calculated as follows:

(D.2-7)

D.2.3 Implementation of the Grayscale Standard Display Function
Carrying this example into the even more specific case of a printer with an 8-bit input leads to the following table, which defines the OD's to be generated for each of the 256 possible P-Values.
Table D.2-1. Optical Densities for Each P-Value for an 8-Bit Printer
P-Value	Optical Density (OD)		P-Value	Optical Density (OD)		P-Value	Optical Density (OD)		P-Value	Optical Density (OD)
0	3.000		1	2.936		2	2.880		3	2.828
4	2.782		5	2.739		6	2.700		7	2.662
8	2.628		9	2.595		10	2.564		11	2.534
12	2.506		13	2.479		14	2.454		15	2.429
16	2.405		17	2.382		18	2.360		19	2.338
20	2.317		21	2.297		22	2.277		23	2.258
24	2.239		25	2.221		26	2.203		27	2.185
28	2.168		29	2.152		30	2.135		31	2.119
32	2.103		33	2.088		34	2.073		35	2.058
36	2.043		37	2.028		38	2.014		39	2.000
40	1.986		41	1.973		42	1.959		43	1.946
44	1.933		45	1.920		46	1.907		47	1.894
48	1.882		49	1.870		50	1.857		51	1.845
52	1.833		53	1.821		54	1.810		55	1.798
56	1.787		57	1.775		58	1.764		59	1.753
60	1.742		61	1.731		62	1.720		63	1.709
64	1.698		65	1.688		66	1.677		67	1.667
68	1.656		69	1.646		70	1.636		71	1.626
72	1.616		73	1.605		74	1.595		75	1.586
76	1.576		77	1.566		78	1.556		79	1.547
80	1.537		81	1.527		82	1.518		83	1.508
84	1.499		85	1.490		86	1.480		87	1.471
88	1.462		89	1.453		90	1.444		91	1.434
92	1.425		93	1.416		94	1.407		95	1.398
96	1.390		97	1.381		98	1.372		99	1.363
100	1.354		101	1.346		102	1.337		103	1.328
104	1.320		105	1.311		106	1.303		107	1.294
108	1.286		109	1.277		110	1.269		111	1.260
112	1.252		113	1.244		114	1.235		115	1.227
116	1.219		117	1.211		118	1.202		119	1.194
120	1.186		121	1.178		122	1.170		123	1.162
124	1.154		125	1.146		126	1.138		127	1.130
128	1.122		129	1.114		130	1.106		131	1.098
132	1.090		133	1.082		134	1.074		135	1.066
136	1.058		137	1.051		138	1.043		139	1.035
140	1.027		141	1.020		142	1.012		143	1.004
144	0.996		145	0.989		146	0.981		147	0.973
148	0.966		149	0.958		150	0.951		151	0.943
152	0.935		153	0.928		154	0.920		155	0.913
156	0.905		157	0.898		158	0.890		159	0.883
160	0.875		161	0.868		162	0.860		163	0.853
164	0.845		165	0.838		166	0.831		167	0.823
168	0.816		169	0.808		170	0.801		171	0.794
172	0.786		173	0.779		174	0.772		175	0.764
176	0.757		177	0.750		178	0.742		179	0.735
180	0.728		181	0.721		182	0.713		183	0.706
184	0.699		185	0.692		186	0.684		187	0.677
188	0.670		189	0.663		190	0.656		191	0.648
192	0.641		193	0.634		194	0.627		195	0.620
196	0.613		197	0.606		198	0.598		199	0.591
200	0.584		201	0.577		202	0.570		203	0.563
204	0.556		205	0.549		206	0.542		207	0.534
208	0.527		209	0.520		210	0.513		211	0.506
212	0.499		213	0.492		214	0.485		215	0.478
216	0.471		217	0.464		218	0.457		219	0.450
220	0.443		221	0.436		222	0.429		223	0.422
224	0.415		225	0.408		226	0.401		227	0.394
228	0.387		229	0.380		230	0.373		231	0.366
232	0.359		233	0.352		234	0.345		235	0.338
236	0.331		237	0.324		238	0.317		239	0.311
240	0.304		241	0.297		242	0.290		243	0.283
244	0.276		245	0.269		246	0.262		247	0.255
248	0.248		249	0.241		250	0.234		251	0.228
252	0.221		253	0.214		254	0.207		255	0.200

Plotting these values gives the curve of Figure D.2-3.

Figure D.2-3. Plot of OD vs P-Value for an 8-Bit Printer
D.2.4 Measures of Conformance
As an example, a bar pattern with 32 optical densities was printed on transmissive media (film). Beforehand, the printer had been set up to print over a density range from 0.2 (Dmin) to 3.0 (Dmax) and had been pre-configured by the manufacturer to use the Grayscale Standard Display Function, converted by the manufacturer into the table of target density values vs. P-Values described earlier.
The test pattern that was used for this was an 8-bit image consisting essentially of 32 horizontal bars. The 32 P-Values used for the bars were as follows: 0, 8, 16, 25, 33, 41, 49, 58, 66, 74, 82, 90,99, 107, 115, 123, 132, 140, 148, 156, 165, 173, 181, 189, 197, 206, 214,222, 230, 239, 247, 255.
For a given film, the 32 bars' optical densities were measured (near the middle of the film), converted to Luminances (using the standard parameters of light-box Luminance and reflected ambient light described earlier),and converted to Just-Noticeable Difference Indices by mathematically computing j(L) from L(j), where L(j) is the Grayscale Standard Display Function of Luminance L as a function of the Just-Noticeable Difference Index j. For each of the 31 intervals between consecutive measured values, a calculated value of "JNDs per increment in P-Values" was obtained by dividing the difference in Just-Noticeable Difference Index by the difference in P-Values for that interval. (In these calculations, density, L, and j are all floating-point variables. No rounding to integer values is done, so no truncation error is introduced.)
In this example, the film's data could be reasonably well fit by a horizontal straight line. That is, the calculated "JNDs per increment in P-Values was essentially constant at 2.4. A mathematical fit yielded a slight non-zero slope (specifically, dropping from 2.5 to 2.3 as the P-Value went from 0 to 255), but the 0.2 total difference was considerably smaller than the noise that was present in the 31 individual values of "JNDs per increment in P-Value" so is of doubtful significance. (The "noise" referred to here consists of the random, non-repeatable variations that are seen if a new set of measured data (e.g., from a second print of the same test pattern) is compared with a previous set of measurements.)
No visual tests were done to see if a slope that small could be detected by a human observer in side-by-side film comparisons.
Incidentally, if one considers just the 32 original absolute measured densities (rather than differential values based on small differences), one finds, in this case, quite reasonable agreement between the target and measured optical densities (within the manufacturer's norms for density accuracy, at a given density). But if one uses any metric that is based on differential information over small intervals, the results must be considered more cautiously, since they can be strongly affected by (and may be dominated by) various imperfections that are independent of a device's "true" (or averaged over many cases) characteristic behavior.
D.3 Reflective Display Systems
This last example illustrates how conformance with the Grayscale Standard Display Function may be achieved for a thermal-dye-transfer paper printer/office-light system. The thermal-dye-transfer printer produces black-and-white grayscale prints on a semi-glossy 8-inch x 10-inch heavy-gauge paper. The print is illuminated uniformly by fluorescent lamps so that the minimum reflective density produces a Luminance of 150 cd/m2. The hypothetical transformation operator is assumed to have equal input and output digitization resolution of 8 bits.
D.3.1 Measuring the System Characteristic Curve
A print with a 64-step grayscale tablet was printed for DDLs 4, 8, 12, ...,248, 252, 255. The reflection optical densities (from 0.08 to 2.80) were measured with a densitometer. The Luminance levels corresponding to the measured optical densities and illumination conditions are plotted in Figure D.3-1.

Figure D.3-1. Measured and interpolated Characteristic Curve and Grayscale Standard Display Function for a printer producing reflective hard-copies
D.3.2 Application of the Grayscale Standard Display Function
This last example illustrates how conformance with the Grayscale Standard Display Function may be achieved for a thermal-dye-transfer paper printer/office-light system. The thermal-dye-transfer printer produces black-and-white grayscale prints on a semi-glossy 8-inch x 10-inch heavy-gauge paper. The print is illuminated uniformly by fluorescent lamps so that the minimum reflective density produces a Luminance of 150 cd/m2. The hypothetical transformation operator is assumed to have equal input and output digitization resolution of 8 bits.
D.3.3 Implementation of the Grayscale Standard Display Function
The measured Characteristic Curve is interpolated for the available DDLs yielding 256 Luminance levels LI,m. The Grayscale Standard Display Function is also interpolated between JNDmin and JNDmax (DJND = [JNDmax - JNDmin]/255) yielding 256 Standard Luminance levels LI,STD.
For every LI,STD, the closest LJ,m is determined. The data pair I,J defines the transformation between Dinput and Doutput (Table D.3-1 and Figure D.3-2) by which the Luminance response of the Display System is made to approximates the Grayscale Standard Display Function.

Figure D.3-2. Transformation for modifying the Characteristic Curve of the printer to a Display Function that approximates the Grayscale Standard Display Function
Table D.3-1. Look-Up Table for Calibrating Reflection Hardcopy System
P-Value	DDL		P-Value	DDL		P-Value	DDL		P-Value	DDL
0	6		1	9		2	12		3	15
4	18		5	20		6	27		7	29
8	30		9	31		10	31		11	32
12	33		13	33		14	34		15	36
16	38		17	40		18	41		19	42
20	43		21	44		22	45		23	59
24	60		25	61		26	62		27	62
28	63		29	63		30	64		31	64
32	65		33	65		34	65		35	66
36	66		37	67		38	67		39	68
40	70		41	74		42	75		43	76
44	78		45	84		46	85		47	86
48	87		49	87		50	88		51	89
52	89		53	91		54	92		55	94
56	95		57	96		58	97		59	97
60	98		61	99		62	99		63	100
64	101		65	102		66	103		67	104
68	105		69	106		70	107		71	108
72	109		73	110		74	112		75	114
76	116		77	118		78	119		79	120
80	121		81	122		82	122		83	123
84	123		85	124		86	125		87	125
88	126		89	126		90	127		91	127
92	128		93	129		94	130		95	131
96	133		97	134		98	135		99	136
100	136		101	137		102	138		103	138
104	139		105	139		106	140		107	141
108	143		109	145		110	147		111	148
112	149		113	150		114	151		115	152
116	153		117	154		118	154		119	155
120	156		121	156		122	157		123	158
124	159		125	160		126	160		127	162
128	163		129	164		130	165		131	166
132	167		133	168		134	169		135	170
136	170		137	171		138	172		139	172
140	173		141	174		142	175		143	175
144	176		145	177		146	178		147	179
148	179		149	180		150	181		151	182
152	182		153	183		154	184		155	184
156	185		157	186		158	186		159	187
160	187		161	188		162	188		163	189
164	189		165	190		166	190		167	190
168	191		169	191		170	192		171	192
172	192		173	193		174	194		175	194
176	195		177	195		178	196		179	197
180	198		181	199		182	199		183	200
184	200		185	201		186	202		187	202
188	203		189	203		190	204		191	204
192	205		193	205		194	206		195	207
196	207		197	208		198	209		199	210
200	211		201	212		202	213		203	214
204	214		205	215		206	216		207	216
208	217		209	218		210	219		211	219
212	220		213	220		214	221		215	222
216	222		217	223		218	223		219	224
220	224		221	225		222	226		223	226
224	227		225	228		226	228		227	230
228	231		229	232		230	234		231	235
232	236		233	238		234	238		235	239
236	240		237	241		238	242		239	242
240	243		241	244		242	245		243	246
244	247		245	248		246	249		247	250
248	250		249	251		250	251		251	252
252	252		253	253		254	253		255	254

D.3.4 Measures of Conformance
The FIT and LUM metrics as proposed in Annex C are applied to determine the macroscopic and microscopic approximation of the LJ,m to the LI,STD. Figure D.3-3 shows the perceptually linearized Display Function superimposed on the Grayscale Standard Display Function and Figure D.3-4 summarizes the results of the two metrics. FIT provides as best fit of the JNDs/Luminance interval a straight line almost perfectly parallel to the horizontal axis indicating good global fit of the transformed Display Function with the Grayscale Standard Display Function. The RMSE computed by LUM is relatively large indicating more pronounced local deviations from the Grayscale Standard Display Function as, for example, with the soft-copy Display System illustrated in Section D.1. At least in part, the larger RMSE is due to the fact that the input and output digitization resolution for the transform are equal. The transformation table (Table D.3-1) and Figure D.3-2 show that several P-Values lead to the same Luminance levels on the transformed Display Function. In fact, only 205 of the 255 Luminance intervals lead to JNDs for the Standard Target.

Figure D.3-3. Transformed Display Function and superimposed Grayscale Standard Display Function for a reflective hard-copy Display System. The transformed Display Function for this Display System matches the Grayscale Standard Display Function and the two curves are superimposed and indistinguishable.

Figure D.3-4. Measures of conformance for a reflective hard-copy Display System with equal input and output digitization resolution of 8 bits
E Realizable JND Range of a Display Under Ambient Light (Informative)
Dynamic rangeis an often used measures of the information content that can be presented by a Display System. However, there are many definitions of dynamic range, and most such definitions do not take into account real world conditions that affect the actual amount of information that can be conveyed by a gray scale pixel. For example, Poynton [E1] refers to the contrast ratioof a gray scale display device as the ratio of display intensity between the brightest white and the darkest black of the particular display device in question. However, this definition of dynamic range applies to ideal viewing conditions. Real world conditions such as veiling glare, noise, spatial frequency content of the image, power supply saturation, and ambient lighting in a cathode ray tube (CRT) based viewing situation can degrade the measured dynamic range of the system significantly [E2, E3]. Because of all of these variables dynamic range is an ill-defined concept for a Display System.
Note
Veiling Glare is the phenomenon wherein internal light reflections inside the CRT creates a "background lighting" thus reducing the contrast range of the CRT device.
The methods used to determine the degree to which the Display Function of a Display System approximates the Grayscale Standard Display Function can also be used to define two measures that might better characterize the potential capabilities of a Display System to convey information content. Two measures, the theoretically achievable JNDs and the realized JNDs, are useful for comparing Display Systems [E4].
The number of theoretically achievable JNDs is simply the number of JNDs predicted by the visual model given the Luminance Range of the Display System used. The number of theoretically achievable JNDs of a Display System may be found from Table B-1 by counting the number of JNDs in the table that fall between the measured minimum and maximum Luminance of the Display System.
This number of JNDs may not actually be achievable due to resolution limitations of other portions of the Display System, in particular, the quantization resolution given by the finite number of bits per pixel driving the Display System. For example, Table B-1 may show that a particular Display System is capable of delivering 352 JNDs. However, if only 8 bits per pixel are presented to the Display System, the number of JNDs achievable cannot exceed 2 8= 256 JNDs because of the quantizing effect. In actual fact, the number of JNDs realized in a Display System will always be smaller than or equal to the lower of the theoretically achievable JNDs and the quantization limit. This is because some of the quantized values input to the display may not line up with the input value required to achieve the next JND.
The more useful number of realized JNDs, describes how many JNDs are actually achieved given the specifics of the Display System (i.e., the number of gray levels of contrast resolution and the distribution of Luminance values). This definition gives a measure of the information that can actually be conveyed by the system to a human observer, in essence, an informational dynamic range. This number is calculated beginning at the minimum Luminance of the Display System, and then stepping one JND in Luminance from the current Luminance value, and choosing the smallest increment in DDL value that achieves a step at least that large. Repeating this through all the available DDLs will produce a sequence of steps, all at least 1 JND apart, and the length of this sequence of steps is then the number of realizable JNDs of the Display System.
The methods of PS3.14 cannot precisely duplicate all of the real world sources of degradation in a Display System. However, this uniform method of determining the realizable number of JNDs should give a measure of the actual performance of a particular Display System that would be experienced by a human observer when using the Display System in a real world situation such as the viewing of radiological images in medicine.
References
[E1] Poynton, C. "Frequently Asked Questions about Gamma",Internet ftp://ftp.inforamp.net/pub/users/poynton/doc/colour/gammaFAQ.pdf
[E2] Roehrig, H., Blume, H., Ji, T. and Browne, M.; "Performance Tests and Quality Control of Cathode Ray Tube Displays"; J. Digital Imaging, Vol. 3, No. 3, August 1990; pp. 134-145.
[E3] Gray, J.; "Use of the SMPTE Test Pattern in Picture Archiving and Communication Systems"; J. Digital Imaging, Vol. 5, No. 1, February 1992; pp. 54-58.
[E4] Hemminger, B., Muller, K., "Performance Metric for evaluating conformance of medical image displays with the ACR/NEMA display function standard", SPIE Medical Imaging 1997, editor Yongmin Kim, vol 3031-25, 1997.