PS3.6
DICOM PS3.6 2019a - Data Dictionary
PS3.6: DICOM PS3.6 2019a - Data Dictionary
Copyright © 2019 NEMA
A DICOM® publication
Table of Contents
Notice and Disclaimer	0
Foreword	0
1. Scope and Field of Application	0
2. Normative References	0
3. Definitions	0
4. Symbols and Abbreviations	0
5. Conventions	0
6. Registry of DICOM Data Elements	0
7. Registry of DICOM File Meta Elements	0
8. Registry of DICOM Directory Structuring Elements	0
A. Registry of DICOM Unique Identifiers (UIDs) (Normative)	0
B. Well-Known Color Palettes (Normative)	0
B.1. Standard Color Palettes	0
B.1.1. Hot Iron Color Palette	0
B.1.1.1. Hot Iron Color Palette Description (Informative)	0
B.1.1.2. Hot Iron Color Palette Definition	0
B.1.2. PET Color Palette	0
B.1.2.1. PET Color Palette Description (Informative)	0
B.1.2.2. PET Color Palette Definition	0
B.1.3. Hot Metal Blue Color Palette	0
B.1.3.1. Hot Metal Blue Color Palette Description (Informative)	0
B.1.3.2. Hot Metal Blue Color Palette Definition	0
B.1.4. PET 20 Step Color Palette	0
B.1.4.1. PET 20 Step Color Palette Description (Informative)	0
B.1.4.2. PET 20 Step Color Palette Definition	0
B.1.5. Spring Color Palette	0
B.1.5.1. Spring Color Palette Description (Informative)	0
B.1.5.2. Spring Color Palette Definition	0
B.1.6. Summer Color Palette	0
B.1.6.1. Summer Color Palette Description (Informative)	0
B.1.6.2. Summer Color Palette Definition	0
B.1.7. Fall Color Palette	0
B.1.7.1. Fall Color Palette Description (Informative)	0
B.1.7.2. Fall Color Palette Definition	0
B.1.8. Winter Color Palette	0
B.1.8.1. Winter Color Palette Description (Informative)	0
B.1.8.2. Winter Color Palette Definition	0
B.2. Localized Standard Color Palette Description Values	0
B.2.1. French	0
B.2.2. German	0
List of Figures
B.1.1.1-1. Nuclear Medicine image with and without Hot Iron Palette applied.	0
B.1.2.1-1. PET image with PET Palette superimposed over grayscale CT image.	0
B.1.3.1-1. PET image with Hot Metal Blue Palette applied.	0
B.1.4.1-1. PET image with PET 20 Step Palette applied.	0
B.1.5.1-1. MR image with Spring LUT Palette applied.	0
B.1.6.1-1. MR image with Summer LUT Palette applied.	0
B.1.7.1-1. MR image with Fall LUT Palette applied.	0
B.1.8.1-1. MR image with Winter LUT Palette applied.	0
List of Tables
6-1. Registry of DICOM Data Elements	0
7-1. Registry of DICOM File Meta Elements	0
8-1. Registry of DICOM Directory Structuring Elements	0
A-1. UID Values	0
A-2. Well-known Frames of Reference	0
A-3. Context Group UID Values	0
A-4. Template UID Values	0
B.1-1. Standard Color Palettes	0
B.1.1.2-1. Hot Iron Color Palette Descriptor	0
B.1.1.2-2. Hot Iron Color Palette Data	0
B.1.2.2-1. PET Color Palette Descriptor	0
B.1.2.2-2. PET Color Palette Data	0
B.1.3.2-1. Hot Metal Blue Color Palette Descriptor	0
B.1.3.2-2. Hot Metal Blue Color Palette Data	0
B.1.4.2-1. PET 20 Step Color Palette Descriptor	0
B.1.4.2-2. PET 20 Step Color Palette Data	0
B.1.5.2-1. Spring Color Palette Descriptor	0
B.1.5.2-2. Spring Segmented Color Palette Data	0
B.1.6.2-1. Summer Color Palette Descriptor	0
B.1.6.2-2. Summer Segmented Color Palette Data	0
B.1.7.2-1. Fall Color Palette Descriptor	0
B.1.7.2-2. Fall Segmented Color Palette Data	0
B.1.8.2-1. Winter Color Palette Descriptor	0
B.1.8.2-2. Winter Segmented Color Palette Data	0
B.2.1-1. French Standard Color Palette Description Values	0
B.2.2-1. German Standard Color Palette Description Values	0
Notice and Disclaimer
The information in this publication was considered technically sound by the consensus of persons engaged in the development and approval of the document at the time it was developed. Consensus does not necessarily mean that there is unanimous agreement among every person participating in the development of this document.
NEMA standards and guideline publications, of which the document contained herein is one, are developed through a voluntary consensus standards development process. This process brings together volunteers and/or seeks out the views of persons who have an interest in the topic covered by this publication. While NEMA administers the process and establishes rules to promote fairness in the development of consensus, it does not write the document and it does not independently test, evaluate, or verify the accuracy or completeness of any information or the soundness of any judgments contained in its standards and guideline publications.
NEMA disclaims liability for any personal injury, property, or other damages of any nature whatsoever, whether special, indirect, consequential, or compensatory, directly or indirectly resulting from the publication, use of, application, or reliance on this document. NEMA disclaims and makes no guaranty or warranty, expressed or implied, as to the accuracy or completeness of any information published herein, and disclaims and makes no warranty that the information in this document will fulfill any of your particular purposes or needs. NEMA does not undertake to guarantee the performance of any individual manufacturer or seller's products or services by virtue of this standard or guide.
In publishing and making this document available, NEMA is not undertaking to render professional or other services for or on behalf of any person or entity, nor is NEMA undertaking to perform any duty owed by any person or entity to someone else. Anyone using this document should rely on his or her own independent judgment or, as appropriate, seek the advice of a competent professional in determining the exercise of reasonable care in any given circumstances. Information and other standards on the topic covered by this publication may be available from other sources, which the user may wish to consult for additional views or information not covered by this publication.
NEMA has no power, nor does it undertake to police or enforce compliance with the contents of this document. NEMA does not certify, test, or inspect products, designs, or installations for safety or health purposes. Any certification or other statement of compliance with any health or safety-related information in this document shall not be attributable to NEMA and is solely the responsibility of the certifier or maker of the statement.
Foreword
This DICOM Standard was developed according to the procedures of the DICOM Standards Committee.
The DICOM Standard is structured as a multi-part document using the guidelines established in [ISO/IEC Directives, Part 2].
PS3.1 should be used as the base reference for the current parts of this Standard.
1 Scope and Field of Application
This Part of the DICOM Standard is PS 3.6 of a multi-part standard produced to facilitate the interchange of information between digital imaging computer systems in medical environments. This interchange will enhance diagnostic imaging and potentially other clinical applications. The multi-part DICOM Standard covers the protocols and data that shall be supplied to achieve this interchange of information.
This Part of the Standard contains the registry of all DICOM Data Elements and all DICOM Unique Identifiers that are defined within the DICOM Standard.
DICOM® is the registered trademark of the National Electrical Manufacturers Association for its standards publications relating to digital communications of medical information, all rights reserved.
HL7® and CDA® are the registered trademarks of Health Level Seven International, all rights reserved.
SNOMED®, SNOMED Clinical Terms®, SNOMED CT® are the registered trademarks of the International Health Terminology Standards Development Organisation (IHTSDO), all rights reserved.
LOINC® is the registered trademark of Regenstrief Institute, Inc, all rights reserved.
2 Normative References
The following standards contain provisions that, through references in this text, constitute provisions of this Standard. At the time of publication, the editions indicated were valid. All standards are subject to revision, and parties to agreements based on this Standard are encouraged to investigate the possibilities of applying the most recent editions of the standards indicated below.
[ACR-NEMA 300] ACR-NEMA. 1988. Digital Imaging and Communications.
[ASTM E2339-11] ASTM. 2011. Standard Practice for Digital Imaging and Communication in Nondestructive Evaluation (DICONDE).
[ISO 8649] ISO. 1988. Information processing systems - Open Systems Interconnection - Service definition for the Association Control Service Element (ACSE).
[ISO 8822] ISO. 1988. Information processing systems - Open Systems Interconnection - Connection oriented presentation service definition.
[ISO/IEC Directives, Part 2] ISO/IEC. 2016/05. 7.0. Rules for the structure and drafting of International Standards. http://​www.iec.ch/​members_experts/​refdocs/​iec/​isoiecdir-2%7Bed7.0%7Den.pdf .
[NEMA IIC 1 v02] NEMA. 2012. Digital Imaging and Communications in Security (DICOS) Information Object Definitions (IODs).
3 Definitions
For the purposes of this Standard, the following definitions apply.
3.1 DICOM Introduction and Overview Definition
This Part of the Standard makes use of the following term defined in PS3.1:
Attribute	Attribute.
Service-Object Pair Class (SOP Class)	Service-Object Pair Class (SOP Class).
3.2 DICOM Information Object Definition
This Part of the Standard makes use of the following term defined in PS3.3:
Attribute Tag	AttributeTag.
3.3 DICOM Data Structures and Encoding Definitions
This Part of the Standard makes use of the following terms defined in PS3.5:
Data Element	Data Element
Data Element Tag	Data Element Tag
Element Number	Element Number
Group Number	Group Number
Repeating Group	Repeating Group
Retired Data Element	Retired Data Element
Standard Data Element	Standard Data Element
Value Multiplicity (VM)	Value Multiplicity (VM)
Value Representation (VR)	Value Representation (VR)
3.4 DICOM Data Dictionary
The following definition is commonly used in this Standard:
Tag	A unique identifier for an element of information composed of an ordered pair of numbers (a Group Number followed by an Element Number), which is used to identify Attributes and corresponding Data Elements.
3.5 DICOM Service Class Specifications Definitions
This Part of the Standard makes use of the following terms defined in PS3.4:
Service-Object Pair Instance (SOP Instance)	Service-Object Pair Instance (SOP Instance)
4 Symbols and Abbreviations
The following symbols and abbreviations are used in this Standard.
ACR	American College of Radiology
DICOM	Digital Imaging and Communications in Medicine
DICONDE	Digital Imaging and Communication in Nondestructive Evaluation
DICOS	Digital Imaging and Communication for Security
IOD	Information Object Definition
ISO	International Standards Organization
JIRA	Japan Medical Imaging and Radiological Systems Industries Association
NEMA	National Electrical Manufacturers Association
OSI	Open Systems Interconnection
TCP/IP	Transmission Control Protocol/Internet Protocol
UID	Unique Identifier
VM	Value Multiplicity
VR	Value Representation
5 Conventions
Word(s) are capitalized in this document to help the reader understand that these word(s) have been previously defined in Section 3 and are to be interpreted with that meaning.
A Data Element Tag is represented as (gggg,eeee), where gggg equates to the Group Number and eeee equates to the Element Number within that Group. Data Element Tags are represented in hexadecimal notation as specified for each named Data Element in this Standard.
Where an "x" is shown in a group or element number, e.g. (ggxx,eeee), "x" means any value from 0 through F inclusive. The resulting group number is still required to be even, as defined by PS3.5, since these are Standard Data Element Tags.
"RET" is used to indicate that the corresponding Data Element, SOP Class, or Transfer Syntax has been retired. Retired items are shown italicized. For retired items, the edition of the Standard in parentheses is the edition in which the item last appeared before it was retired. When the name of a retired Data Element has been reused, the retired element has the qualifier "(Retired)" added, or "(Trial)" in the cases in which the Data Element was used in a Draft For Trial Implementation but not standardized.
Note
The use of retired items is supported in this version of DICOM. However, new implementations are strongly encouraged to implement alternative Data Elements, SOP Classes or Transfer Syntaxes.
"Note n" is used to indicate that further information is provided at the end of the corresponding table. The "n" is the consecutive number of the note. This information is not inserted directly into the tables in order to preserve their simple structure, e.g., for automatic processing of the contents.
6 Registry of DICOM Data Elements
Note
For data elements that were present in ACR-NEMA 1.0 and 2.0 and that have been retired, the specifications of Value Representation and Value Multiplicity provided are recommendations for the purpose of interpreting their values in objects created in accordance with earlier editions of this Standard. These recommendations are suggested as most appropriate for a particular data element; however, there is no guarantee that historical objects will not violate some requirements or specified VR and/or VM.
Table 6-1. Registry of DICOM Data Elements
Tag	Name	Keyword	VR	VM
(0008,0001)	Length to End	Length​To​End	UL	1	RET
(0008,0005)	Specific Character Set	Specific​Character​Set	CS	1-n
(0008,0006)	Language Code Sequence	Language​Code​Sequence	SQ	1
(0008,0008)	Image Type	Image​Type	CS	2-n
(0008,0010)	Recognition Code	Recognition​Code	SH	1	RET
(0008,0012)	Instance Creation Date	Instance​Creation​Date	DA	1
(0008,0013)	Instance Creation Time	Instance​Creation​Time	TM	1
(0008,0014)	Instance Creator UID	Instance​Creator​UID	UI	1
(0008,0015)	Instance Coercion DateTime	Instance​Coercion​Date​Time	DT	1
(0008,0016)	SOP Class UID	SOP​Class​UID	UI	1
(0008,0018)	SOP Instance UID	SOP​Instance​UID	UI	1
(0008,001A)	Related General SOP Class UID	Related​General​SOP​Class​UID	UI	1-n
(0008,001B)	Original Specialized SOP Class UID	Original​Specialized​SOP​Class​UID	UI	1
(0008,0020)	Study Date	Study​Date	DA	1
(0008,0021)	Series Date	Series​Date	DA	1
(0008,0022)	Acquisition Date	Acquisition​Date	DA	1
(0008,0023)	Content Date	Content​Date	DA	1
(0008,0024)	Overlay Date	Overlay​Date	DA	1	RET (2004)
(0008,0025)	Curve Date	Curve​Date	DA	1	RET (2004)
(0008,002A)	Acquisition DateTime	Acquisition​Date​Time	DT	1
(0008,0030)	Study Time	Study​Time	TM	1
(0008,0031)	Series Time	Series​Time	TM	1
(0008,0032)	Acquisition Time	Acquisition​Time	TM	1
(0008,0033)	Content Time	Content​Time	TM	1
(0008,0034)	Overlay Time	Overlay​Time	TM	1	RET (2004)
(0008,0035)	Curve Time	Curve​Time	TM	1	RET
(0008,0040)	Data Set Type	Data​Set​Type	US	1	RET (2004)
(0008,0041)	Data Set Subtype	Data​Set​Subtype	LO	1	RET (2004)
(0008,0042)	Nuclear Medicine Series Type	Nuclear​Medicine​Series​Type	CS	1	RET (2004)
(0008,0050)	Accession Number	Accession​Number	SH	1
(0008,0051)	Issuer of Accession Number Sequence	Issuer​Of​Accession​Number​Sequence	SQ	1
(0008,0052)	Query/Retrieve Level	Query​Retrieve​Level	CS	1
(0008,0053)	Query/Retrieve View	Query​Retrieve​View	CS	1
(0008,0054)	Retrieve AE Title	Retrieve​AE​Title	AE	1-n
(0008,0055)	Station AE Title	Station​AE​Title	AE	1
(0008,0056)	Instance Availability	Instance​Availability	CS	1
(0008,0058)	Failed SOP Instance UID List	Failed​SOP​Instance​UID​List	UI	1-n
(0008,0060)	Modality	Modality	CS	1
(0008,0061)	Modalities in Study	Modalities​In​Study	CS	1-n
(0008,0062)	SOP Classes in Study	SOP​Classes​In​Study	UI	1-n
(0008,0063)	Anatomic Regions in Study Code Sequence	Anatomic​Regions​In​StudyC​ode​Sequence	SQ	1
(0008,0064)	Conversion Type	Conversion​Type	CS	1
(0008,0068)	Presentation Intent Type	Presentation​Intent​Type	CS	1
(0008,0070)	Manufacturer	Manufacturer	LO	1
(0008,0080)	Institution Name	Institution​Name	LO	1
(0008,0081)	Institution Address	Institution​Address	ST	1
(0008,0082)	Institution Code Sequence	Institution​Code​Sequence	SQ	1
(0008,0090)	Referring Physician's Name	Referring​Physician​Name	PN	1
(0008,0092)	Referring Physician's Address	Referring​Physician​Address	ST	1
(0008,0094)	Referring Physician's Telephone Numbers	Referring​Physician​Telephone​Numbers	SH	1-n
(0008,0096)	Referring Physician Identification Sequence	Referring​Physician​Identification​Sequence	SQ	1
(0008,009C)	Consulting Physician's Name	Consulting​Physician​Name	PN	1-n
(0008,009D)	Consulting Physician Identification Sequence	Consulting​Physician​Identification​Sequence	SQ	1
(0008,0100)	Code Value	Code​Value	SH	1
(0008,0101)	Extended Code Value	Extended​Code​Value	LO	1	DICOS
(0008,0102)	Coding Scheme Designator	Coding​Scheme​Designator	SH	1
(0008,0103)	Coding Scheme Version	Coding​Scheme​Version	SH	1
(0008,0104)	Code Meaning	Code​Meaning	LO	1
(0008,0105)	Mapping Resource	Mapping​Resource	CS	1
(0008,0106)	Context Group Version	Context​Group​Version	DT	1
(0008,0107)	Context Group Local Version	Context​Group​Local​Version	DT	1
(0008,0108)	Extended Code Meaning	Extended​Code​Meaning	LT	1	DICOS
(0008,0109)	Coding Scheme Resources Sequence	Coding​Scheme​Resources​Sequence	SQ	1
(0008,010A)	Coding Scheme URL Type	Coding​Scheme​URL​Type	CS	1
(0008,010B)	Context Group Extension Flag	Context​Group​Extension​Flag	CS	1
(0008,010C)	Coding Scheme UID	Coding​Scheme​UID	UI	1
(0008,010D)	Context Group Extension Creator UID	Context​Group​Extension​Creator​UID	UI	1
(0008,010E)	Coding Scheme URL	Coding​Scheme​URL	UR	1
(0008,010F)	Context Identifier	Context​Identifier	CS	1
(0008,0110)	Coding Scheme Identification Sequence	Coding​Scheme​Identification​Sequence	SQ	1
(0008,0112)	Coding Scheme Registry	Coding​Scheme​Registry	LO	1
(0008,0114)	Coding Scheme External ID	Coding​Scheme​External​ID	ST	1
(0008,0115)	Coding Scheme Name	Coding​Scheme​Name	ST	1
(0008,0116)	Coding Scheme Responsible Organization	Coding​Scheme​Responsible​Organization	ST	1
(0008,0117)	Context UID	Context​UID	UI	1
(0008,0118)	Mapping Resource UID	Mapping​Resource​UID	UI	1
(0008,0119)	Long Code Value	Long​Code​Value	UC	1
(0008,0120)	URN Code Value	URN​Code​Value	UR	1
(0008,0121)	Equivalent Code Sequence	Equivalent​Code​Sequence	SQ	1
(0008,0122)	Mapping Resource Name	Mapping​Resource​Name	LO	1
(0008,0123)	Context Group Identification Sequence	Context​Group​Identification​Sequence	SQ	1
(0008,0124)	Mapping Resource Identification Sequence	Mapping​Resource​Identification​Sequence	SQ	1
(0008,0201)	Timezone Offset From UTC	Timezone​Offset​From​UTC	SH	1
(0008,0220)	Responsible Group Code Sequence	Responsible​Group​Code​Sequence	SQ	1
(0008,0221)	Equipment Modality	Equipment​Modality	CS	1
(0008,0222)	Manufacturer's Related Model Group	Manufacturer​Related​Model​Group	LO	1
(0008,0300)	Private Data Element Characteristics Sequence	Private​Data​Element​Characteristics​Sequence	SQ	1
(0008,0301)	Private Group Reference	Private​Group​Reference	US	1
(0008,0302)	Private Creator Reference	Private​Creator​Reference	LO	1
(0008,0303)	Block Identifying Information Status	Block​Identifying​Information​Status	CS	1
(0008,0304)	Nonidentifying Private Elements	Nonidentifying​Private​Elements	US	1-n
(0008,0306)	Identifying Private Elements	Identifying​Private​Elements	US	1-n
(0008,0305)	Deidentification Action Sequence	Deidentification​Action​Sequence	SQ	1
(0008,0307)	Deidentification Action	Deidentification​Action	CS	1
(0008,0308)	Private Data Element	Private​Data​Element	US	1
(0008,0309)	Private Data Element Value Multiplicity	Private​Data​Element​Value​Multiplicity	UL	1-3
(0008,030A)	Private Data Element Value Representation	Private​Data​Element​Value​Representation	CS	1
(0008,030B)	Private Data Element Number of Items	Private​Data​Element​Number​Of​Items	UL	1-2
(0008,030C)	Private Data Element Name	Private​Data​Element​Name	UC	1
(0008,030D)	Private Data Element Keyword	Private​Data​Element​Keyword	UC	1
(0008,030E)	Private Data Element Description	Private​Data​Element​Description	UT	1
(0008,030F)	Private Data Element Encoding	Private​Data​Element​Encoding	UT	1
(0008,0310)	Private Data Element Definition Sequence	Private​Data​Element​Definition​Sequence	SQ	1
(0008,1000)	Network ID	Network​ID	AE	1	RET
(0008,1010)	Station Name	Station​Name	SH	1
(0008,1030)	Study Description	Study​Description	LO	1
(0008,1032)	Procedure Code Sequence	Procedure​Code​Sequence	SQ	1
(0008,103E)	Series Description	Series​Description	LO	1
(0008,103F)	Series Description Code Sequence	Series​Description​Code​Sequence	SQ	1
(0008,1040)	Institutional Department Name	Institutional​Department​Name	LO	1
(0008,1041)	Institutional Department Type Code Sequence	Institutional​Department​Type​Code​Sequence	SQ	1
(0008,1048)	Physician(s) of Record	Physicians​Of​Record	PN	1-n
(0008,1049)	Physician(s) of Record Identification Sequence	Physicians​Of​Record​Identification​Sequence	SQ	1
(0008,1050)	Performing Physician's Name	Performing​Physician​Name	PN	1-n
(0008,1052)	Performing Physician Identification Sequence	Performing​Physician​Identification​Sequence	SQ	1
(0008,1060)	Name of Physician(s) Reading Study	Name​Of​Physicians​Reading​Study	PN	1-n
(0008,1062)	Physician(s) Reading Study Identification Sequence	Physicians​Reading​Study​Identification​Sequence	SQ	1
(0008,1070)	Operators' Name	Operators​Name	PN	1-n
(0008,1072)	Operator Identification Sequence	Operator​Identification​Sequence	SQ	1
(0008,1080)	Admitting Diagnoses Description	Admitting​Diagnoses​Description	LO	1-n
(0008,1084)	Admitting Diagnoses Code Sequence	Admitting​Diagnoses​Code​Sequence	SQ	1
(0008,1090)	Manufacturer's Model Name	Manufacturer​Model​Name	LO	1
(0008,1100)	Referenced Results Sequence	Referenced​Results​Sequence	SQ	1	RET (2004)
(0008,1110)	Referenced Study Sequence	Referenced​Study​Sequence	SQ	1
(0008,1111)	Referenced Performed Procedure Step Sequence	Referenced​Performed​Procedure​Step​Sequence	SQ	1
(0008,1115)	Referenced Series Sequence	Referenced​Series​Sequence	SQ	1
(0008,1120)	Referenced Patient Sequence	Referenced​Patient​Sequence	SQ	1
(0008,1125)	Referenced Visit Sequence	Referenced​Visit​Sequence	SQ	1
(0008,1130)	Referenced Overlay Sequence	Referenced​Overlay​Sequence	SQ	1	RET (2004)
(0008,1134)	Referenced Stereometric Instance Sequence	Referenced​Stereometric​Instance​Sequence	SQ	1
(0008,113A)	Referenced Waveform Sequence	Referenced​Waveform​Sequence	SQ	1
(0008,1140)	Referenced Image Sequence	Referenced​Image​Sequence	SQ	1
(0008,1145)	Referenced Curve Sequence	Referenced​Curve​Sequence	SQ	1	RET (2004)
(0008,114A)	Referenced Instance Sequence	Referenced​Instance​Sequence	SQ	1
(0008,114B)	Referenced Real World Value Mapping Instance Sequence	Referenced​Real​World​Value​Mapping​Instance​Sequence	SQ	1
(0008,1150)	Referenced SOP Class UID	Referenced​SOP​Class​UID	UI	1
(0008,1155)	Referenced SOP Instance UID	Referenced​SOP​Instance​UID	UI	1
(0008,1156)	Definition Source Sequence	Definition​Source​Sequence	SQ	1
(0008,115A)	SOP Classes Supported	SOP​Classes​Supported	UI	1-n
(0008,1160)	Referenced Frame Number	Referenced​Frame​Number	IS	1-n
(0008,1161)	Simple Frame List	Simple​Frame​List	UL	1-n
(0008,1162)	Calculated Frame List	Calculated​Frame​List	UL	3-3n
(0008,1163)	Time Range	Time​Range	FD	2
(0008,1164)	Frame Extraction Sequence	Frame​Extraction​Sequence	SQ	1
(0008,1167)	Multi-frame Source SOP Instance UID	Multi​Frame​Source​SOP​Instance​UID	UI	1
(0008,1190)	Retrieve URL	Retrieve​URL	UR	1
(0008,1195)	Transaction UID	Transaction​UID	UI	1
(0008,1196)	Warning Reason	Warning​Reason	US	1
(0008,1197)	Failure Reason	Failure​Reason	US	1
(0008,1198)	Failed SOP Sequence	Failed​SOP​Sequence	SQ	1
(0008,1199)	Referenced SOP Sequence	Referenced​SOP​Sequence	SQ	1
(0008,119A)	Other Failures Sequence	Other​Failures​Sequence	SQ	1
(0008,1200)	Studies Containing Other Referenced Instances Sequence	Studies​Containing​Other​Referenced​Instances​Sequence	SQ	1
(0008,1250)	Related Series Sequence	Related​Series​Sequence	SQ	1
(0008,2110)	Lossy Image Compression (Retired)	Lossy​Image​Compression​Retired	CS	1	RET (1993)
(0008,2111)	Derivation Description	Derivation​Description	ST	1
(0008,2112)	Source Image Sequence	Source​Image​Sequence	SQ	1
(0008,2120)	Stage Name	Stage​Name	SH	1
(0008,2122)	Stage Number	Stage​Number	IS	1
(0008,2124)	Number of Stages	Number​Of​Stages	IS	1
(0008,2127)	View Name	View​Name	SH	1
(0008,2128)	View Number	View​Number	IS	1
(0008,2129)	Number of Event Timers	Number​Of​Event​Timers	IS	1
(0008,212A)	Number of Views in Stage	Number​Of​Views​In​Stage	IS	1
(0008,2130)	Event Elapsed Time(s)	Event​Elapsed​Times	DS	1-n
(0008,2132)	Event Timer Name(s)	Event​Timer​Names	LO	1-n
(0008,2133)	Event Timer Sequence	Event​Timer​Sequence	SQ	1
(0008,2134)	Event Time Offset	Event​Time​Offset	FD	1
(0008,2135)	Event Code Sequence	Event​Code​Sequence	SQ	1
(0008,2142)	Start Trim	Start​Trim	IS	1
(0008,2143)	Stop Trim	Stop​Trim	IS	1
(0008,2144)	Recommended Display Frame Rate	Recommended​Display​Frame​Rate	IS	1
(0008,2200)	Transducer Position	Transducer​Position	CS	1	RET (1993)
(0008,2204)	Transducer Orientation	Transducer​Orientation	CS	1	RET (1993)
(0008,2208)	Anatomic Structure	Anatomic​Structure	CS	1	RET (1993)
(0008,2218)	Anatomic Region Sequence	Anatomic​Region​Sequence	SQ	1
(0008,2220)	Anatomic Region Modifier Sequence	Anatomic​Region​Modifier​Sequence	SQ	1
(0008,2228)	Primary Anatomic Structure Sequence	Primary​Anatomic​Structure​Sequence	SQ	1
(0008,2229)	Anatomic Structure, Space or Region Sequence	Anatomic​Structure​Space​Or​Region​Sequence	SQ	1	RET (2017c)
(0008,2230)	Primary Anatomic Structure Modifier Sequence	Primary​Anatomic​Structure​Modifier​Sequence	SQ	1
(0008,2240)	Transducer Position Sequence	Transducer​Position​Sequence	SQ	1	RET (2004)
(0008,2242)	Transducer Position Modifier Sequence	Transducer​Position​Modifier​Sequence	SQ	1	RET (2004)
(0008,2244)	Transducer Orientation Sequence	Transducer​Orientation​Sequence	SQ	1	RET (2004)
(0008,2246)	Transducer Orientation Modifier Sequence	Transducer​Orientation​Modifier​Sequence	SQ	1	RET (2004)
(0008,2251)	Anatomic Structure Space Or Region Code Sequence (Trial)	Anatomic​Structure​Space​Or​Region​Code​Sequence​Trial	SQ	1	RET (2007)
(0008,2253)	Anatomic Portal Of Entrance Code Sequence (Trial)	Anatomic​Portal​Of​Entrance​Code​Sequence​Trial	SQ	1	RET (2007)
(0008,2255)	Anatomic Approach Direction Code Sequence (Trial)	Anatomic​Approach​Direction​Code​Sequence​Trial	SQ	1	RET (2007)
(0008,2256)	Anatomic Perspective Description (Trial)	Anatomic​Perspective​Description​Trial	ST	1	RET (2007)
(0008,2257)	Anatomic Perspective Code Sequence (Trial)	Anatomic​Perspective​Code​Sequence​Trial	SQ	1	RET (2007)
(0008,2258)	Anatomic Location Of Examining Instrument Description (Trial)	Anatomic​Location​Of​Examining​Instrument​Description​Trial	ST	1	RET (2007)
(0008,2259)	Anatomic Location Of Examining Instrument Code Sequence (Trial)	Anatomic​Location​Of​Examining​Instrument​Code​Sequence​Trial	SQ	1	RET (2007)
(0008,225A)	Anatomic Structure Space Or Region Modifier Code Sequence (Trial)	Anatomic​Structure​Space​Or​Region​Modifier​Code​Sequence​Trial	SQ	1	RET (2007)
(0008,225C)	On Axis Background Anatomic Structure Code Sequence (Trial)	On​Axis​Background​Anatomic​Structure​Code​Sequence​Trial	SQ	1	RET (2007)
(0008,3001)	Alternate Representation Sequence	Alternate​Representation​Sequence	SQ	1
(0008,3010)	Irradiation Event UID	Irradiation​Event​UID	UI	1-n
(0008,3011)	Source Irradiation Event Sequence	Source​Irradiation​Event​Sequence	SQ	1
(0008,3012)	Radiopharmaceutical Administration Event UID	Radiopharmaceutical​Administration​Event​UID	UI	1
(0008,4000)	Identifying Comments	Identifying​Comments	LT	1	RET
(0008,9007)	Frame Type	Frame​Type	CS	4
(0008,9092)	Referenced Image Evidence Sequence	Referenced​Image​Evidence​Sequence	SQ	1
(0008,9121)	Referenced Raw Data Sequence	Referenced​Raw​Data​Sequence	SQ	1
(0008,9123)	Creator-Version UID	Creator​Version​UID	UI	1
(0008,9124)	Derivation Image Sequence	Derivation​Image​Sequence	SQ	1
(0008,9154)	Source Image Evidence Sequence	Source​Image​Evidence​Sequence	SQ	1
(0008,9205)	Pixel Presentation	Pixel​Presentation	CS	1
(0008,9206)	Volumetric Properties	Volumetric​Properties	CS	1
(0008,9207)	Volume Based Calculation Technique	Volume​Based​Calculation​Technique	CS	1
(0008,9208)	Complex Image Component	Complex​Image​Component	CS	1
(0008,9209)	Acquisition Contrast	Acquisition​Contrast	CS	1
(0008,9215)	Derivation Code Sequence	Derivation​Code​Sequence	SQ	1
(0008,9237)	Referenced Presentation State Sequence	Referenced​Presentation​State​Sequence	SQ	1
(0008,9410)	Referenced Other Plane Sequence	Referenced​Other​Plane​Sequence	SQ	1
(0008,9458)	Frame Display Sequence	Frame​Display​Sequence	SQ	1
(0008,9459)	Recommended Display Frame Rate in Float	Recommended​Display​Frame​Rate​In​Float	FL	1
(0008,9460)	Skip Frame Range Flag	Skip​Frame​Range​Flag	CS	1
(0010,0010)	Patient's Name	Patient​Name	PN	1
(0010,0020)	Patient ID	Patient​ID	LO	1
(0010,0021)	Issuer of Patient ID	Issuer​Of​Patient​ID	LO	1
(0010,0022)	Type of Patient ID	Type​Of​Patient​ID	CS	1
(0010,0024)	Issuer of Patient ID Qualifiers Sequence	Issuer​Of​Patient​ID​Qualifiers​Sequence	SQ	1
(0010,0026)	Source Patient Group Identification Sequence	Source​Patient​Group​Identification​Sequence	SQ	1
(0010,0027)	Group of Patients Identification Sequence	Group​Of​Patients​Identification​Sequence	SQ	1
(0010,0028)	Subject Relative Position in Image	Subject​Relative​Position​In​Image	US	3
(0010,0030)	Patient's Birth Date	Patient​Birth​Date	DA	1
(0010,0032)	Patient's Birth Time	Patient​Birth​Time	TM	1
(0010,0033)	Patient's Birth Date in Alternative Calendar	Patient​Birth​Date​In​Alternative​Calendar	LO	1
(0010,0034)	Patient's Death Date in Alternative Calendar	Patient​Death​Date​In​Alternative​Calendar	LO	1
(0010,0035)	Patient's Alternative Calendar	Patient​Alternative​Calendar	CS	1
(0010,0040)	Patient's Sex	Patient​Sex	CS	1
(0010,0050)	Patient's Insurance Plan Code Sequence	Patient​Insurance​Plan​Code​Sequence	SQ	1
(0010,0101)	Patient's Primary Language Code Sequence	Patient​Primary​Language​Code​Sequence	SQ	1
(0010,0102)	Patient's Primary Language Modifier Code Sequence	Patient​Primary​Language​Modifier​Code​Sequence	SQ	1
(0010,0200)	Quality Control Subject	Quality​Control​Subject	CS	1
(0010,0201)	Quality Control Subject Type Code Sequence	Quality​Control​Subject​Type​Code​Sequence	SQ	1
(0010,0212)	Strain Description	Strain​Description	UC	1
(0010,0213)	Strain Nomenclature	Strain​Nomenclature	LO	1
(0010,0214)	Strain Stock Number	Strain​Stock​Number	LO	1
(0010,0215)	Strain Source Registry Code Sequence	Strain​Source​Registry​Code​Sequence	SQ	1
(0010,0216)	Strain Stock Sequence	Strain​Stock​Sequence	SQ	1
(0010,0217)	Strain Source	Strain​Source	LO	1
(0010,0218)	Strain Additional Information	Strain​Additional​Information	UT	1
(0010,0219)	Strain Code Sequence	Strain​Code​Sequence	SQ	1
(0010,0221)	Genetic Modifications Sequence	Genetic​Modifications​​Sequence	SQ	1
(0010,0222)	Genetic Modifications Description	Genetic​Modifications​Description	UC	1
(0010,0223)	Genetic Modifications Nomenclature	Genetic​Modifications​Nomenclature	LO	1
(0010,0229)	Genetic Modifications Code Sequence	Genetic​Modifications​Code​Sequence	SQ	1
(0010,1000)	Other Patient IDs	Other​Patient​IDs	LO	1-n	RET (2017a)
(0010,1001)	Other Patient Names	Other​Patient​Names	PN	1-n
(0010,1002)	Other Patient IDs Sequence	Other​Patient​IDs​Sequence	SQ	1
(0010,1005)	Patient's Birth Name	Patient​Birth​Name	PN	1
(0010,1010)	Patient's Age	Patient​Age	AS	1
(0010,1020)	Patient's Size	Patient​Size	DS	1
(0010,1021)	Patient's Size Code Sequence	Patient​Size​Code​Sequence	SQ	1
(0010,1022)	Patient's Body Mass Index	Patient​Body​Mass​Index	DS	1
(0010,1023)	Measured AP Dimension	Measured​APDimension	DS	1
(0010,1024)	Measured Lateral Dimension	Measured​Lateral​Dimension	DS	1
(0010,1030)	Patient's Weight	Patient​Weight	DS	1
(0010,1040)	Patient's Address	Patient​Address	LO	1
(0010,1050)	Insurance Plan Identification	Insurance​Plan​Identification	LO	1-n	RET
(0010,1060)	Patient's Mother's Birth Name	Patient​Mother​Birth​Name	PN	1
(0010,1080)	Military Rank	Military​Rank	LO	1
(0010,1081)	Branch of Service	Branch​Of​Service	LO	1
(0010,1090)	Medical Record Locator	Medical​Record​Locator	LO	1	RET (2017a)
(0010,1100)	Referenced Patient Photo Sequence	Referenced​Patient​Photo​Sequence	SQ	1
(0010,2000)	Medical Alerts	Medical​Alerts	LO	1-n
(0010,2110)	Allergies	Allergies	LO	1-n
(0010,2150)	Country of Residence	Country​Of​Residence	LO	1
(0010,2152)	Region of Residence	Region​Of​Residence	LO	1
(0010,2154)	Patient's Telephone Numbers	Patient​Telephone​Numbers	SH	1-n
(0010,2155)	Patient's Telecom Information	Patient​Telecom​Information​	LT	1
(0010,2160)	Ethnic Group	Ethnic​Group	SH	1
(0010,2180)	Occupation	Occupation	SH	1
(0010,21A0)	Smoking Status	Smoking​Status	CS	1
(0010,21B0)	Additional Patient History	Additional​Patient​History	LT	1
(0010,21C0)	Pregnancy Status	Pregnancy​Status	US	1
(0010,21D0)	Last Menstrual Date	Last​Menstrual​Date	DA	1
(0010,21F0)	Patient's Religious Preference	Patient​Religious​Preference	LO	1
(0010,2201)	Patient Species Description	Patient​Species​Description	LO	1
(0010,2202)	Patient Species Code Sequence	Patient​Species​Code​Sequence	SQ	1
(0010,2203)	Patient's Sex Neutered	Patient​Sex​Neutered	CS	1
(0010,2210)	Anatomical Orientation Type	Anatomical​Orientation​Type	CS	1
(0010,2292)	Patient Breed Description	Patient​Breed​Description	LO	1
(0010,2293)	Patient Breed Code Sequence	Patient​Breed​Code​Sequence	SQ	1
(0010,2294)	Breed Registration Sequence	Breed​Registration​Sequence	SQ	1
(0010,2295)	Breed Registration Number	Breed​Registration​Number	LO	1
(0010,2296)	Breed Registry Code Sequence	Breed​Registry​Code​Sequence	SQ	1
(0010,2297)	Responsible Person	Responsible​Person	PN	1
(0010,2298)	Responsible Person Role	Responsible​Person​Role	CS	1
(0010,2299)	Responsible Organization	Responsible​Organization	LO	1
(0010,4000)	Patient Comments	Patient​Comments	LT	1
(0010,9431)	Examined Body Thickness	Examined​Body​Thickness	FL	1
(0012,0010)	Clinical Trial Sponsor Name	Clinical​Trial​Sponsor​Name	LO	1
(0012,0020)	Clinical Trial Protocol ID	Clinical​Trial​Protocol​ID	LO	1
(0012,0021)	Clinical Trial Protocol Name	Clinical​Trial​Protocol​Name	LO	1
(0012,0030)	Clinical Trial Site ID	Clinical​Trial​Site​ID	LO	1
(0012,0031)	Clinical Trial Site Name	Clinical​Trial​Site​Name	LO	1
(0012,0040)	Clinical Trial Subject ID	Clinical​Trial​Subject​ID	LO	1
(0012,0042)	Clinical Trial Subject Reading ID	Clinical​Trial​Subject​Reading​ID	LO	1
(0012,0050)	Clinical Trial Time Point ID	Clinical​Trial​Time​Point​ID	LO	1
(0012,0051)	Clinical Trial Time Point Description	Clinical​Trial​Time​Point​Description	ST	1
(0012,0052)	Longitudinal Temporal Offset from Event	Longitudinal​Temporal​Offset​From​Event	FD	1
(0012,0053)	Longitudinal Temporal Event Type	Longitudinal​Temporal​Event​Type	CS	1
(0012,0060)	Clinical Trial Coordinating Center Name	Clinical​Trial​Coordinating​Center​Name	LO	1
(0012,0062)	Patient Identity Removed	Patient​Identity​Removed	CS	1
(0012,0063)	De-identification Method	Deidentification​Method	LO	1-n
(0012,0064)	De-identification Method Code Sequence	Deidentification​Method​Code​Sequence	SQ	1
(0012,0071)	Clinical Trial Series ID	Clinical​Trial​Series​ID	LO	1
(0012,0072)	Clinical Trial Series Description	Clinical​Trial​Series​Description	LO	1
(0012,0081)	Clinical Trial Protocol Ethics Committee Name	Clinical​Trial​Protocol​Ethics​Committee​Name	LO	1
(0012,0082)	Clinical Trial Protocol Ethics Committee Approval Number	Clinical​Trial​Protocol​Ethics​Committee​Approval​Number	LO	1
(0012,0083)	Consent for Clinical Trial Use Sequence	Consent​For​Clinical​Trial​Use​Sequence	SQ	1
(0012,0084)	Distribution Type	Distribution​Type	CS	1
(0012,0085)	Consent for Distribution Flag	Consent​For​Distribution​Flag	CS	1
(0012,0086)	Ethics Committee Approval Effectiveness Start Date	Ethics​Committee​Approval​Effectiveness​Start​Date	DA	1
(0012,0087)	Ethics Committee Approval Effectiveness End Date	Ethics​Committee​Approval​Effectiveness​End​Date	DA	1
(0014,0023)	CAD File Format	CAD​File​Format	ST	1	RET (2011)
(0014,0024)	Component Reference System	Component​Reference​System	ST	1	RET (2011)
(0014,0025)	Component Manufacturing Procedure	Component​Manufacturing​Procedure	ST	1	DICONDE
(0014,0028)	Component Manufacturer	Component​Manufacturer	ST	1	DICONDE
(0014,0030)	Material Thickness	Material​Thickness	DS	1-n	DICONDE
(0014,0032)	Material Pipe Diameter	Material​Pipe​Diameter	DS	1-n	DICONDE
(0014,0034)	Material Isolation Diameter	Material​Isolation​Diameter	DS	1-n	DICONDE
(0014,0042)	Material Grade	Material​Grade	ST	1	DICONDE
(0014,0044)	Material Properties Description	Material​Properties​Description	ST	1	DICONDE
(0014,0045)	Material Properties File Format (Retired)	Material​Properties​File​Format​Retired	ST	1	RET (2011)
(0014,0046)	Material Notes	Material​Notes	LT	1	DICONDE
(0014,0050)	Component Shape	Component​Shape	CS	1	DICONDE
(0014,0052)	Curvature Type	Curvature​Type	CS	1	DICONDE
(0014,0054)	Outer Diameter	Outer​Diameter	DS	1	DICONDE
(0014,0056)	Inner Diameter	Inner​Diameter	DS	1	DICONDE
(0014,0100)	Component Welder IDs	Component​Welder​IDs	LO	1-n	DICONDE
(0014,0101)	Secondary Approval Status	Secondary​Approval​Status	CS	1	DICONDE
(0014,0102)	Secondary Review Date	Secondary​Review​Date	DA	1	DICONDE
(0014,0103)	Secondary Review Time	Secondary​Review​Time	TM	1	DICONDE
(0014,0104)	Secondary Reviewer Name	Secondary​Reviewer​Name	PN	1	DICONDE
(0014,0105)	Repair ID	Repair​ID	ST	1	DICONDE
(0014,0106)	Multiple Component Approval Sequence	Multiple​Component​Approval​Sequence	SQ	1	DICONDE
(0014,0107)	Other Approval Status	Other​Approval​Status	CS	1-n	DICONDE
(0014,0108)	Other Secondary Approval Status	Other​Secondary​Approval​Status	CS	1-n	DICONDE
(0014,1010)	Actual Environmental Conditions	Actual​Environmental​Conditions	ST	1	DICONDE
(0014,1020)	Expiry Date	Expiry​Date	DA	1	DICONDE
(0014,1040)	Environmental Conditions	Environmental​Conditions	ST	1	DICONDE
(0014,2002)	Evaluator Sequence	Evaluator​Sequence	SQ	1	DICONDE
(0014,2004)	Evaluator Number	Evaluator​Number	IS	1	DICONDE
(0014,2006)	Evaluator Name	Evaluator​Name	PN	1	DICONDE
(0014,2008)	Evaluation Attempt	Evaluation​Attempt	IS	1	DICONDE
(0014,2012)	Indication Sequence	Indication​Sequence	SQ	1	DICONDE
(0014,2014)	Indication Number	Indication​Number	IS	1	DICONDE
(0014,2016)	Indication Label	Indication​Label	SH	1	DICONDE
(0014,2018)	Indication Description	Indication​Description	ST	1	DICONDE
(0014,201A)	Indication Type	Indication​Type	CS	1-n	DICONDE
(0014,201C)	Indication Disposition	Indication​Disposition	CS	1	DICONDE
(0014,201E)	Indication ROI Sequence	Indication​ROI​Sequence	SQ	1	DICONDE
(0014,2030)	Indication Physical Property Sequence	Indication​Physical​Property​Sequence	SQ	1	DICONDE
(0014,2032)	Property Label	Property​Label	SH	1	DICONDE
(0014,2202)	Coordinate System Number of Axes	Coordinate​System​Number​Of​Axes	IS	1	DICONDE
(0014,2204)	Coordinate System Axes Sequence	Coordinate​System​Axes​Sequence	SQ	1	DICONDE
(0014,2206)	Coordinate System Axis Description	Coordinate​System​Axis​Description	ST	1	DICONDE
(0014,2208)	Coordinate System Data Set Mapping	Coordinate​System​Data​Set​Mapping	CS	1	DICONDE
(0014,220A)	Coordinate System Axis Number	Coordinate​System​Axis​Number	IS	1	DICONDE
(0014,220C)	Coordinate System Axis Type	Coordinate​System​Axis​Type	CS	1	DICONDE
(0014,220E)	Coordinate System Axis Units	Coordinate​System​Axis​Units	CS	1	DICONDE
(0014,2210)	Coordinate System Axis Values	Coordinate​System​Axis​Values	OB	1	DICONDE
(0014,2220)	Coordinate System Transform Sequence	Coordinate​System​Transform​Sequence	SQ	1	DICONDE
(0014,2222)	Transform Description	Transform​Description	ST	1	DICONDE
(0014,2224)	Transform Number of Axes	Transform​Number​Of​Axes	IS	1	DICONDE
(0014,2226)	Transform Order of Axes	Transform​Order​Of​Axes	IS	1-n	DICONDE
(0014,2228)	Transformed Axis Units	Transformed​Axis​Units	CS	1	DICONDE
(0014,222A)	Coordinate System Transform Rotation and Scale Matrix	Coordinate​System​Transform​Rotation​And​Scale​Matrix	DS	1-n	DICONDE
(0014,222C)	Coordinate System Transform Translation Matrix	Coordinate​System​Transform​Translation​Matrix	DS	1-n	DICONDE
(0014,3011)	Internal Detector Frame Time	Internal​Detector​Frame​Time	DS	1	DICONDE
(0014,3012)	Number of Frames Integrated	Number​Of​Frames​Integrated	DS	1	DICONDE
(0014,3020)	Detector Temperature Sequence	Detector​Temperature​Sequence	SQ	1	DICONDE
(0014,3022)	Sensor Name	Sensor​Name	ST	1	DICONDE
(0014,3024)	Horizontal Offset of Sensor	Horizontal​Offset​Of​Sensor	DS	1	DICONDE
(0014,3026)	Vertical Offset of Sensor	Vertical​Offset​Of​Sensor	DS	1	DICONDE
(0014,3028)	Sensor Temperature	Sensor​Temperature	DS	1	DICONDE
(0014,3040)	Dark Current Sequence	Dark​Current​Sequence	SQ	1	DICONDE
(0014,3050)	Dark Current Counts	Dark​Current​Counts	OB or OW	1	DICONDE
(0014,3060)	Gain Correction Reference Sequence	Gain​Correction​Reference​Sequence	SQ	1	DICONDE
(0014,3070)	Air Counts	Air​Counts	OB or OW	1	DICONDE
(0014,3071)	KV Used in Gain Calibration	KV​Used​In​Gain​Calibration	DS	1	DICONDE
(0014,3072)	MA Used in Gain Calibration	MA​Used​In​Gain​Calibration	DS	1	DICONDE
(0014,3073)	Number of Frames Used for Integration	Number​Of​Frames​Used​For​Integration	DS	1	DICONDE
(0014,3074)	Filter Material Used in Gain Calibration	Filter​Material​Used​In​Gain​Calibration	LO	1	DICONDE
(0014,3075)	Filter Thickness Used in Gain Calibration	Filter​Thickness​Used​In​Gain​Calibration	DS	1	DICONDE
(0014,3076)	Date of Gain Calibration	Date​Of​Gain​Calibration	DA	1	DICONDE
(0014,3077)	Time of Gain Calibration	Time​Of​Gain​Calibration	TM	1	DICONDE
(0014,3080)	Bad Pixel Image	Bad​Pixel​Image	OB	1	DICONDE
(0014,3099)	Calibration Notes	Calibration​Notes	LT	1	DICONDE
(0014,4002)	Pulser Equipment Sequence	Pulser​Equipment​Sequence	SQ	1	DICONDE
(0014,4004)	Pulser Type	Pulser​Type	CS	1	DICONDE
(0014,4006)	Pulser Notes	Pulser​Notes	LT	1	DICONDE
(0014,4008)	Receiver Equipment Sequence	Receiver​Equipment​Sequence	SQ	1	DICONDE
(0014,400A)	Amplifier Type	Amplifier​Type	CS	1	DICONDE
(0014,400C)	Receiver Notes	Receiver​Notes	LT	1	DICONDE
(0014,400E)	Pre-Amplifier Equipment Sequence	Pre​Amplifier​Equipment​Sequence	SQ	1	DICONDE
(0014,400F)	Pre-Amplifier Notes	Pre​Amplifier​Notes	LT	1	DICONDE
(0014,4010)	Transmit Transducer Sequence	Transmit​Transducer​Sequence	SQ	1	DICONDE
(0014,4011)	Receive Transducer Sequence	Receive​Transducer​Sequence	SQ	1	DICONDE
(0014,4012)	Number of Elements	Number​Of​Elements	US	1	DICONDE
(0014,4013)	Element Shape	Element​Shape	CS	1	DICONDE
(0014,4014)	Element Dimension A	Element​Dimension​A	DS	1	DICONDE
(0014,4015)	Element Dimension B	Element​Dimension​B	DS	1	DICONDE
(0014,4016)	Element Pitch A	Element​Pitch​A	DS	1	DICONDE
(0014,4017)	Measured Beam Dimension A	Measured​Beam​Dimension​A	DS	1	DICONDE
(0014,4018)	Measured Beam Dimension B	Measured​Beam​Dimension​B	DS	1	DICONDE
(0014,4019)	Location of Measured Beam Diameter	Location​Of​Measured​Beam​Diameter	DS	1	DICONDE
(0014,401A)	Nominal Frequency	Nominal​Frequency	DS	1	DICONDE
(0014,401B)	Measured Center Frequency	Measured​Center​Frequency	DS	1	DICONDE
(0014,401C)	Measured Bandwidth	Measured​Bandwidth	DS	1	DICONDE
(0014,401D)	Element Pitch B	Element​Pitch​B	DS	1	DICONDE
(0014,4020)	Pulser Settings Sequence	Pulser​Settings​Sequence	SQ	1	DICONDE
(0014,4022)	Pulse Width	Pulse​Width	DS	1	DICONDE
(0014,4024)	Excitation Frequency	Excitation​Frequency	DS	1	DICONDE
(0014,4026)	Modulation Type	Modulation​Type	CS	1	DICONDE
(0014,4028)	Damping	Damping	DS	1	DICONDE
(0014,4030)	Receiver Settings Sequence	Receiver​Settings​Sequence	SQ	1	DICONDE
(0014,4031)	Acquired Soundpath Length	Acquired​Soundpath​Length	DS	1	DICONDE
(0014,4032)	Acquisition Compression Type	Acquisition​Compression​Type	CS	1	DICONDE
(0014,4033)	Acquisition Sample Size	Acquisition​Sample​Size	IS	1	DICONDE
(0014,4034)	Rectifier Smoothing	Rectifier​Smoothing	DS	1	DICONDE
(0014,4035)	DAC Sequence	DAC​Sequence	SQ	1	DICONDE
(0014,4036)	DAC Type	DAC​Type	CS	1	DICONDE
(0014,4038)	DAC Gain Points	DAC​Gain​Points	DS	1-n	DICONDE
(0014,403A)	DAC Time Points	DAC​Time​Points	DS	1-n	DICONDE
(0014,403C)	DAC Amplitude	DAC​Amplitude	DS	1-n	DICONDE
(0014,4040)	Pre-Amplifier Settings Sequence	Pre​Amplifier​Settings​Sequence	SQ	1	DICONDE
(0014,4050)	Transmit Transducer Settings Sequence	Transmit​Transducer​Settings​Sequence	SQ	1	DICONDE
(0014,4051)	Receive Transducer Settings Sequence	Receive​Transducer​Settings​Sequence	SQ	1	DICONDE
(0014,4052)	Incident Angle	Incident​Angle	DS	1	DICONDE
(0014,4054)	Coupling Technique	Coupling​Technique	ST	1	DICONDE
(0014,4056)	Coupling Medium	Coupling​Medium	ST	1	DICONDE
(0014,4057)	Coupling Velocity	Coupling​Velocity	DS	1	DICONDE
(0014,4058)	Probe Center Location X	Probe​Center​Location​X	DS	1	DICONDE
(0014,4059)	Probe Center Location Z	Probe​Center​Location​Z	DS	1	DICONDE
(0014,405A)	Sound Path Length	Sound​Path​Length	DS	1	DICONDE
(0014,405C)	Delay Law Identifier	Delay​Law​Identifier	ST	1	DICONDE
(0014,4060)	Gate Settings Sequence	Gate​Settings​Sequence	SQ	1	DICONDE
(0014,4062)	Gate Threshold	Gate​Threshold	DS	1	DICONDE
(0014,4064)	Velocity of Sound	Velocity​Of​Sound	DS	1	DICONDE
(0014,4070)	Calibration Settings Sequence	Calibration​Settings​Sequence	SQ	1	DICONDE
(0014,4072)	Calibration Procedure	Calibration​Procedure	ST	1	DICONDE
(0014,4074)	Procedure Version	Procedure​Version	SH	1	DICONDE
(0014,4076)	Procedure Creation Date	Procedure​Creation​Date	DA	1	DICONDE
(0014,4078)	Procedure Expiration Date	Procedure​Expiration​Date	DA	1	DICONDE
(0014,407A)	Procedure Last Modified Date	Procedure​Last​Modified​Date	DA	1	DICONDE
(0014,407C)	Calibration Time	Calibration​Time	TM	1-n	DICONDE
(0014,407E)	Calibration Date	Calibration​Date	DA	1-n	DICONDE
(0014,4080)	Probe Drive Equipment Sequence	Probe​Drive​Equipment​Sequence	SQ	1	DICONDE
(0014,4081)	Drive Type	Drive​Type	CS	1	DICONDE
(0014,4082)	Probe Drive Notes	Probe​Drive​Notes	LT	1	DICONDE
(0014,4083)	Drive Probe Sequence	Drive​Probe​Sequence	SQ	1	DICONDE
(0014,4084)	Probe Inductance	Probe​Inductance	DS	1	DICONDE
(0014,4085)	Probe Resistance	Probe​Resistance	DS	1	DICONDE
(0014,4086)	Receive Probe Sequence	Receive​Probe​Sequence	SQ	1	DICONDE
(0014,4087)	Probe Drive Settings Sequence	Probe​Drive​Settings​Sequence	SQ	1	DICONDE
(0014,4088)	Bridge Resistors	Bridge​Resistors	DS	1	DICONDE
(0014,4089)	Probe Orientation Angle	Probe​Orientation​Angle	DS	1	DICONDE
(0014,408B)	User Selected Gain Y	User​Selected​Gain​Y	DS	1	DICONDE
(0014,408C)	User Selected Phase	User​Selected​Phase	DS	1	DICONDE
(0014,408D)	User Selected Offset X	User​Selected​Offset​X	DS	1	DICONDE
(0014,408E)	User Selected Offset Y	User​Selected​Offset​Y	DS	1	DICONDE
(0014,4091)	Channel Settings Sequence	Channel​Settings​Sequence	SQ	1	DICONDE
(0014,4092)	Channel Threshold	Channel​Threshold	DS	1	DICONDE
(0014,409A)	Scanner Settings Sequence	Scanner​Settings​Sequence	SQ	1	DICONDE
(0014,409B)	Scan Procedure	Scan​Procedure	ST	1	DICONDE
(0014,409C)	Translation Rate X	Translation​Rate​X	DS	1	DICONDE
(0014,409D)	Translation Rate Y	Translation​Rate​Y	DS	1	DICONDE
(0014,409F)	Channel Overlap	Channel​Overlap	DS	1	DICONDE
(0014,40A0)	Image Quality Indicator Type	Image​Quality​Indicator​Type	LO	1	DICONDE
(0014,40A1)	Image Quality Indicator Material	Image​Quality​Indicator​Material	LO	1	DICONDE
(0014,40A2)	Image Quality Indicator Size	Image​Quality​Indicator​Size	LO	1	DICONDE
(0014,5002)	LINAC Energy	LINAC​Energy	IS	1	DICONDE
(0014,5004)	LINAC Output	LINAC​Output	IS	1	DICONDE
(0014,5100)	Active Aperture	Active​Aperture	US	1	DICONDE
(0014,5101)	Total Aperture	Total​Aperture	DS	1	DICONDE
(0014,5102)	Aperture Elevation	Aperture​Elevation	DS	1	DICONDE
(0014,5103)	Main Lobe Angle	Main​Lobe​Angle	DS	1	DICONDE
(0014,5104)	Main Roof Angle	Main​Roof​Angle	DS	1	DICONDE
(0014,5105)	Connector Type	Connector​Type	CS	1	DICONDE
(0014,5106)	Wedge Model Number	Wedge​Model​Number	SH	1	DICONDE
(0014,5107)	Wedge Angle Float	Wedge​Angle​Float	DS	1	DICONDE
(0014,5108)	Wedge Roof Angle	Wedge​Roof​Angle	DS	1	DICONDE
(0014,5109)	Wedge Element 1 Position	Wedge​Element​1​Position	CS	1	DICONDE
(0014,510A)	Wedge Material Velocity	Wedge​Material​Velocity	DS	1	DICONDE
(0014,510B)	Wedge Material	Wedge​Material	SH	1	DICONDE
(0014,510C)	Wedge Offset Z	Wedge​Offset​Z	DS	1	DICONDE
(0014,510D)	Wedge Origin Offset X	Wedge​Origin​Offset​X	DS	1	DICONDE
(0014,510E)	Wedge Time Delay	Wedge​Time​Delay	DS	1	DICONDE
(0014,510F)	Wedge Name	Wedge​Name	SH	1	DICONDE
(0014,5110)	Wedge Manufacturer Name	Wedge​Manufacturer​Name	SH	1	DICONDE
(0014,5111)	Wedge Description	Wedge​Description	LO	1	DICONDE
(0014,5112)	Nominal Beam Angle	Nominal​Beam​Angle	DS	1	DICONDE
(0014,5113)	Wedge Offset X	Wedge​Offset​X	DS	1	DICONDE
(0014,5114)	Wedge Offset Y	Wedge​Offset​Y	DS	1	DICONDE
(0014,5115)	Wedge Total Length	Wedge​Total​Length	DS	1	DICONDE
(0014,5116)	Wedge In Contact Length	Wedge​In​Contact​Length	DS	1	DICONDE
(0014,5117)	Wedge Front Gap	Wedge​Front​Gap	DS	1	DICONDE
(0014,5118)	Wedge Total Height	Wedge​Total​Height	DS	1	DICONDE
(0014,5119)	Wedge Front Height	Wedge​Front​Height	DS	1	DICONDE
(0014,511A)	Wedge Rear Height	Wedge​Rear​Height	DS	1	DICONDE
(0014,511B)	Wedge Total Width	Wedge​Total​Width	DS	1	DICONDE
(0014,511C)	Wedge In Contact Width	Wedge​In​Contact​Width	DS	1	DICONDE
(0014,511D)	Wedge Chamfer Height	Wedge​Chamfer​Height	DS	1	DICONDE
(0014,511E)	Wedge Curve	Wedge​Curve	CS	1	DICONDE
(0014,511F)	Radius Along the Wedge	Radius​Along​Wedge	DS	1	DICONDE
(0016,0001)	White Point	White​Point​	DS	1
(0016,0002)	Primary Chromaticities	Primary​Chromaticities	DS	3
(0016,0003)	Battery Level	Battery​Level	UT	1
(0016,0004)	Exposure Time in Seconds	Exposure​Time​In​Seconds	DS	1
(0016,0005)	F-Number	FNumber	DS	1
(0016,0006)	OECF Rows	OECF​Rows	IS	1
(0016,0007)	OECF Columns	OECF​Columns	IS	1
(0016,0008)	OECF Column Names	OECF​Column​Names	UC	1-n
(0016,0009)	OECF Values	OECF​Values	DS	1-n
(0016,000A)	Spatial Frequency Response Rows	Spatial​Frequency​Response​Rows	IS	1
(0016,000B)	Spatial Frequency Response Columns	Spatial​Frequency​Response​Columns	IS	1
(0016,000C)	Spatial Frequency Response Column Names	Spatial​Frequency​Response​Column​Names	UC	1-n
(0016,000D)	Spatial Frequency Response Values	Spatial​Frequency​Response​Values	DS	1-n
(0016,000E)	Color Filter Array Pattern Rows	Color​Filter​Array​Pattern​Rows	IS	1
(0016,000F)	Color Filter Array Pattern Columns	Color​Filter​Array​Pattern​Columns	IS	1
(0016,0010)	Color Filter Array Pattern Values	Color​Filter​Array​Pattern​Values	DS	1-n
(0016,0011)	Flash Firing Status	Flash​Firing​Status	US	1
(0016,0012)	Flash Return Status	Flash​Return​Status	US	1
(0016,0013)	Flash Mode	Flash​Mode	US	1
(0016,0014)	Flash Function Present	Flash​Function​Present	US	1
(0016,0015)	Flash Red Eye Mode	Flash​Red​Eye​Mode	US	1
(0016,0016)	Exposure Program	Exposure​Program	US	1
(0016,0017)	Spectral Sensitivity	Spectral​Sensitivity	UT	1
(0016,0018)	Photographic Sensitivity	Photographic​Sensitivity	IS	1
(0016,0019)	Self Timer Mode	Self​Timer​Mode	IS	1
(0016,001A)	Sensitivity Type	Sensitivity​Type	US	1
(0016,001B)	Standard Output Sensitivity	Standard​Output​Sensitivity	IS	1
(0016,001C)	Recommended Exposure Index	Recommended​Exposure​Index	IS	1
(0016,001D)	ISO Speed​	ISO​Speed​	IS	1
(0016,001E)	ISO Speed​​ Latitude yyy	ISO​Speed​Latitude​yyy	IS	1
(0016,001F)	ISO Speed​​ Latitude zzz	ISO​Speed​Latitude​zzz	IS	1
(0016,0020)	EXIF Version	EXIF​Version	UT	1
(0016,0021)	Shutter Speed Value	Shutter​Speed​Value	DS	1
(0016,0022)	Aperture Value	Aperture​Value	DS	1
(0016,0023)	Brightness Value	Brightness​Value	DS	1
(0016,0024)	Exposure Bias Value	Exposure​Bias​Value	DS	1
(0016,0025)	Max Aperture Value	Max​Aperture​Value	DS	1
(0016,0026)	Subject Distance	Subject​Distance	DS	1
(0016,0027)	Metering Mode	Metering​Mode	US	1
(0016,0028)	Light Source	Light​Source	US	1
(0016,0029)	Focal Length	Focal​Length	DS	1
(0016,002A)	Subject Area	Subject​Area	IS	2-4
(0016,002B)	Maker Note	Maker​Note	OB	1
(0016,0030)	Temperature	Temperature	DS	1
(0016,0031)	Humidity	Humidity	DS	1
(0016,0032)	Pressure	Pressure	DS	1
(0016,0033)	Water Depth	Water​Depth	DS	1
(0016,0034)	Acceleration	Acceleration	DS	1
(0016,0035)	Camera Elevation Angle	Camera​Elevation​Angle	DS	1
(0016,0036)	Flash Energy	Flash​Energy	DS	1-2
(0016,0037)	Subject Location	Subject​Location	IS	2
(0016,0038)	Photographic Exposure Index	Photographic​Exposure​Index	DS	1
(0016,0039)	Sensing Method	Sensing​Method	US	1
(0016,003A)	File Source	File​Source	US	1
(0016,003B)	Scene Type	Scene​Type	US	1
(0016,0041)	Custom Rendered	Custom​Rendered	US	1
(0016,0042)	Exposure Mode	Exposure​Mode	US	1
(0016,0043)	White Balance	White​Balance	US	1
(0016,0044)	Digital Zoom Ratio	Digital​Zoom​Ratio	DS	1
(0016,0045)	Focal Length In 35mm​ Film	Focal​Length​In​35mm​Film	IS	1
(0016,0046)	Scene Capture Type	Scene​Capture​Type	US	1
(0016,0047)	Gain Control	Gain​Control	US	1
(0016,0048)	Contrast	Contrast	US	1
(0016,0049)	Saturation	Saturation	US	1
(0016,004A)	Sharpness	Sharpness	US	1
(0016,004B)	Device Setting Description	Device​Setting​Description	OB	1
(0016,004C)	Subject Distance Range	Subject​Distance​Range	US	1
(0016,004D)	Camera Owner Name	Camera​Owner​Name	UT	1
(0016,004E)	Lens Specification	Lens​Specification	DS	4
(0016,004F)	Lens Make	Lens​Make	UT	1
(0016,0050)	Lens Model	Lens​Model	UT	1
(0016,0051)	Lens Serial Number	Lens​Serial​Number	UT	1
(0016,0061)	Interoperability Index	Interoperability​Index	CS	1
(0016,0062)	Interoperability Version	Interoperability​Version	OB	1
(0016,0070)	GPS Version ID	GPS​Version​ID	OB	1
(0016,0071)	GPS Latitude​ Ref	GPS​Latitude​Ref	CS	1
(0016,0072)	GPS Latitude​	GPS​Latitude​	DS	3
(0016,0073)	GPS Longitude Ref	GPS​Longitude​​Ref	CS	1
(0016,0074)	GPS Longitude	GPS​Longitude	DS	3
(0016,0075)	GPS Altitude​ Ref	GPS​Altitude​Ref	US	1
(0016,0076)	GPS Altitude​	GPS​Altitude​	DS	1
(0016,0077)	GPS Time​ Stamp	GPS​Time​Stamp	DT	1
(0016,0078)	GPS Satellites	GPS​Satellites	UT	1
(0016,0079)	GPS Status	GPS​Status	CS	1
(0016,007A)	GPS Measure ​Mode	GPS​Measure​​Mode	CS	1
(0016,007B)	GPS DOP	GPS​DOP	DS	1
(0016,007C)	GPS Speed​ Ref	GPS​Speed​Ref	CS	1
(0016,007D)	GPS Speed​	GPS​Speed​	DS	1
(0016,007E)	GPS Track ​Ref	GPS​Track​​Ref	CS	1
(0016,007F)	GPS Track	GPS​Track	DS	1
(0016,0080)	GPS Img​ Direction Ref	GPS​Img​Direction​Ref	CS	1
(0016,0081)	GPS Img​ Direction	GPS​Img​Direction	DS	1
(0016,0082)	GPS Map​ Datum	GPS​Map​​Datum	UT	1
(0016,0083)	GPS Dest ​Latitude Ref	GPS​Dest​Latitude​Ref	CS	1
(0016,0084)	GPS Dest​ Latitude	GPS​Dest​Latitude	DS	3
(0016,0085)	GPS Dest​ Longitude Ref	GPS​Dest​Longitude​Ref	CS	1
(0016,0086)	GPS Dest ​Longitude	GPS​Dest​​Longitude	DS	3
(0016,0087)	GPS Dest ​Bearing Ref	GPS​Dest​Bearing​Ref	CS	1
(0016,0088)	GPS Dest ​Bearing	GPS​Dest​Bearing	DS	1
(0016,0089)	GPS Dest ​Distance Ref	GPS​Dest​Distance​Ref	CS	1
(0016,008A)	GPS Dest​ Distance	GPS​Dest​Distance	DS	1
(0016,008B)	GPS Processing​ Method	GPS​Processing​Method	OB	1
(0016,008C)	GPS Area ​Information	GPS​Area​Information	OB	1
(0016,008D)	GPS Date​ Stamp	GPS​Date​Stamp	DT	1
(0016,008E)	GPS Differential	GPS​Differential	IS	1
(0018,0010)	Contrast/Bolus Agent	Contrast​Bolus​Agent	LO	1
(0018,0012)	Contrast/Bolus Agent Sequence	Contrast​Bolus​Agent​Sequence	SQ	1
(0018,0013)	Contrast/Bolus T1 Relaxivity	Contrast​Bolus​​T1​Relaxivity	FL	1
(0018,0014)	Contrast/Bolus Administration Route Sequence	Contrast​Bolus​Administration​Route​Sequence	SQ	1
(0018,0015)	Body Part Examined	Body​Part​Examined	CS	1
(0018,0020)	Scanning Sequence	Scanning​Sequence	CS	1-n
(0018,0021)	Sequence Variant	Sequence​Variant	CS	1-n
(0018,0022)	Scan Options	Scan​Options	CS	1-n
(0018,0023)	MR Acquisition Type	MR​Acquisition​Type	CS	1
(0018,0024)	Sequence Name	Sequence​Name	SH	1
(0018,0025)	Angio Flag	Angio​Flag	CS	1
(0018,0026)	Intervention Drug Information Sequence	Intervention​Drug​Information​Sequence	SQ	1
(0018,0027)	Intervention Drug Stop Time	Intervention​Drug​Stop​Time	TM	1
(0018,0028)	Intervention Drug Dose	Intervention​Drug​Dose	DS	1
(0018,0029)	Intervention Drug Code Sequence	Intervention​Drug​Code​Sequence	SQ	1
(0018,002A)	Additional Drug Sequence	Additional​Drug​Sequence	SQ	1
(0018,0030)	Radionuclide	Radionuclide	LO	1-n	RET (1993)
(0018,0031)	Radiopharmaceutical	Radio​pharmaceutical	LO	1
(0018,0032)	Energy Window Centerline	Energy​Window​Centerline	DS	1	RET (1993)
(0018,0033)	Energy Window Total Width	Energy​Window​Total​Width	DS	1-n	RET (1993)
(0018,0034)	Intervention Drug Name	Intervention​Drug​Name	LO	1
(0018,0035)	Intervention Drug Start Time	Intervention​Drug​Start​Time	TM	1
(0018,0036)	Intervention Sequence	Intervention​Sequence	SQ	1
(0018,0037)	Therapy Type	Therapy​Type	CS	1	RET (2001)
(0018,0038)	Intervention Status	Intervention​Status	CS	1
(0018,0039)	Therapy Description	Therapy​Description	CS	1	RET (2003)
(0018,003A)	Intervention Description	Intervention​Description	ST	1
(0018,0040)	Cine Rate	Cine​Rate	IS	1
(0018,0042)	Initial Cine Run State	Initial​Cine​Run​State	CS	1
(0018,0050)	Slice Thickness	Slice​Thickness	DS	1
(0018,0060)	KVP	KVP	DS	1
(0018,0061)			DS	1	RET (2015c)
(0018,0070)	Counts Accumulated	Counts​Accumulated	IS	1
(0018,0071)	Acquisition Termination Condition	Acquisition​Termination​Condition	CS	1
(0018,0072)	Effective Duration	Effective​Duration	DS	1
(0018,0073)	Acquisition Start Condition	Acquisition​Start​Condition	CS	1
(0018,0074)	Acquisition Start Condition Data	Acquisition​Start​Condition​Data	IS	1
(0018,0075)	Acquisition Termination Condition Data	Acquisition​Termination​Condition​Data	IS	1
(0018,0080)	Repetition Time	Repetition​Time	DS	1
(0018,0081)	Echo Time	Echo​Time	DS	1
(0018,0082)	Inversion Time	Inversion​Time	DS	1
(0018,0083)	Number of Averages	Number​Of​Averages	DS	1
(0018,0084)	Imaging Frequency	Imaging​Frequency	DS	1
(0018,0085)	Imaged Nucleus	Imaged​Nucleus	SH	1
(0018,0086)	Echo Number(s)	Echo​Numbers	IS	1-n
(0018,0087)	Magnetic Field Strength	Magnetic​Field​Strength	DS	1
(0018,0088)	Spacing Between Slices	Spacing​Between​Slices	DS	1
(0018,0089)	Number of Phase Encoding Steps	Number​Of​Phase​Encoding​Steps	IS	1
(0018,0090)	Data Collection Diameter	Data​Collection​Diameter	DS	1
(0018,0091)	Echo Train Length	Echo​Train​Length	IS	1
(0018,0093)	Percent Sampling	Percent​Sampling	DS	1
(0018,0094)	Percent Phase Field of View	Percent​Phase​Field​Of​View	DS	1
(0018,0095)	Pixel Bandwidth	Pixel​Bandwidth	DS	1
(0018,1000)	Device Serial Number	Device​Serial​Number	LO	1
(0018,1002)	Device UID	Device​UID	UI	1
(0018,1003)	Device ID	Device​ID	LO	1
(0018,1004)	Plate ID	Plate​ID	LO	1
(0018,1005)	Generator ID	Generator​ID	LO	1
(0018,1006)	Grid ID	Grid​ID	LO	1
(0018,1007)	Cassette ID	Cassette​ID	LO	1
(0018,1008)	Gantry ID	Gantry​ID	LO	1
(0018,1009)	Unique Device Identifier	Unique​Device​Identifier	UT	1
(0018,100A)	UDI Sequence	UDI​Sequence	SQ	1
(0018,1010)	Secondary Capture Device ID	Secondary​Capture​Device​ID	LO	1
(0018,1011)	Hardcopy Creation Device ID	Hardcopy​Creation​Device​ID	LO	1	RET (2007)
(0018,1012)	Date of Secondary Capture	Date​Of​Secondary​Capture	DA	1
(0018,1014)	Time of Secondary Capture	Time​Of​Secondary​Capture	TM	1
(0018,1016)	Secondary Capture Device Manufacturer	Secondary​Capture​Device​Manufacturer	LO	1
(0018,1017)	Hardcopy Device Manufacturer	Hardcopy​Device​Manufacturer	LO	1	RET (2007)
(0018,1018)	Secondary Capture Device Manufacturer's Model Name	Secondary​Capture​Device​Manufacturer​Model​Name	LO	1
(0018,1019)	Secondary Capture Device Software Versions	Secondary​Capture​Device​Software​Versions	LO	1-n
(0018,101A)	Hardcopy Device Software Version	Hardcopy​Device​Software​Version	LO	1-n	RET (2007)
(0018,101B)	Hardcopy Device Manufacturer's Model Name	Hardcopy​Device​Manufacturer​Model​Name	LO	1	RET (2007)
(0018,1020)	Software Versions	Software​Versions	LO	1-n
(0018,1022)	Video Image Format Acquired	Video​Image​Format​Acquired	SH	1
(0018,1023)	Digital Image Format Acquired	Digital​Image​Format​Acquired	LO	1
(0018,1030)	Protocol Name	Protocol​Name	LO	1
(0018,1040)	Contrast/Bolus Route	Contrast​Bolus​Route	LO	1
(0018,1041)	Contrast/Bolus Volume	Contrast​Bolus​Volume	DS	1
(0018,1042)	Contrast/Bolus Start Time	Contrast​Bolus​Start​Time	TM	1
(0018,1043)	Contrast/Bolus Stop Time	Contrast​Bolus​Stop​Time	TM	1
(0018,1044)	Contrast/Bolus Total Dose	Contrast​Bolus​Total​Dose	DS	1
(0018,1045)	Syringe Counts	Syringe​Counts	IS	1
(0018,1046)	Contrast Flow Rate	Contrast​Flow​Rate	DS	1-n
(0018,1047)	Contrast Flow Duration	Contrast​Flow​Duration	DS	1-n
(0018,1048)	Contrast/Bolus Ingredient	Contrast​Bolus​Ingredient	CS	1
(0018,1049)	Contrast/Bolus Ingredient Concentration	Contrast​Bolus​Ingredient​Concentration	DS	1
(0018,1050)	Spatial Resolution	Spatial​Resolution	DS	1
(0018,1060)	Trigger Time	Trigger​Time	DS	1
(0018,1061)	Trigger Source or Type	Trigger​Source​Or​Type	LO	1
(0018,1062)	Nominal Interval	Nominal​Interval	IS	1
(0018,1063)	Frame Time	Frame​Time	DS	1
(0018,1064)	Cardiac Framing Type	Cardiac​Framing​Type	LO	1
(0018,1065)	Frame Time Vector	Frame​Time​Vector	DS	1-n
(0018,1066)	Frame Delay	Frame​Delay	DS	1
(0018,1067)	Image Trigger Delay	Image​Trigger​Delay	DS	1
(0018,1068)	Multiplex Group Time Offset	Multiplex​Group​Time​Offset	DS	1
(0018,1069)	Trigger Time Offset	Trigger​Time​Offset	DS	1
(0018,106A)	Synchronization Trigger	Synchronization​Trigger	CS	1
(0018,106C)	Synchronization Channel	Synchronization​Channel	US	2
(0018,106E)	Trigger Sample Position	Trigger​Sample​Position	UL	1
(0018,1070)	Radiopharmaceutical Route	Radio​pharmaceutical​Route	LO	1
(0018,1071)	Radiopharmaceutical Volume	Radio​pharmaceutical​Volume	DS	1
(0018,1072)	Radiopharmaceutical Start Time	Radio​pharmaceutical​Start​Time	TM	1
(0018,1073)	Radiopharmaceutical Stop Time	Radio​pharmaceutical​Stop​Time	TM	1
(0018,1074)	Radionuclide Total Dose	Radionuclide​Total​Dose	DS	1
(0018,1075)	Radionuclide Half Life	Radionuclide​Half​Life	DS	1
(0018,1076)	Radionuclide Positron Fraction	Radionuclide​Positron​Fraction	DS	1
(0018,1077)	Radiopharmaceutical Specific Activity	Radio​pharmaceutical​Specific​Activity	DS	1
(0018,1078)	Radiopharmaceutical Start DateTime	Radio​pharmaceutical​Start​Date​Time	DT	1
(0018,1079)	Radiopharmaceutical Stop DateTime	Radio​pharmaceutical​Stop​Date​Time	DT	1
(0018,1080)	Beat Rejection Flag	Beat​Rejection​Flag	CS	1
(0018,1081)	Low R-R Value	Low​RR​Value	IS	1
(0018,1082)	High R-R Value	High​RR​Value	IS	1
(0018,1083)	Intervals Acquired	Intervals​Acquired	IS	1
(0018,1084)	Intervals Rejected	Intervals​Rejected	IS	1
(0018,1085)	PVC Rejection	PVC​Rejection	LO	1
(0018,1086)	Skip Beats	Skip​Beats	IS	1
(0018,1088)	Heart Rate	Heart​Rate	IS	1
(0018,1090)	Cardiac Number of Images	Cardiac​Number​Of​Images	IS	1
(0018,1094)	Trigger Window	Trigger​Window	IS	1
(0018,1100)	Reconstruction Diameter	Reconstruction​Diameter	DS	1
(0018,1110)	Distance Source to Detector	Distance​Source​To​Detector	DS	1
(0018,1111)	Distance Source to Patient	Distance​Source​To​Patient	DS	1
(0018,1114)	Estimated Radiographic Magnification Factor	Estimated​Radiographic​Magnification​Factor	DS	1
(0018,1120)	Gantry/Detector Tilt	Gantry​Detector​Tilt	DS	1
(0018,1121)	Gantry/Detector Slew	Gantry​Detector​Slew	DS	1
(0018,1130)	Table Height	Table​Height	DS	1
(0018,1131)	Table Traverse	Table​Traverse	DS	1
(0018,1134)	Table Motion	Table​Motion	CS	1
(0018,1135)	Table Vertical Increment	Table​Vertical​Increment	DS	1-n
(0018,1136)	Table Lateral Increment	Table​Lateral​Increment	DS	1-n
(0018,1137)	Table Longitudinal Increment	Table​Longitudinal​Increment	DS	1-n
(0018,1138)	Table Angle	Table​Angle	DS	1
(0018,113A)	Table Type	Table​Type	CS	1
(0018,1140)	Rotation Direction	Rotation​Direction	CS	1
(0018,1141)	Angular Position	Angular​Position	DS	1	RET (2007)
(0018,1142)	Radial Position	Radial​Position	DS	1-n
(0018,1143)	Scan Arc	Scan​Arc	DS	1
(0018,1144)	Angular Step	Angular​Step	DS	1
(0018,1145)	Center of Rotation Offset	Center​Of​Rotation​Offset	DS	1
(0018,1146)	Rotation Offset	Rotation​Offset	DS	1-n	RET (1993)
(0018,1147)	Field of View Shape	Field​Of​View​Shape	CS	1
(0018,1149)	Field of View Dimension(s)	Field​Of​View​Dimensions	IS	1-2
(0018,1150)	Exposure Time	Exposure​Time	IS	1
(0018,1151)	X-Ray Tube Current	X​Ray​Tube​Current	IS	1
(0018,1152)	Exposure	Exposure	IS	1
(0018,1153)	Exposure in µAs	Exposure​Inu​As	IS	1
(0018,1154)	Average Pulse Width	Average​Pulse​Width	DS	1
(0018,1155)	Radiation Setting	Radiation​Setting	CS	1
(0018,1156)	Rectification Type	Rectification​Type	CS	1
(0018,115A)	Radiation Mode	Radiation​Mode	CS	1
(0018,115E)	Image and Fluoroscopy Area Dose Product	Image​And​Fluoroscopy​Area​Dose​Product	DS	1
(0018,1160)	Filter Type	Filter​Type	SH	1
(0018,1161)	Type of Filters	Type​Of​Filters	LO	1-n
(0018,1162)	Intensifier Size	Intensifier​Size	DS	1
(0018,1164)	Imager Pixel Spacing	Imager​Pixel​Spacing	DS	2
(0018,1166)	Grid	Grid	CS	1-n
(0018,1170)	Generator Power	Generator​Power	IS	1
(0018,1180)	Collimator/grid Name	Collimator​Grid​Name	SH	1
(0018,1181)	Collimator Type	Collimator​Type	CS	1
(0018,1182)	Focal Distance	Focal​Distance	IS	1-2
(0018,1183)	X Focus Center	X​Focus​Center	DS	1-2
(0018,1184)	Y Focus Center	Y​Focus​Center	DS	1-2
(0018,1190)	Focal Spot(s)	Focal​Spots	DS	1-n
(0018,1191)	Anode Target Material	Anode​Target​Material	CS	1
(0018,11A0)	Body Part Thickness	Body​Part​Thickness	DS	1
(0018,11A2)	Compression Force	Compression​Force	DS	1
(0018,11A3)	Compression Pressure	Compression​​Pressure	DS	1
(0018,11A4)	Paddle Description	Paddle​Description	LO	1
(0018,11A5)	Compression Contact Area	Compression​​Contact​​Area	DS	1
(0018,1200)	Date of Last Calibration	Date​Of​Last​Calibration	DA	1-n
(0018,1201)	Time of Last Calibration	Time​Of​Last​Calibration	TM	1-n
(0018,1202)	DateTime of Last Calibration	Date​Time​Of​Last​Calibration	DT	1
(0018,1210)	Convolution Kernel	Convolution​Kernel	SH	1-n
(0018,1240)	Upper/Lower Pixel Values	Upper​Lower​Pixel​Values	IS	1-n	RET
(0018,1242)	Actual Frame Duration	Actual​Frame​Duration	IS	1
(0018,1243)	Count Rate	Count​Rate	IS	1
(0018,1244)	Preferred Playback Sequencing	Preferred​Playback​Sequencing	US	1
(0018,1250)	Receive Coil Name	Receive​Coil​Name	SH	1
(0018,1251)	Transmit Coil Name	Transmit​Coil​Name	SH	1
(0018,1260)	Plate Type	Plate​Type	SH	1
(0018,1261)	Phosphor Type	Phosphor​Type	LO	1
(0018,1271)	Water Equivalent Diameter	Water​Equivalent​Diameter	FD	1
(0018,1272)	Water Equivalent Diameter Calculation Method Code Sequence	Water​Equivalent​Diameter​Calculation​Method​Code​Sequence	SQ	1
(0018,1300)	Scan Velocity	Scan​Velocity	DS	1
(0018,1301)	Whole Body Technique	Whole​Body​Technique	CS	1-n
(0018,1302)	Scan Length	Scan​Length	IS	1
(0018,1310)	Acquisition Matrix	Acquisition​Matrix	US	4
(0018,1312)	In-plane Phase Encoding Direction	In​Plane​Phase​Encoding​Direction	CS	1
(0018,1314)	Flip Angle	Flip​Angle	DS	1
(0018,1315)	Variable Flip Angle Flag	Variable​Flip​Angle​Flag	CS	1
(0018,1316)	SAR	SAR	DS	1
(0018,1318)	dB/dt	d​Bdt	DS	1
(0018,1320)	B1rms	B1rms	FL	1
(0018,1400)	Acquisition Device Processing Description	Acquisition​Device​Processing​Description	LO	1
(0018,1401)	Acquisition Device Processing Code	Acquisition​Device​Processing​Code	LO	1
(0018,1402)	Cassette Orientation	Cassette​Orientation	CS	1
(0018,1403)	Cassette Size	Cassette​Size	CS	1
(0018,1404)	Exposures on Plate	Exposures​On​Plate	US	1
(0018,1405)	Relative X-Ray Exposure	Relative​X​Ray​Exposure	IS	1
(0018,1411)	Exposure Index	Exposure​Index	DS	1
(0018,1412)	Target Exposure Index	Target​Exposure​Index	DS	1
(0018,1413)	Deviation Index	Deviation​Index	DS	1
(0018,1450)	Column Angulation	Column​Angulation	DS	1
(0018,1460)	Tomo Layer Height	Tomo​Layer​Height	DS	1
(0018,1470)	Tomo Angle	Tomo​Angle	DS	1
(0018,1480)	Tomo Time	Tomo​Time	DS	1
(0018,1490)	Tomo Type	Tomo​Type	CS	1
(0018,1491)	Tomo Class	Tomo​Class	CS	1
(0018,1495)	Number of Tomosynthesis Source Images	Number​Of​Tomosynthesis​Source​Images	IS	1
(0018,1500)	Positioner Motion	Positioner​Motion	CS	1
(0018,1508)	Positioner Type	Positioner​Type	CS	1
(0018,1510)	Positioner Primary Angle	Positioner​Primary​Angle	DS	1
(0018,1511)	Positioner Secondary Angle	Positioner​Secondary​Angle	DS	1
(0018,1520)	Positioner Primary Angle Increment	Positioner​Primary​Angle​Increment	DS	1-n
(0018,1521)	Positioner Secondary Angle Increment	Positioner​Secondary​Angle​Increment	DS	1-n
(0018,1530)	Detector Primary Angle	Detector​Primary​Angle	DS	1
(0018,1531)	Detector Secondary Angle	Detector​Secondary​Angle	DS	1
(0018,1600)	Shutter Shape	Shutter​Shape	CS	1-3
(0018,1602)	Shutter Left Vertical Edge	Shutter​Left​Vertical​Edge	IS	1
(0018,1604)	Shutter Right Vertical Edge	Shutter​Right​Vertical​Edge	IS	1
(0018,1606)	Shutter Upper Horizontal Edge	Shutter​Upper​Horizontal​Edge	IS	1
(0018,1608)	Shutter Lower Horizontal Edge	Shutter​Lower​Horizontal​Edge	IS	1
(0018,1610)	Center of Circular Shutter	Center​Of​Circular​Shutter	IS	2
(0018,1612)	Radius of Circular Shutter	Radius​Of​Circular​Shutter	IS	1
(0018,1620)	Vertices of the Polygonal Shutter	Vertices​Of​The​Polygonal​Shutter	IS	2-2n
(0018,1622)	Shutter Presentation Value	Shutter​Presentation​Value	US	1
(0018,1623)	Shutter Overlay Group	Shutter​Overlay​Group	US	1
(0018,1624)	Shutter Presentation Color CIELab Value	Shutter​Presentation​Color​CIE​Lab​Value	US	3
(0018,1700)	Collimator Shape	Collimator​Shape	CS	1-3
(0018,1702)	Collimator Left Vertical Edge	Collimator​Left​Vertical​Edge	IS	1
(0018,1704)	Collimator Right Vertical Edge	Collimator​Right​Vertical​Edge	IS	1
(0018,1706)	Collimator Upper Horizontal Edge	Collimator​Upper​Horizontal​Edge	IS	1
(0018,1708)	Collimator Lower Horizontal Edge	Collimator​Lower​Horizontal​Edge	IS	1
(0018,1710)	Center of Circular Collimator	Center​Of​Circular​Collimator	IS	2
(0018,1712)	Radius of Circular Collimator	Radius​Of​Circular​Collimator	IS	1
(0018,1720)	Vertices of the Polygonal Collimator	Vertices​Of​The​Polygonal​Collimator	IS	2-2n
(0018,1800)	Acquisition Time Synchronized	Acquisition​Time​Synchronized	CS	1
(0018,1801)	Time Source	Time​Source	SH	1
(0018,1802)	Time Distribution Protocol	Time​Distribution​Protocol	CS	1
(0018,1803)	NTP Source Address	NTP​Source​Address	LO	1
(0018,2001)	Page Number Vector	Page​Number​Vector	IS	1-n
(0018,2002)	Frame Label Vector	Frame​Label​Vector	SH	1-n
(0018,2003)	Frame Primary Angle Vector	Frame​Primary​Angle​Vector	DS	1-n
(0018,2004)	Frame Secondary Angle Vector	Frame​Secondary​Angle​Vector	DS	1-n
(0018,2005)	Slice Location Vector	Slice​Location​Vector	DS	1-n
(0018,2006)	Display Window Label Vector	Display​Window​Label​Vector	SH	1-n
(0018,2010)	Nominal Scanned Pixel Spacing	Nominal​Scanned​Pixel​Spacing	DS	2
(0018,2020)	Digitizing Device Transport Direction	Digitizing​Device​Transport​Direction	CS	1
(0018,2030)	Rotation of Scanned Film	Rotation​Of​Scanned​Film	DS	1
(0018,2041)	Biopsy Target Sequence	Biopsy​Target​Sequence	SQ	1
(0018,2042)	Target UID	Target​UID	UI	1
(0018,2043)	Localizing Cursor Position	Localizing​Cursor​Position	FL	2
(0018,2044)	Calculated Target Position	Calculated​Target​Position	FL	3
(0018,2045)	Target Label	Target​Label	SH	1
(0018,2046)	Displayed Z Value	Displayed​Z​Value	FL	1
(0018,3100)	IVUS Acquisition	IVUS​Acquisition	CS	1
(0018,3101)	IVUS Pullback Rate	IVUS​Pullback​Rate	DS	1
(0018,3102)	IVUS Gated Rate	IVUS​Gated​Rate	DS	1
(0018,3103)	IVUS Pullback Start Frame Number	IVUS​Pullback​Start​Frame​Number	IS	1
(0018,3104)	IVUS Pullback Stop Frame Number	IVUS​Pullback​Stop​Frame​Number	IS	1
(0018,3105)	Lesion Number	Lesion​Number	IS	1-n
(0018,4000)	Acquisition Comments	Acquisition​Comments	LT	1	RET
(0018,5000)	Output Power	Output​Power	SH	1-n
(0018,5010)	Transducer Data	Transducer​Data	LO	1-n
(0018,5012)	Focus Depth	Focus​Depth	DS	1
(0018,5020)	Processing Function	Processing​Function	LO	1
(0018,5021)	Postprocessing Function	Postprocessing​Function	LO	1	RET (2007)
(0018,5022)	Mechanical Index	Mechanical​Index	DS	1
(0018,5024)	Bone Thermal Index	Bone​Thermal​Index	DS	1
(0018,5026)	Cranial Thermal Index	Cranial​Thermal​Index	DS	1
(0018,5027)	Soft Tissue Thermal Index	Soft​Tissue​Thermal​Index	DS	1
(0018,5028)	Soft Tissue-focus Thermal Index	Soft​Tissue​Focus​Thermal​Index	DS	1
(0018,5029)	Soft Tissue-surface Thermal Index	Soft​Tissue​Surface​Thermal​Index	DS	1
(0018,5030)	Dynamic Range	Dynamic​Range	DS	1	RET
(0018,5040)	Total Gain	Total​Gain	DS	1	RET
(0018,5050)	Depth of Scan Field	Depth​Of​Scan​Field	IS	1
(0018,5100)	Patient Position	Patient​Position	CS	1
(0018,5101)	View Position	View​Position	CS	1
(0018,5104)	Projection Eponymous Name Code Sequence	Projection​Eponymous​Name​Code​Sequence	SQ	1
(0018,5210)	Image Transformation Matrix	Image​Transformation​Matrix	DS	6	RET (2003)
(0018,5212)	Image Translation Vector	Image​Translation​Vector	DS	3	RET (2003)
(0018,6000)	Sensitivity	Sensitivity	DS	1
(0018,6011)	Sequence of Ultrasound Regions	Sequence​Of​Ultrasound​Regions	SQ	1
(0018,6012)	Region Spatial Format	Region​Spatial​Format	US	1
(0018,6014)	Region Data Type	Region​Data​Type	US	1
(0018,6016)	Region Flags	Region​Flags	UL	1
(0018,6018)	Region Location Min X0	Region​Location​Min​X0	UL	1
(0018,601A)	Region Location Min Y0	Region​Location​Min​Y0	UL	1
(0018,601C)	Region Location Max X1	Region​Location​Max​X1	UL	1
(0018,601E)	Region Location Max Y1	Region​Location​Max​Y1	UL	1
(0018,6020)	Reference Pixel X0	Reference​Pixel​X0	SL	1
(0018,6022)	Reference Pixel Y0	Reference​Pixel​Y0	SL	1
(0018,6024)	Physical Units X Direction	Physical​Units​X​Direction	US	1
(0018,6026)	Physical Units Y Direction	Physical​Units​Y​Direction	US	1
(0018,6028)	Reference Pixel Physical Value X	Reference​Pixel​Physical​Value​X	FD	1
(0018,602A)	Reference Pixel Physical Value Y	Reference​Pixel​Physical​Value​Y	FD	1
(0018,602C)	Physical Delta X	Physical​Delta​X	FD	1
(0018,602E)	Physical Delta Y	Physical​Delta​Y	FD	1
(0018,6030)	Transducer Frequency	Transducer​Frequency	UL	1
(0018,6031)	Transducer Type	Transducer​Type	CS	1
(0018,6032)	Pulse Repetition Frequency	Pulse​Repetition​Frequency	UL	1
(0018,6034)	Doppler Correction Angle	Doppler​Correction​Angle	FD	1
(0018,6036)	Steering Angle	Steering​Angle	FD	1
(0018,6038)	Doppler Sample Volume X Position (Retired)	Doppler​Sample​Volume​X​Position​Retired	UL	1	RET (2001)
(0018,6039)	Doppler Sample Volume X Position	Doppler​Sample​Volume​X​Position	SL	1
(0018,603A)	Doppler Sample Volume Y Position (Retired)	Doppler​Sample​Volume​Y​Position​Retired	UL	1	RET (2001)
(0018,603B)	Doppler Sample Volume Y Position	Doppler​Sample​Volume​Y​Position	SL	1
(0018,603C)	TM-Line Position X0 (Retired)	TM​Line​Position​X0Retired	UL	1	RET (2001)
(0018,603D)	TM-Line Position X0	TM​Line​Position​X0	SL	1
(0018,603E)	TM-Line Position Y0 (Retired)	TM​Line​Position​Y0Retired	UL	1	RET (2001)
(0018,603F)	TM-Line Position Y0	TM​Line​Position​Y0	SL	1
(0018,6040)	TM-Line Position X1 (Retired)	TM​Line​Position​X1Retired	UL	1	RET (2001)
(0018,6041)	TM-Line Position X1	TM​Line​Position​X1	SL	1
(0018,6042)	TM-Line Position Y1 (Retired)	TM​Line​Position​Y1Retired	UL	1	RET (2001)
(0018,6043)	TM-Line Position Y1	TM​Line​Position​Y1	SL	1
(0018,6044)	Pixel Component Organization	Pixel​Component​Organization	US	1
(0018,6046)	Pixel Component Mask	Pixel​Component​Mask	UL	1
(0018,6048)	Pixel Component Range Start	Pixel​Component​Range​Start	UL	1
(0018,604A)	Pixel Component Range Stop	Pixel​Component​Range​Stop	UL	1
(0018,604C)	Pixel Component Physical Units	Pixel​Component​Physical​Units	US	1
(0018,604E)	Pixel Component Data Type	Pixel​Component​Data​Type	US	1
(0018,6050)	Number of Table Break Points	Number​Of​Table​Break​Points	UL	1
(0018,6052)	Table of X Break Points	Table​Of​X​Break​Points	UL	1-n
(0018,6054)	Table of Y Break Points	Table​Of​Y​Break​Points	FD	1-n
(0018,6056)	Number of Table Entries	Number​Of​Table​Entries	UL	1
(0018,6058)	Table of Pixel Values	Table​Of​Pixel​Values	UL	1-n
(0018,605A)	Table of Parameter Values	Table​Of​Parameter​Values	FL	1-n
(0018,6060)	R Wave Time Vector	R​Wave​Time​Vector	FL	1-n
(0018,7000)	Detector Conditions Nominal Flag	Detector​Conditions​Nominal​Flag	CS	1
(0018,7001)	Detector Temperature	Detector​Temperature	DS	1
(0018,7004)	Detector Type	Detector​Type	CS	1
(0018,7005)	Detector Configuration	Detector​Configuration	CS	1
(0018,7006)	Detector Description	Detector​Description	LT	1
(0018,7008)	Detector Mode	Detector​Mode	LT	1
(0018,700A)	Detector ID	Detector​ID	SH	1
(0018,700C)	Date of Last Detector Calibration	Date​Of​Last​Detector​Calibration	DA	1
(0018,700E)	Time of Last Detector Calibration	Time​Of​Last​Detector​Calibration	TM	1
(0018,7010)	Exposures on Detector Since Last Calibration	Exposures​On​Detector​Since​Last​Calibration	IS	1
(0018,7011)	Exposures on Detector Since Manufactured	Exposures​On​Detector​Since​Manufactured	IS	1
(0018,7012)	Detector Time Since Last Exposure	Detector​Time​Since​Last​Exposure	DS	1
(0018,7014)	Detector Active Time	Detector​Active​Time	DS	1
(0018,7016)	Detector Activation Offset From Exposure	Detector​Activation​Offset​From​Exposure	DS	1
(0018,701A)	Detector Binning	Detector​Binning	DS	2
(0018,7020)	Detector Element Physical Size	Detector​Element​Physical​Size	DS	2
(0018,7022)	Detector Element Spacing	Detector​Element​Spacing	DS	2
(0018,7024)	Detector Active Shape	Detector​Active​Shape	CS	1
(0018,7026)	Detector Active Dimension(s)	Detector​Active​Dimensions	DS	1-2
(0018,7028)	Detector Active Origin	Detector​Active​Origin	DS	2
(0018,702A)	Detector Manufacturer Name	Detector​Manufacturer​Name	LO	1
(0018,702B)	Detector Manufacturer's Model Name	Detector​Manufacturer​Model​Name	LO	1
(0018,7030)	Field of View Origin	Field​Of​View​Origin	DS	2
(0018,7032)	Field of View Rotation	Field​Of​View​Rotation	DS	1
(0018,7034)	Field of View Horizontal Flip	Field​Of​View​Horizontal​Flip	CS	1
(0018,7036)	Pixel Data Area Origin Relative To FOV	Pixel​Data​Area​Origin​Relative​To​FOV	FL	2
(0018,7038)	Pixel Data Area Rotation Angle Relative To FOV	Pixel​Data​Area​Rotation​Angle​Relative​To​FOV	FL	1
(0018,7040)	Grid Absorbing Material	Grid​Absorbing​Material	LT	1
(0018,7041)	Grid Spacing Material	Grid​Spacing​Material	LT	1
(0018,7042)	Grid Thickness	Grid​Thickness	DS	1
(0018,7044)	Grid Pitch	Grid​Pitch	DS	1
(0018,7046)	Grid Aspect Ratio	Grid​Aspect​Ratio	IS	2
(0018,7048)	Grid Period	Grid​Period	DS	1
(0018,704C)	Grid Focal Distance	Grid​Focal​Distance	DS	1
(0018,7050)	Filter Material	Filter​Material	CS	1-n
(0018,7052)	Filter Thickness Minimum	Filter​Thickness​Minimum	DS	1-n
(0018,7054)	Filter Thickness Maximum	Filter​Thickness​Maximum	DS	1-n
(0018,7056)	Filter Beam Path Length Minimum	Filter​Beam​Path​Length​Minimum	FL	1-n
(0018,7058)	Filter Beam Path Length Maximum	Filter​Beam​Path​Length​Maximum	FL	1-n
(0018,7060)	Exposure Control Mode	Exposure​Control​Mode	CS	1
(0018,7062)	Exposure Control Mode Description	Exposure​Control​Mode​Description	LT	1
(0018,7064)	Exposure Status	Exposure​Status	CS	1
(0018,7065)	Phototimer Setting	Phototimer​Setting	DS	1
(0018,8150)	Exposure Time in µS	Exposure​Time​InuS	DS	1
(0018,8151)	X-Ray Tube Current in µA	X​Ray​Tube​Current​InuA	DS	1
(0018,9004)	Content Qualification	Content​Qualification	CS	1
(0018,9005)	Pulse Sequence Name	Pulse​Sequence​Name	SH	1
(0018,9006)	MR Imaging Modifier Sequence	MR​Imaging​Modifier​Sequence	SQ	1
(0018,9008)	Echo Pulse Sequence	Echo​Pulse​Sequence	CS	1
(0018,9009)	Inversion Recovery	Inversion​Recovery	CS	1
(0018,9010)	Flow Compensation	Flow​Compensation	CS	1
(0018,9011)	Multiple Spin Echo	Multiple​Spin​Echo	CS	1
(0018,9012)	Multi-planar Excitation	Multi​Planar​Excitation	CS	1
(0018,9014)	Phase Contrast	Phase​Contrast	CS	1
(0018,9015)	Time of Flight Contrast	Time​Of​Flight​Contrast	CS	1
(0018,9016)	Spoiling	Spoiling	CS	1
(0018,9017)	Steady State Pulse Sequence	Steady​State​Pulse​Sequence	CS	1
(0018,9018)	Echo Planar Pulse Sequence	Echo​Planar​Pulse​Sequence	CS	1
(0018,9019)	Tag Angle First Axis	Tag​Angle​First​Axis	FD	1
(0018,9020)	Magnetization Transfer	Magnetization​Transfer	CS	1
(0018,9021)	T2 Preparation	T2Preparation	CS	1
(0018,9022)	Blood Signal Nulling	Blood​Signal​Nulling	CS	1
(0018,9024)	Saturation Recovery	Saturation​Recovery	CS	1
(0018,9025)	Spectrally Selected Suppression	Spectrally​Selected​Suppression	CS	1
(0018,9026)	Spectrally Selected Excitation	Spectrally​Selected​Excitation	CS	1
(0018,9027)	Spatial Pre-saturation	Spatial​Presaturation	CS	1
(0018,9028)	Tagging	Tagging	CS	1
(0018,9029)	Oversampling Phase	Oversampling​Phase	CS	1
(0018,9030)	Tag Spacing First Dimension	Tag​Spacing​First​Dimension	FD	1
(0018,9032)	Geometry of k-Space Traversal	Geometry​Of​K​Space​Traversal	CS	1
(0018,9033)	Segmented k-Space Traversal	Segmented​K​Space​Traversal	CS	1
(0018,9034)	Rectilinear Phase Encode Reordering	Rectilinear​Phase​Encode​Reordering	CS	1
(0018,9035)	Tag Thickness	Tag​Thickness	FD	1
(0018,9036)	Partial Fourier Direction	Partial​Fourier​Direction	CS	1
(0018,9037)	Cardiac Synchronization Technique	Cardiac​Synchronization​Technique	CS	1
(0018,9041)	Receive Coil Manufacturer Name	Receive​Coil​Manufacturer​Name	LO	1
(0018,9042)	MR Receive Coil Sequence	MR​Receive​Coil​Sequence	SQ	1
(0018,9043)	Receive Coil Type	Receive​Coil​Type	CS	1
(0018,9044)	Quadrature Receive Coil	Quadrature​Receive​Coil	CS	1
(0018,9045)	Multi-Coil Definition Sequence	Multi​Coil​Definition​Sequence	SQ	1
(0018,9046)	Multi-Coil Configuration	Multi​Coil​Configuration	LO	1
(0018,9047)	Multi-Coil Element Name	Multi​Coil​Element​Name	SH	1
(0018,9048)	Multi-Coil Element Used	Multi​Coil​Element​Used	CS	1
(0018,9049)	MR Transmit Coil Sequence	MR​Transmit​Coil​Sequence	SQ	1
(0018,9050)	Transmit Coil Manufacturer Name	Transmit​Coil​Manufacturer​Name	LO	1
(0018,9051)	Transmit Coil Type	Transmit​Coil​Type	CS	1
(0018,9052)	Spectral Width	Spectral​Width	FD	1-2
(0018,9053)	Chemical Shift Reference	Chemical​Shift​Reference	FD	1-2
(0018,9054)	Volume Localization Technique	Volume​Localization​Technique	CS	1
(0018,9058)	MR Acquisition Frequency Encoding Steps	MR​Acquisition​Frequency​Encoding​Steps	US	1
(0018,9059)	De-coupling	Decoupling	CS	1
(0018,9060)	De-coupled Nucleus	Decoupled​Nucleus	CS	1-2
(0018,9061)	De-coupling Frequency	Decoupling​Frequency	FD	1-2
(0018,9062)	De-coupling Method	Decoupling​Method	CS	1
(0018,9063)	De-coupling Chemical Shift Reference	Decoupling​Chemical​Shift​Reference	FD	1-2
(0018,9064)	k-space Filtering	K​Space​Filtering	CS	1
(0018,9065)	Time Domain Filtering	Time​Domain​Filtering	CS	1-2
(0018,9066)	Number of Zero Fills	Number​Of​Zero​Fills	US	1-2
(0018,9067)	Baseline Correction	Baseline​Correction	CS	1
(0018,9069)	Parallel Reduction Factor In-plane	Parallel​Reduction​Factor​In​Plane	FD	1
(0018,9070)	Cardiac R-R Interval Specified	Cardiac​RR​Interval​Specified	FD	1
(0018,9073)	Acquisition Duration	Acquisition​Duration	FD	1
(0018,9074)	Frame Acquisition DateTime	Frame​Acquisition​Date​Time	DT	1
(0018,9075)	Diffusion Directionality	Diffusion​Directionality	CS	1
(0018,9076)	Diffusion Gradient Direction Sequence	Diffusion​Gradient​Direction​Sequence	SQ	1
(0018,9077)	Parallel Acquisition	Parallel​Acquisition	CS	1
(0018,9078)	Parallel Acquisition Technique	Parallel​Acquisition​Technique	CS	1
(0018,9079)	Inversion Times	Inversion​Times	FD	1-n
(0018,9080)	Metabolite Map Description	Metabolite​Map​Description	ST	1
(0018,9081)	Partial Fourier	Partial​Fourier	CS	1
(0018,9082)	Effective Echo Time	Effective​Echo​Time	FD	1
(0018,9083)	Metabolite Map Code Sequence	Metabolite​Map​Code​Sequence	SQ	1
(0018,9084)	Chemical Shift Sequence	Chemical​Shift​Sequence	SQ	1
(0018,9085)	Cardiac Signal Source	Cardiac​Signal​Source	CS	1
(0018,9087)	Diffusion b-value	Diffusion​B​Value	FD	1
(0018,9089)	Diffusion Gradient Orientation	Diffusion​Gradient​Orientation	FD	3
(0018,9090)	Velocity Encoding Direction	Velocity​Encoding​Direction	FD	3
(0018,9091)	Velocity Encoding Minimum Value	Velocity​Encoding​Minimum​Value	FD	1
(0018,9092)	Velocity Encoding Acquisition Sequence	Velocity​Encoding​Acquisition​Sequence	SQ	1
(0018,9093)	Number of k-Space Trajectories	Number​Of​K​Space​Trajectories	US	1
(0018,9094)	Coverage of k-Space	Coverage​Of​K​Space	CS	1
(0018,9095)	Spectroscopy Acquisition Phase Rows	Spectroscopy​Acquisition​Phase​Rows	UL	1
(0018,9096)	Parallel Reduction Factor In-plane (Retired)	Parallel​Reduction​Factor​In​Plane​Retired	FD	1	RET (2007)
(0018,9098)	Transmitter Frequency	Transmitter​Frequency	FD	1-2
(0018,9100)	Resonant Nucleus	Resonant​Nucleus	CS	1-2
(0018,9101)	Frequency Correction	Frequency​Correction	CS	1
(0018,9103)	MR Spectroscopy FOV/Geometry Sequence	MR​Spectroscopy​FOV​Geometry​Sequence	SQ	1
(0018,9104)	Slab Thickness	Slab​Thickness	FD	1
(0018,9105)	Slab Orientation	Slab​Orientation	FD	3
(0018,9106)	Mid Slab Position	Mid​Slab​Position	FD	3
(0018,9107)	MR Spatial Saturation Sequence	MR​Spatial​Saturation​Sequence	SQ	1
(0018,9112)	MR Timing and Related Parameters Sequence	MR​Timing​And​Related​Parameters​Sequence	SQ	1
(0018,9114)	MR Echo Sequence	MR​Echo​Sequence	SQ	1
(0018,9115)	MR Modifier Sequence	MR​Modifier​Sequence	SQ	1
(0018,9117)	MR Diffusion Sequence	MR​Diffusion​Sequence	SQ	1
(0018,9118)	Cardiac Synchronization Sequence	Cardiac​Synchronization​Sequence	SQ	1
(0018,9119)	MR Averages Sequence	MR​Averages​Sequence	SQ	1
(0018,9125)	MR FOV/Geometry Sequence	MRFOV​Geometry​Sequence	SQ	1
(0018,9126)	Volume Localization Sequence	Volume​Localization​Sequence	SQ	1
(0018,9127)	Spectroscopy Acquisition Data Columns	Spectroscopy​Acquisition​Data​Columns	UL	1
(0018,9147)	Diffusion Anisotropy Type	Diffusion​Anisotropy​Type	CS	1
(0018,9151)	Frame Reference DateTime	Frame​Reference​Date​Time	DT	1
(0018,9152)	MR Metabolite Map Sequence	MR​Metabolite​Map​Sequence	SQ	1
(0018,9155)	Parallel Reduction Factor out-of-plane	Parallel​Reduction​Factor​Out​Of​Plane	FD	1
(0018,9159)	Spectroscopy Acquisition Out-of-plane Phase Steps	Spectroscopy​Acquisition​Out​Of​Plane​Phase​Steps	UL	1
(0018,9166)	Bulk Motion Status	Bulk​Motion​Status	CS	1	RET (2007)
(0018,9168)	Parallel Reduction Factor Second In-plane	Parallel​Reduction​Factor​Second​In​Plane	FD	1
(0018,9169)	Cardiac Beat Rejection Technique	Cardiac​Beat​Rejection​Technique	CS	1
(0018,9170)	Respiratory Motion Compensation Technique	Respiratory​Motion​Compensation​Technique	CS	1
(0018,9171)	Respiratory Signal Source	Respiratory​Signal​Source	CS	1
(0018,9172)	Bulk Motion Compensation Technique	Bulk​Motion​Compensation​Technique	CS	1
(0018,9173)	Bulk Motion Signal Source	Bulk​Motion​Signal​Source	CS	1
(0018,9174)	Applicable Safety Standard Agency	Applicable​Safety​Standard​Agency	CS	1
(0018,9175)	Applicable Safety Standard Description	Applicable​Safety​Standard​Description	LO	1
(0018,9176)	Operating Mode Sequence	Operating​Mode​Sequence	SQ	1
(0018,9177)	Operating Mode Type	Operating​Mode​Type	CS	1
(0018,9178)	Operating Mode	Operating​Mode	CS	1
(0018,9179)	Specific Absorption Rate Definition	Specific​Absorption​Rate​Definition	CS	1
(0018,9180)	Gradient Output Type	Gradient​Output​Type	CS	1
(0018,9181)	Specific Absorption Rate Value	Specific​Absorption​Rate​Value	FD	1
(0018,9182)	Gradient Output	Gradient​Output	FD	1
(0018,9183)	Flow Compensation Direction	Flow​Compensation​Direction	CS	1
(0018,9184)	Tagging Delay	Tagging​Delay	FD	1
(0018,9185)	Respiratory Motion Compensation Technique Description	Respiratory​Motion​Compensation​Technique​Description	ST	1
(0018,9186)	Respiratory Signal Source ID	Respiratory​Signal​Source​ID	SH	1
(0018,9195)	Chemical Shift Minimum Integration Limit in Hz	Chemical​Shift​Minimum​Integration​Limit​In​Hz	FD	1	RET (2003)
(0018,9196)	Chemical Shift Maximum Integration Limit in Hz	Chemical​Shift​Maximum​Integration​Limit​In​Hz	FD	1	RET (2003)
(0018,9197)	MR Velocity Encoding Sequence	MR​Velocity​Encoding​Sequence	SQ	1
(0018,9198)	First Order Phase Correction	First​Order​Phase​Correction	CS	1
(0018,9199)	Water Referenced Phase Correction	Water​Referenced​Phase​Correction	CS	1
(0018,9200)	MR Spectroscopy Acquisition Type	MR​Spectroscopy​Acquisition​Type	CS	1
(0018,9214)	Respiratory Cycle Position	Respiratory​Cycle​Position	CS	1
(0018,9217)	Velocity Encoding Maximum Value	Velocity​Encoding​Maximum​Value	FD	1
(0018,9218)	Tag Spacing Second Dimension	Tag​Spacing​Second​Dimension	FD	1
(0018,9219)	Tag Angle Second Axis	Tag​Angle​Second​Axis	SS	1
(0018,9220)	Frame Acquisition Duration	Frame​Acquisition​Duration	FD	1
(0018,9226)	MR Image Frame Type Sequence	MR​Image​Frame​Type​Sequence	SQ	1
(0018,9227)	MR Spectroscopy Frame Type Sequence	MR​Spectroscopy​Frame​Type​Sequence	SQ	1
(0018,9231)	MR Acquisition Phase Encoding Steps in-plane	MR​Acquisition​Phase​Encoding​Steps​In​Plane	US	1
(0018,9232)	MR Acquisition Phase Encoding Steps out-of-plane	MR​Acquisition​Phase​Encoding​Steps​Out​Of​Plane	US	1
(0018,9234)	Spectroscopy Acquisition Phase Columns	Spectroscopy​Acquisition​Phase​Columns	UL	1
(0018,9236)	Cardiac Cycle Position	Cardiac​Cycle​Position	CS	1
(0018,9239)	Specific Absorption Rate Sequence	Specific​Absorption​Rate​Sequence	SQ	1
(0018,9240)	RF Echo Train Length	RF​Echo​Train​Length	US	1
(0018,9241)	Gradient Echo Train Length	Gradient​Echo​Train​Length	US	1
(0018,9250)	Arterial Spin Labeling Contrast	Arterial​Spin​Labeling​Contrast	CS	1
(0018,9251)	MR Arterial Spin Labeling Sequence	MR​Arterial​Spin​Labeling​Sequence	SQ	1
(0018,9252)	ASL Technique Description	ASL​Technique​Description	LO	1
(0018,9253)	ASL Slab Number	ASL​Slab​Number	US	1
(0018,9254)	ASL Slab Thickness	ASL​Slab​Thickness	FD	1
(0018,9255)	ASL Slab Orientation	ASL​Slab​Orientation	FD	3
(0018,9256)	ASL Mid Slab Position	ASL​Mid​Slab​Position	FD	3
(0018,9257)	ASL Context	ASL​Context	CS	1
(0018,9258)	ASL Pulse Train Duration	ASL​Pulse​Train​Duration	UL	1
(0018,9259)	ASL Crusher Flag	ASL​Crusher​Flag	CS	1
(0018,925A)	ASL Crusher Flow Limit	ASL​Crusher​Flow​Limit	FD	1
(0018,925B)	ASL Crusher Description	ASL​Crusher​Description	LO	1
(0018,925C)	ASL Bolus Cut-off Flag	ASL​Bolus​Cutoff​Flag	CS	1
(0018,925D)	ASL Bolus Cut-off Timing Sequence	ASL​Bolus​Cutoff​Timing​Sequence	SQ	1
(0018,925E)	ASL Bolus Cut-off Technique	ASL​Bolus​Cutoff​Technique	LO	1
(0018,925F)	ASL Bolus Cut-off Delay Time	ASL​Bolus​Cutoff​Delay​Time	UL	1
(0018,9260)	ASL Slab Sequence	ASL​Slab​Sequence	SQ	1
(0018,9295)	Chemical Shift Minimum Integration Limit in ppm	Chemical​Shift​Minimum​Integration​Limit​Inppm	FD	1
(0018,9296)	Chemical Shift Maximum Integration Limit in ppm	Chemical​Shift​Maximum​Integration​Limit​Inppm	FD	1
(0018,9297)	Water Reference Acquisition	Water​Reference​Acquisition	CS	1
(0018,9298)	Echo Peak Position	Echo​Peak​Position	IS	1
(0018,9301)	CT Acquisition Type Sequence	CT​Acquisition​Type​Sequence	SQ	1
(0018,9302)	Acquisition Type	Acquisition​Type	CS	1
(0018,9303)	Tube Angle	Tube​Angle	FD	1
(0018,9304)	CT Acquisition Details Sequence	CT​Acquisition​Details​Sequence	SQ	1
(0018,9305)	Revolution Time	Revolution​Time	FD	1
(0018,9306)	Single Collimation Width	Single​Collimation​Width	FD	1
(0018,9307)	Total Collimation Width	Total​Collimation​Width	FD	1
(0018,9308)	CT Table Dynamics Sequence	CT​Table​Dynamics​Sequence	SQ	1
(0018,9309)	Table Speed	Table​Speed	FD	1
(0018,9310)	Table Feed per Rotation	Table​Feed​Per​Rotation	FD	1
(0018,9311)	Spiral Pitch Factor	Spiral​Pitch​Factor	FD	1
(0018,9312)	CT Geometry Sequence	CT​Geometry​Sequence	SQ	1
(0018,9313)	Data Collection Center (Patient)	Data​Collection​Center​Patient	FD	3
(0018,9314)	CT Reconstruction Sequence	CT​Reconstruction​Sequence	SQ	1
(0018,9315)	Reconstruction Algorithm	Reconstruction​Algorithm	CS	1
(0018,9316)	Convolution Kernel Group	Convolution​Kernel​Group	CS	1
(0018,9317)	Reconstruction Field of View	Reconstruction​Field​Of​View	FD	2
(0018,9318)	Reconstruction Target Center (Patient)	Reconstruction​Target​Center​Patient	FD	3
(0018,9319)	Reconstruction Angle	Reconstruction​Angle	FD	1
(0018,9320)	Image Filter	Image​Filter	SH	1
(0018,9321)	CT Exposure Sequence	CT​Exposure​Sequence	SQ	1
(0018,9322)	Reconstruction Pixel Spacing	Reconstruction​Pixel​Spacing	FD	2
(0018,9323)	Exposure Modulation Type	Exposure​Modulation​Type	CS	1-n
(0018,9324)	Estimated Dose Saving	Estimated​Dose​Saving	FD	1
(0018,9325)	CT X-Ray Details Sequence	CT​X​Ray​Details​Sequence	SQ	1
(0018,9326)	CT Position Sequence	CT​Position​Sequence	SQ	1
(0018,9327)	Table Position	Table​Position	FD	1
(0018,9328)	Exposure Time in ms	Exposure​Time​Inms	FD	1
(0018,9329)	CT Image Frame Type Sequence	CT​Image​Frame​Type​Sequence	SQ	1
(0018,9330)	X-Ray Tube Current in mA	X​Ray​Tube​Current​InmA	FD	1
(0018,9332)	Exposure in mAs	Exposure​Inm​As	FD	1
(0018,9333)	Constant Volume Flag	Constant​Volume​Flag	CS	1
(0018,9334)	Fluoroscopy Flag	Fluoroscopy​Flag	CS	1
(0018,9335)	Distance Source to Data Collection Center	Distance​Source​To​Data​Collection​Center	FD	1
(0018,9337)	Contrast/Bolus Agent Number	Contrast​Bolus​Agent​Number	US	1
(0018,9338)	Contrast/Bolus Ingredient Code Sequence	Contrast​Bolus​Ingredient​Code​Sequence	SQ	1
(0018,9340)	Contrast Administration Profile Sequence	Contrast​Administration​Profile​Sequence	SQ	1
(0018,9341)	Contrast/Bolus Usage Sequence	Contrast​Bolus​Usage​Sequence	SQ	1
(0018,9342)	Contrast/Bolus Agent Administered	Contrast​Bolus​Agent​Administered	CS	1
(0018,9343)	Contrast/Bolus Agent Detected	Contrast​Bolus​Agent​Detected	CS	1
(0018,9344)	Contrast/Bolus Agent Phase	Contrast​Bolus​Agent​Phase	CS	1
(0018,9345)	CTDIvol	CTD​Ivol	FD	1
(0018,9346)	CTDI Phantom Type Code Sequence	CTDI​Phantom​Type​Code​Sequence	SQ	1
(0018,9351)	Calcium Scoring Mass Factor Patient	Calcium​Scoring​Mass​Factor​Patient	FL	1
(0018,9352)	Calcium Scoring Mass Factor Device	Calcium​Scoring​Mass​Factor​Device	FL	3
(0018,9353)	Energy Weighting Factor	Energy​Weighting​Factor	FL	1
(0018,9360)	CT Additional X-Ray Source Sequence	CT​Additional​X​Ray​Source​Sequence	SQ	1
(0018,9361)	Multi-energy CT Acquisition	Multienergy​CT​Acquisition	CS	1
(0018,9362)	Multi-energy CT Acquisition Sequence	Multienergy​CT​Acquisition​Sequence	SQ	1
(0018,9363)	Multi-energy CT Processing Sequence	Multienergy​CT​Processing​Sequence	SQ	1
(0018,9364)	Multi-energy CT Characteristics Sequence	Multienergy​CT​Characteristics​Sequence	SQ	1
(0018,9365)	Multi-energy CT X-Ray Source Sequence	Multienergy​CT​XRay​Source​Sequence	SQ	1
(0018,9366)	X-Ray Source Index	XRay​Source​Index	US	1
(0018,9367)	X-Ray Source ID	XRay​Source​ID	UC	1
(0018,9368)	Multi-energy Source Technique	Multienergy​Source​Technique	CS	1
(0018,9369)	Source Start Date Time	Source​Start​Date​Time	DT	1
(0018,936A)	Source End Date Time	Source​End​Date​Time	DT	1
(0018,936B)	Switching Phase Number	Switching​Phase​Number	US	1
(0018,936C)	Switching Phase Nominal Duration	Switching​Phase​Nominal​Duration	DS	1
(0018,936D)	Switching Phase Transition Duration	Switching​Phase​Transition​Duration	DS	1
(0018,936E)	Effective Bin Energy	Effective​Bin​Energy	DS	1
(0018,936F)	Multi-energy CT X-Ray Detector Sequence	Multienergy​CT​XRay​Detector​Sequence	SQ	1
(0018,9370)	X-Ray Detector Index	XRay​Detector​Index	US	1
(0018,9371)	X-Ray Detector ID	XRay​Detector​ID	UC	1
(0018,9372)	Multi-energy Detector Type	Multienergy​Detector​Type	CS	1
(0018,9373)	X-Ray Detector Label	XRay​Detector​Label	ST	1
(0018,9374)	Nominal Max Energy	Nominal​Max​Energy	DS	1
(0018,9375)	Nominal Min Energy	Nominal​Min​Energy	DS	1
(0018,9376)	Referenced X-Ray Detector Index	Referenced​XRay​Detector​Index	US	1-n
(0018,9377)	Referenced X-Ray Source Index	Referenced​XRay​Source​Index	US	1-n
(0018,9378)	Referenced Path Index	Referenced​Path​Index	US	1-n
(0018,9379)	Multi-energy CT Path Sequence	Multienergy​CT​Path​Sequence	SQ	1
(0018,937A)	Multi-energy CT Path Index	Multienergy​CT​Path​Index	US	1
(0018,937B)	Multi-energy Acquisition Description	Multienergy​Acquisition​Description	UT	1
(0018,937C)	Monoenergetic Energy Equivalent	Monoenergetic​Energy​Equivalent	FD	1
(0018,937D)	Material Code Sequence	Material​Code​Sequence	SQ	1
(0018,937E)	Decomposition Method	Decomposition​Method	CS	1
(0018,937F)	Decomposition Description	Decomposition​Description	UT	1
(0018,9380)	Decomposition Algorithm Identification Sequence	Decomposition​Algorithm​Identification​Sequence	SQ	1
(0018,9381)	Decomposition Material Sequence	Decomposition​Material​Sequence	SQ	1
(0018,9382)	Material Attenuation Sequence	Material​Attenuation​Sequence	SQ	1
(0018,9383)	Photon Energy	Photon​Energy	DS	1
(0018,9384)	X-Ray Mass Attenuation Coefficient	XRay​Mass​Attenuation​Coefficient	DS	1
(0018,9401)	Projection Pixel Calibration Sequence	Projection​Pixel​Calibration​Sequence	SQ	1
(0018,9402)	Distance Source to Isocenter	Distance​Source​To​Isocenter	FL	1
(0018,9403)	Distance Object to Table Top	Distance​Object​To​Table​Top	FL	1
(0018,9404)	Object Pixel Spacing in Center of Beam	Object​Pixel​Spacing​In​Center​Of​Beam	FL	2
(0018,9405)	Positioner Position Sequence	Positioner​Position​Sequence	SQ	1
(0018,9406)	Table Position Sequence	Table​Position​Sequence	SQ	1
(0018,9407)	Collimator Shape Sequence	Collimator​Shape​Sequence	SQ	1
(0018,9410)	Planes in Acquisition	Planes​In​Acquisition	CS	1
(0018,9412)	XA/XRF Frame Characteristics Sequence	XAXRF​Frame​Characteristics​Sequence	SQ	1
(0018,9417)	Frame Acquisition Sequence	Frame​Acquisition​Sequence	SQ	1
(0018,9420)	X-Ray Receptor Type	X​Ray​Receptor​Type	CS	1
(0018,9423)	Acquisition Protocol Name	Acquisition​Protocol​Name	LO	1
(0018,9424)	Acquisition Protocol Description	Acquisition​Protocol​Description	LT	1
(0018,9425)	Contrast/Bolus Ingredient Opaque	Contrast​Bolus​Ingredient​Opaque	CS	1
(0018,9426)	Distance Receptor Plane to Detector Housing	Distance​Receptor​Plane​To​Detector​Housing	FL	1
(0018,9427)	Intensifier Active Shape	Intensifier​Active​Shape	CS	1
(0018,9428)	Intensifier Active Dimension(s)	Intensifier​Active​Dimensions	FL	1-2
(0018,9429)	Physical Detector Size	Physical​Detector​Size	FL	2
(0018,9430)	Position of Isocenter Projection	Position​Of​Isocenter​Projection	FL	2
(0018,9432)	Field of View Sequence	Field​Of​View​Sequence	SQ	1
(0018,9433)	Field of View Description	Field​Of​View​Description	LO	1
(0018,9434)	Exposure Control Sensing Regions Sequence	Exposure​Control​Sensing​Regions​Sequence	SQ	1
(0018,9435)	Exposure Control Sensing Region Shape	Exposure​Control​Sensing​Region​Shape	CS	1
(0018,9436)	Exposure Control Sensing Region Left Vertical Edge	Exposure​Control​Sensing​Region​Left​Vertical​Edge	SS	1
(0018,9437)	Exposure Control Sensing Region Right Vertical Edge	Exposure​Control​Sensing​Region​Right​Vertical​Edge	SS	1
(0018,9438)	Exposure Control Sensing Region Upper Horizontal Edge	Exposure​Control​Sensing​Region​Upper​Horizontal​Edge	SS	1
(0018,9439)	Exposure Control Sensing Region Lower Horizontal Edge	Exposure​Control​Sensing​Region​Lower​Horizontal​Edge	SS	1
(0018,9440)	Center of Circular Exposure Control Sensing Region	Center​Of​Circular​Exposure​Control​Sensing​Region	SS	2
(0018,9441)	Radius of Circular Exposure Control Sensing Region	Radius​Of​Circular​Exposure​Control​Sensing​Region	US	1
(0018,9442)	Vertices of the Polygonal Exposure Control Sensing Region	Vertices​Of​The​Polygonal​Exposure​Control​Sensing​Region	SS	2-n
(0018,9445)					RET (2004) - See Note 3
(0018,9447)	Column Angulation (Patient)	Column​Angulation​Patient	FL	1
(0018,9449)	Beam Angle	Beam​Angle	FL	1
(0018,9451)	Frame Detector Parameters Sequence	Frame​Detector​Parameters​Sequence	SQ	1
(0018,9452)	Calculated Anatomy Thickness	Calculated​Anatomy​Thickness	FL	1
(0018,9455)	Calibration Sequence	Calibration​Sequence	SQ	1
(0018,9456)	Object Thickness Sequence	Object​Thickness​Sequence	SQ	1
(0018,9457)	Plane Identification	Plane​Identification	CS	1
(0018,9461)	Field of View Dimension(s) in Float	Field​Of​View​Dimensions​In​Float	FL	1-2
(0018,9462)	Isocenter Reference System Sequence	Isocenter​Reference​System​Sequence	SQ	1
(0018,9463)	Positioner Isocenter Primary Angle	Positioner​Isocenter​Primary​Angle	FL	1
(0018,9464)	Positioner Isocenter Secondary Angle	Positioner​Isocenter​Secondary​Angle	FL	1
(0018,9465)	Positioner Isocenter Detector Rotation Angle	Positioner​Isocenter​Detector​Rotation​Angle	FL	1
(0018,9466)	Table X Position to Isocenter	Table​X​Position​To​Isocenter	FL	1
(0018,9467)	Table Y Position to Isocenter	Table​Y​Position​To​Isocenter	FL	1
(0018,9468)	Table Z Position to Isocenter	Table​Z​Position​To​Isocenter	FL	1
(0018,9469)	Table Horizontal Rotation Angle	Table​Horizontal​Rotation​Angle	FL	1
(0018,9470)	Table Head Tilt Angle	Table​Head​Tilt​Angle	FL	1
(0018,9471)	Table Cradle Tilt Angle	Table​Cradle​Tilt​Angle	FL	1
(0018,9472)	Frame Display Shutter Sequence	Frame​Display​Shutter​Sequence	SQ	1
(0018,9473)	Acquired Image Area Dose Product	Acquired​Image​Area​Dose​Product	FL	1
(0018,9474)	C-arm Positioner Tabletop Relationship	C​Arm​Positioner​Tabletop​Relationship	CS	1
(0018,9476)	X-Ray Geometry Sequence	X​Ray​Geometry​Sequence	SQ	1
(0018,9477)	Irradiation Event Identification Sequence	Irradiation​Event​Identification​Sequence	SQ	1
(0018,9504)	X-Ray 3D Frame Type Sequence	X​Ray3D​Frame​Type​Sequence	SQ	1
(0018,9506)	Contributing Sources Sequence	Contributing​Sources​Sequence	SQ	1
(0018,9507)	X-Ray 3D Acquisition Sequence	X​Ray3D​Acquisition​Sequence	SQ	1
(0018,9508)	Primary Positioner Scan Arc	Primary​Positioner​Scan​Arc	FL	1
(0018,9509)	Secondary Positioner Scan Arc	Secondary​Positioner​Scan​Arc	FL	1
(0018,9510)	Primary Positioner Scan Start Angle	Primary​Positioner​Scan​Start​Angle	FL	1
(0018,9511)	Secondary Positioner Scan Start Angle	Secondary​Positioner​Scan​Start​Angle	FL	1
(0018,9514)	Primary Positioner Increment	Primary​Positioner​Increment	FL	1
(0018,9515)	Secondary Positioner Increment	Secondary​Positioner​Increment	FL	1
(0018,9516)	Start Acquisition DateTime	Start​Acquisition​Date​Time	DT	1
(0018,9517)	End Acquisition DateTime	End​Acquisition​Date​Time	DT	1
(0018,9518)	Primary Positioner Increment Sign	Primary​Positioner​Increment​Sign	SS	1
(0018,9519)	Secondary Positioner Increment Sign	Secondary​Positioner​Increment​Sign	SS	1
(0018,9524)	Application Name	Application​Name	LO	1
(0018,9525)	Application Version	Application​Version	LO	1
(0018,9526)	Application Manufacturer	Application​Manufacturer	LO	1
(0018,9527)	Algorithm Type	Algorithm​Type	CS	1
(0018,9528)	Algorithm Description	Algorithm​Description	LO	1
(0018,9530)	X-Ray 3D Reconstruction Sequence	X​Ray3D​Reconstruction​Sequence	SQ	1
(0018,9531)	Reconstruction Description	Reconstruction​Description	LO	1
(0018,9538)	Per Projection Acquisition Sequence	Per​Projection​Acquisition​Sequence	SQ	1
(0018,9541)	Detector Position Sequence	Detector​Position​Sequence	SQ	1
(0018,9542)	X-Ray Acquisition Dose Sequence	XRay​Acquisition​Dose​Sequence	SQ	1
(0018,9543)	X-Ray Source Isocenter Primary Angle	XRay​Source​Isocenter​Primary​Angle	FD	1
(0018,9544)	X-Ray Source Isocenter Secondary Angle	XRay​Source​Isocenter​Secondary​Angle	FD	1
(0018,9545)	Breast Support Isocenter Primary Angle	Breast​Support​Isocenter​Primary​Angle	FD	1
(0018,9546)	Breast Support Isocenter Secondary Angle	Breast​Support​Isocenter​Secondary​Angle	FD	1
(0018,9547)	Breast Support X Position to Isocenter	Breast​Support​XPosition​To​Isocenter	FD	1
(0018,9548)	Breast Support Y Position to Isocenter	Breast​Support​YPosition​To​Isocenter	FD	1
(0018,9549)	Breast Support Z Position to Isocenter	Breast​Support​ZPosition​To​Isocenter	FD	1
(0018,9550)	Detector Isocenter Primary Angle	Detector​Isocenter​Primary​Angle	FD	1
(0018,9551)	Detector Isocenter Secondary Angle	Detector​Isocenter​Secondary​Angle	FD	1
(0018,9552)	Detector X Position to Isocenter	Detector​XPosition​To​Isocenter	FD	1
(0018,9553)	Detector Y Position to Isocenter	Detector​YPosition​To​Isocenter	FD	1
(0018,9554)	Detector Z Position to Isocenter	Detector​ZPosition​To​Isocenter	FD	1
(0018,9555)	X-Ray Grid Sequence	XRay​Grid​Sequence	SQ	1
(0018,9556)	X-Ray Filter Sequence	XRay​Filter​Sequence	SQ	1
(0018,9557)	Detector Active Area TLHC Position	Detector​Active​Area​TLHCPosition	FD	3
(0018,9558)	Detector Active Area Orientation	Detector​Active​Area​Orientation	FD	6
(0018,9559)	Positioner Primary Angle Direction	Positioner​Primary​Angle​Direction	CS	1
(0018,9601)	Diffusion b-matrix Sequence	Diffusion​B​Matrix​Sequence	SQ	1
(0018,9602)	Diffusion b-value XX	Diffusion​B​Value​XX	FD	1
(0018,9603)	Diffusion b-value XY	Diffusion​B​Value​XY	FD	1
(0018,9604)	Diffusion b-value XZ	Diffusion​B​Value​XZ	FD	1
(0018,9605)	Diffusion b-value YY	Diffusion​B​Value​YY	FD	1
(0018,9606)	Diffusion b-value YZ	Diffusion​B​Value​YZ	FD	1
(0018,9607)	Diffusion b-value ZZ	Diffusion​B​Value​ZZ	FD	1
(0018,9621)	Functional MR Sequence	Functional​MR​Sequence	SQ	1
(0018,9622)	Functional Settling Phase Frames Present	Functional​Settling​Phase​Frames​Present	CS	1
(0018,9623)	Functional Sync Pulse	Functional​Sync​Pulse	DT	1
(0018,9624)	Settling Phase Frame	Settling​Phase​Frame	CS	1
(0018,9701)	Decay Correction DateTime	Decay​Correction​Date​Time	DT	1
(0018,9715)	Start Density Threshold	Start​Density​Threshold	FD	1
(0018,9716)	Start Relative Density Difference Threshold	Start​Relative​Density​Difference​Threshold	FD	1
(0018,9717)	Start Cardiac Trigger Count Threshold	Start​Cardiac​Trigger​Count​Threshold	FD	1
(0018,9718)	Start Respiratory Trigger Count Threshold	Start​Respiratory​Trigger​Count​Threshold	FD	1
(0018,9719)	Termination Counts Threshold	Termination​Counts​Threshold	FD	1
(0018,9720)	Termination Density Threshold	Termination​Density​Threshold	FD	1
(0018,9721)	Termination Relative Density Threshold	Termination​Relative​Density​Threshold	FD	1
(0018,9722)	Termination Time Threshold	Termination​Time​Threshold	FD	1
(0018,9723)	Termination Cardiac Trigger Count Threshold	Termination​Cardiac​Trigger​Count​Threshold	FD	1
(0018,9724)	Termination Respiratory Trigger Count Threshold	Termination​Respiratory​Trigger​Count​Threshold	FD	1
(0018,9725)	Detector Geometry	Detector​Geometry	CS	1
(0018,9726)	Transverse Detector Separation	Transverse​Detector​Separation	FD	1
(0018,9727)	Axial Detector Dimension	Axial​Detector​Dimension	FD	1
(0018,9729)	Radiopharmaceutical Agent Number	Radio​pharmaceutical​Agent​Number	US	1
(0018,9732)	PET Frame Acquisition Sequence	PET​Frame​Acquisition​Sequence	SQ	1
(0018,9733)	PET Detector Motion Details Sequence	PET​Detector​Motion​Details​Sequence	SQ	1
(0018,9734)	PET Table Dynamics Sequence	PET​Table​Dynamics​Sequence	SQ	1
(0018,9735)	PET Position Sequence	PET​Position​Sequence	SQ	1
(0018,9736)	PET Frame Correction Factors Sequence	PET​Frame​Correction​Factors​Sequence	SQ	1
(0018,9737)	Radiopharmaceutical Usage Sequence	Radio​pharmaceutical​Usage​Sequence	SQ	1
(0018,9738)	Attenuation Correction Source	Attenuation​Correction​Source	CS	1
(0018,9739)	Number of Iterations	Number​Of​Iterations	US	1
(0018,9740)	Number of Subsets	Number​Of​Subsets	US	1
(0018,9749)	PET Reconstruction Sequence	PET​Reconstruction​Sequence	SQ	1
(0018,9751)	PET Frame Type Sequence	PET​Frame​Type​Sequence	SQ	1
(0018,9755)	Time of Flight Information Used	Time​Of​Flight​Information​Used	CS	1
(0018,9756)	Reconstruction Type	Reconstruction​Type	CS	1
(0018,9758)	Decay Corrected	Decay​Corrected	CS	1
(0018,9759)	Attenuation Corrected	Attenuation​Corrected	CS	1
(0018,9760)	Scatter Corrected	Scatter​Corrected	CS	1
(0018,9761)	Dead Time Corrected	Dead​Time​Corrected	CS	1
(0018,9762)	Gantry Motion Corrected	Gantry​Motion​Corrected	CS	1
(0018,9763)	Patient Motion Corrected	Patient​Motion​Corrected	CS	1
(0018,9764)	Count Loss Normalization Corrected	Count​Loss​Normalization​Corrected	CS	1
(0018,9765)	Randoms Corrected	Randoms​Corrected	CS	1
(0018,9766)	Non-uniform Radial Sampling Corrected	Non​Uniform​Radial​Sampling​Corrected	CS	1
(0018,9767)	Sensitivity Calibrated	Sensitivity​Calibrated	CS	1
(0018,9768)	Detector Normalization Correction	Detector​Normalization​Correction	CS	1
(0018,9769)	Iterative Reconstruction Method	Iterative​Reconstruction​Method	CS	1
(0018,9770)	Attenuation Correction Temporal Relationship	Attenuation​Correction​Temporal​Relationship	CS	1
(0018,9771)	Patient Physiological State Sequence	Patient​Physiological​State​Sequence	SQ	1
(0018,9772)	Patient Physiological State Code Sequence	Patient​Physiological​State​Code​Sequence	SQ	1
(0018,9801)	Depth(s) of Focus	Depths​Of​Focus	FD	1-n
(0018,9803)	Excluded Intervals Sequence	Excluded​Intervals​Sequence	SQ	1
(0018,9804)	Exclusion Start DateTime	Exclusion​Start​Date​Time	DT	1
(0018,9805)	Exclusion Duration	Exclusion​Duration	FD	1
(0018,9806)	US Image Description Sequence	US​Image​Description​Sequence	SQ	1
(0018,9807)	Image Data Type Sequence	Image​Data​Type​Sequence	SQ	1
(0018,9808)	Data Type	Data​Type	CS	1
(0018,9809)	Transducer Scan Pattern Code Sequence	Transducer​Scan​Pattern​Code​Sequence	SQ	1
(0018,980B)	Aliased Data Type	Aliased​Data​Type	CS	1
(0018,980C)	Position Measuring Device Used	Position​Measuring​Device​Used	CS	1
(0018,980D)	Transducer Geometry Code Sequence	Transducer​Geometry​Code​Sequence	SQ	1
(0018,980E)	Transducer Beam Steering Code Sequence	Transducer​Beam​Steering​Code​Sequence	SQ	1
(0018,980F)	Transducer Application Code Sequence	Transducer​Application​Code​Sequence	SQ	1
(0018,9810)	Zero Velocity Pixel Value	Zero​Velocity​Pixel​Value	US or SS	1
(0018,9900)	Reference Location Label	Reference​Location​L​abel	LO	1
(0018,9901)	Reference Location Description	Reference​Location​Description	UT	1
(0018,9902)	Reference Basis Code Sequence	Reference​Basis​Code​Sequence	SQ	1
(0018,9903)	Reference Geometry Code Sequence	Reference​Geometry​Code​Sequence	SQ	1
(0018,9904)	Offset Distance	Offset​Distance	DS	1
(0018,9905)	Offset Direction	Offset​Direction	CS	1
(0018,9906)	Potential Scheduled Protocol Code Sequence	Potential​Scheduled​Protocol​Code​Sequence	SQ	1
(0018,9907)	Potential Requested Procedure Code Sequence	Potential​Requested​Procedure​Code​Sequence	SQ	1
(0018,9908)	Potential Reasons for Procedure	Potential​Reasons​For​Procedure	UC	1-n
(0018,9909)	Potential Reasons for Procedure Code Sequence	Potential​Reasons​For​Procedure​Code​Sequence	SQ	1
(0018,990A)	Potential Diagnostic Tasks	Potential​Diagnostic​Tasks	UC	1-n
(0018,990B)	Contraindications Code Sequence	Contraindications​Code​Sequence	SQ	1
(0018,990C)	Referenced Defined Protocol Sequence	Referenced​Defined​Protocol​Sequence	SQ	1
(0018,990D)	Referenced Performed Protocol Sequence	Referenced​Performed​Protocol​Sequence	SQ	1
(0018,990E)	Predecessor Protocol Sequence	Predecessor​Protocol​Sequence	SQ	1
(0018,990F)	Protocol Planning Information	Protocol​Planning​Information	UT	1
(0018,9910)	Protocol Design Rationale	Protocol​Design​Rationale	UT	1
(0018,9911)	Patient Specification Sequence	Patient​Specification​Sequence	SQ	1
(0018,9912)	Model Specification Sequence	Model​Specification​Sequence	SQ	1
(0018,9913)	Parameters Specification Sequence	Parameters​Specification​Sequence	SQ	1
(0018,9914)	Instruction Sequence	Instruction​Sequence	SQ	1
(0018,9915)	Instruction Index	Instruction​Index	US	1
(0018,9916)	Instruction Text	Instruction​Text	LO	1
(0018,9917)	Instruction Description	Instruction​Description	UT	1
(0018,9918)	Instruction Performed Flag	Instruction​Performed​Flag	CS	1
(0018,9919)	Instruction Performed DateTime	Instruction​Performed​Date​Time	DT	1
(0018,991A)	Instruction Performance Comment	Instruction​Performance​Comment	UT	1
(0018,991B)	Patient Positioning Instruction Sequence	Patient​Positioning​Instruction​Sequence	SQ	1
(0018,991C)	Positioning Method Code Sequence	Positioning​Method​Code​Sequence	SQ	1
(0018,991D)	Positioning Landmark Sequence	Positioning​Landmark​Sequence	SQ	1
(0018,991E)	Target Frame of Reference UID	Target​Frame​Of​Reference​UID	UI	1
(0018,991F)	Acquisition Protocol Element Specification Sequence	Acquisition​Protocol​Element​Specification​Sequence	SQ	1
(0018,9920)	Acquisition Protocol Element Sequence	Acquisition​Protocol​Element​Sequence	SQ	1
(0018,9921)	Protocol Element Number	Protocol​Element​Number	US	1
(0018,9922)	Protocol Element Name	Protocol​Element​Name	LO	1
(0018,9923)	Protocol Element Characteristics Summary	Protocol​Element​Characteristics​Summary	UT	1
(0018,9924)	Protocol Element Purpose	Protocol​Element​Purpose	UT	1
(0018,9930)	Acquisition Motion	Acquisition​Motion	CS	1
(0018,9931)	Acquisition Start Location Sequence	Acquisition​Start​Location​Sequence	SQ	1
(0018,9932)	Acquisition End Location Sequence	Acquisition​End​Location​Sequence	SQ	1
(0018,9933)	Reconstruction Protocol Element Specification Sequence	Reconstruction​Protocol​Element​Specification​Sequence	SQ	1
(0018,9934)	Reconstruction Protocol Element Sequence	Reconstruction​Protocol​Element​Sequence	SQ	1
(0018,9935)	Storage Protocol Element Specification Sequence	Storage​Protocol​Element​Specification​Sequence	SQ	1
(0018,9936)	Storage Protocol Element Sequence	Storage​Protocol​Element​Sequence	SQ	1
(0018,9937)	Requested Series Description	Requested​Series​Description	LO	1
(0018,9938)	Source Acquisition Protocol Element Number	Source​Acquisition​Protocol​Element​Number	US	1-n
(0018,9939)	Source Acquisition Beam Number	Source​Acquisition​Beam​Number	US	1-n
(0018,993A)	Source Reconstruction Protocol Element Number	Source​Reconstruction​Protocol​Element​Number	US	1-n
(0018,993B)	Reconstruction Start Location Sequence	Reconstruction​Start​Location​Sequence	SQ	1
(0018,993C)	Reconstruction End Location Sequence	Reconstruction​End​Location​Sequence	SQ	1
(0018,993D)	Reconstruction Algorithm Sequence	Reconstruction​Algorithm​Sequence	SQ	1
(0018,993E)	Reconstruction Target Center Location Sequence	Reconstruction​Target​Center​Location​Sequence	SQ	1
(0018,9941)	Image Filter Description	Image​Filter​Description	UT	1
(0018,9942)	CTDIvol Notification Trigger	CTDIvol​Notification​Trigger	FD	1
(0018,9943)	DLP Notification Trigger	DLP​Notification​Trigger	FD	1
(0018,9944)	Auto KVP Selection Type	Auto​KVP​Selection​Type	CS	1
(0018,9945)	Auto KVP Upper Bound	Auto​KVP​Upper​Bound	FD	1
(0018,9946)	Auto KVP Lower Bound	Auto​KVP​Lower​Bound	FD	1
(0018,9947)	Protocol Defined Patient Position	Protocol​Defined​Patient​Position	CS	1
(0018,A001)	Contributing Equipment Sequence	Contributing​Equipment​Sequence	SQ	1
(0018,A002)	Contribution DateTime	Contribution​Date​Time	DT	1
(0018,A003)	Contribution Description	Contribution​Description	ST	1
(0020,000D)	Study Instance UID	Study​Instance​UID	UI	1
(0020,000E)	Series Instance UID	Series​Instance​UID	UI	1
(0020,0010)	Study ID	Study​ID	SH	1
(0020,0011)	Series Number	Series​Number	IS	1
(0020,0012)	Acquisition Number	Acquisition​Number	IS	1
(0020,0013)	Instance Number	Instance​Number	IS	1
(0020,0014)	Isotope Number	Isotope​Number	IS	1	RET (1993)
(0020,0015)	Phase Number	Phase​Number	IS	1	RET (1993)
(0020,0016)	Interval Number	Interval​Number	IS	1	RET (1993)
(0020,0017)	Time Slot Number	Time​Slot​Number	IS	1	RET (1993)
(0020,0018)	Angle Number	Angle​Number	IS	1	RET (1993)
(0020,0019)	Item Number	Item​Number	IS	1
(0020,0020)	Patient Orientation	Patient​Orientation	CS	2
(0020,0022)	Overlay Number	Overlay​Number	IS	1	RET (2004)
(0020,0024)	Curve Number	Curve​Number	IS	1	RET (2004)
(0020,0026)	LUT Number	LUT​Number	IS	1	RET (2004)
(0020,0030)	Image Position	Image​Position	DS	3	RET
(0020,0032)	Image Position (Patient)	Image​Position​Patient	DS	3
(0020,0035)	Image Orientation	Image​Orientation	DS	6	RET
(0020,0037)	Image Orientation (Patient)	Image​Orientation​Patient	DS	6
(0020,0050)	Location	Location	DS	1	RET
(0020,0052)	Frame of Reference UID	Frame​Of​Reference​UID	UI	1
(0020,0060)	Laterality	Laterality	CS	1
(0020,0062)	Image Laterality	Image​Laterality	CS	1
(0020,0070)	Image Geometry Type	Image​Geometry​Type	LO	1	RET
(0020,0080)	Masking Image	Masking​Image	CS	1-n	RET
(0020,00AA)	Report Number	Report​Number	IS	1	RET (2009)
(0020,0100)	Temporal Position Identifier	Temporal​Position​Identifier	IS	1
(0020,0105)	Number of Temporal Positions	Number​Of​Temporal​Positions	IS	1
(0020,0110)	Temporal Resolution	Temporal​Resolution	DS	1
(0020,0200)	Synchronization Frame of Reference UID	Synchronization​Frame​Of​Reference​UID	UI	1
(0020,0242)	SOP Instance UID of Concatenation Source	SOP​Instance​UID​Of​Concatenation​Source	UI	1
(0020,1000)	Series in Study	Series​In​Study	IS	1	RET (2004)
(0020,1001)	Acquisitions in Series	Acquisitions​In​Series	IS	1	RET
(0020,1002)	Images in Acquisition	Images​In​Acquisition	IS	1
(0020,1003)	Images in Series	Images​In​Series	IS	1	RET (2006)
(0020,1004)	Acquisitions in Study	Acquisitions​In​Study	IS	1	RET (2004)
(0020,1005)	Images in Study	Images​In​Study	IS	1	RET (2006)
(0020,1020)	Reference	Reference	LO	1-n	RET
(0020,103F)	Target Position Reference Indicator	Target​Position​Reference​Indicator	LO	1
(0020,1040)	Position Reference Indicator	Position​Reference​Indicator	LO	1
(0020,1041)	Slice Location	Slice​Location	DS	1
(0020,1070)	Other Study Numbers	Other​Study​Numbers	IS	1-n	RET (2004)
(0020,1200)	Number of Patient Related Studies	Number​Of​Patient​Related​Studies	IS	1
(0020,1202)	Number of Patient Related Series	Number​Of​Patient​Related​Series	IS	1
(0020,1204)	Number of Patient Related Instances	Number​Of​Patient​Related​Instances	IS	1
(0020,1206)	Number of Study Related Series	Number​Of​Study​Related​Series	IS	1
(0020,1208)	Number of Study Related Instances	Number​Of​Study​Related​Instances	IS	1
(0020,1209)	Number of Series Related Instances	Number​Of​Series​Related​Instances	IS	1
(0020,31xx)	Source Image IDs	Source​Image​IDs	CS	1-n	RET
(0020,3401)	Modifying Device ID	Modifying​Device​ID	CS	1	RET
(0020,3402)	Modified Image ID	Modified​Image​ID	CS	1	RET
(0020,3403)	Modified Image Date	Modified​Image​Date	DA	1	RET
(0020,3404)	Modifying Device Manufacturer	Modifying​Device​Manufacturer	LO	1	RET
(0020,3405)	Modified Image Time	Modified​Image​Time	TM	1	RET
(0020,3406)	Modified Image Description	Modified​Image​Description	LO	1	RET
(0020,4000)	Image Comments	Image​Comments	LT	1
(0020,5000)	Original Image Identification	Original​Image​Identification	AT	1-n	RET
(0020,5002)	Original Image Identification Nomenclature	Original​Image​Identification​Nomenclature	LO	1-n	RET
(0020,9056)	Stack ID	Stack​ID	SH	1
(0020,9057)	In-Stack Position Number	In​Stack​Position​Number	UL	1
(0020,9071)	Frame Anatomy Sequence	Frame​Anatomy​Sequence	SQ	1
(0020,9072)	Frame Laterality	Frame​Laterality	CS	1
(0020,9111)	Frame Content Sequence	Frame​Content​Sequence	SQ	1
(0020,9113)	Plane Position Sequence	Plane​Position​Sequence	SQ	1
(0020,9116)	Plane Orientation Sequence	Plane​Orientation​Sequence	SQ	1
(0020,9128)	Temporal Position Index	Temporal​Position​Index	UL	1
(0020,9153)	Nominal Cardiac Trigger Delay Time	Nominal​Cardiac​Trigger​Delay​Time	FD	1
(0020,9154)	Nominal Cardiac Trigger Time Prior To R-Peak	Nominal​Cardiac​Trigger​Time​Prior​To​R​Peak	FL	1
(0020,9155)	Actual Cardiac Trigger Time Prior To R-Peak	Actual​Cardiac​Trigger​Time​Prior​To​R​Peak	FL	1
(0020,9156)	Frame Acquisition Number	Frame​Acquisition​Number	US	1
(0020,9157)	Dimension Index Values	Dimension​Index​Values	UL	1-n
(0020,9158)	Frame Comments	Frame​Comments	LT	1
(0020,9161)	Concatenation UID	Concatenation​UID	UI	1
(0020,9162)	In-concatenation Number	In​Concatenation​Number	US	1
(0020,9163)	In-concatenation Total Number	In​Concatenation​Total​Number	US	1
(0020,9164)	Dimension Organization UID	Dimension​Organization​UID	UI	1
(0020,9165)	Dimension Index Pointer	Dimension​Index​Pointer	AT	1
(0020,9167)	Functional Group Pointer	Functional​Group​Pointer	AT	1
(0020,9170)	Unassigned Shared Converted Attributes Sequence	Unassigned​Shared​Converted​Attributes​Sequence	SQ	1
(0020,9171)	Unassigned Per-Frame Converted Attributes Sequence	Unassigned​Per​Frame​Converted​Attributes​Sequence	SQ	1
(0020,9172)	Conversion Source Attributes Sequence	Conversion​Source​Attributes​Sequence	SQ	1
(0020,9213)	Dimension Index Private Creator	Dimension​Index​Private​Creator	LO	1
(0020,9221)	Dimension Organization Sequence	Dimension​Organization​Sequence	SQ	1
(0020,9222)	Dimension Index Sequence	Dimension​Index​Sequence	SQ	1
(0020,9228)	Concatenation Frame Offset Number	Concatenation​Frame​Offset​Number	UL	1
(0020,9238)	Functional Group Private Creator	Functional​Group​Private​Creator	LO	1
(0020,9241)	Nominal Percentage of Cardiac Phase	Nominal​Percentage​Of​Cardiac​Phase	FL	1
(0020,9245)	Nominal Percentage of Respiratory Phase	Nominal​Percentage​Of​Respiratory​Phase	FL	1
(0020,9246)	Starting Respiratory Amplitude	Starting​Respiratory​Amplitude	FL	1
(0020,9247)	Starting Respiratory Phase	Starting​Respiratory​Phase	CS	1
(0020,9248)	Ending Respiratory Amplitude	Ending​Respiratory​Amplitude	FL	1
(0020,9249)	Ending Respiratory Phase	Ending​Respiratory​Phase	CS	1
(0020,9250)	Respiratory Trigger Type	Respiratory​Trigger​Type	CS	1
(0020,9251)	R-R Interval Time Nominal	RR​Interval​Time​Nominal	FD	1
(0020,9252)	Actual Cardiac Trigger Delay Time	Actual​Cardiac​Trigger​Delay​Time	FD	1
(0020,9253)	Respiratory Synchronization Sequence	Respiratory​Synchronization​Sequence	SQ	1
(0020,9254)	Respiratory Interval Time	Respiratory​Interval​Time	FD	1
(0020,9255)	Nominal Respiratory Trigger Delay Time	Nominal​Respiratory​Trigger​Delay​Time	FD	1
(0020,9256)	Respiratory Trigger Delay Threshold	Respiratory​Trigger​Delay​Threshold	FD	1
(0020,9257)	Actual Respiratory Trigger Delay Time	Actual​Respiratory​Trigger​Delay​Time	FD	1
(0020,9301)	Image Position (Volume)	Image​Position​Volume	FD	3
(0020,9302)	Image Orientation (Volume)	Image​Orientation​Volume	FD	6
(0020,9307)	Ultrasound Acquisition Geometry	Ultrasound​Acquisition​Geometry	CS	1
(0020,9308)	Apex Position	Apex​Position	FD	3
(0020,9309)	Volume to Transducer Mapping Matrix	Volume​To​Transducer​Mapping​Matrix	FD	16
(0020,930A)	Volume to Table Mapping Matrix	Volume​To​Table​Mapping​Matrix	FD	16
(0020,930B)	Volume to Transducer Relationship	Volume​To​Transducer​Relationship	CS	1
(0020,930C)	Patient Frame of Reference Source	Patient​Frame​Of​Reference​Source	CS	1
(0020,930D)	Temporal Position Time Offset	Temporal​Position​Time​Offset	FD	1
(0020,930E)	Plane Position (Volume) Sequence	Plane​Position​Volume​Sequence	SQ	1
(0020,930F)	Plane Orientation (Volume) Sequence	Plane​Orientation​Volume​Sequence	SQ	1
(0020,9310)	Temporal Position Sequence	Temporal​Position​Sequence	SQ	1
(0020,9311)	Dimension Organization Type	Dimension​Organization​Type	CS	1
(0020,9312)	Volume Frame of Reference UID	Volume​Frame​Of​Reference​UID	UI	1
(0020,9313)	Table Frame of Reference UID	Table​Frame​Of​Reference​UID	UI	1
(0020,9421)	Dimension Description Label	Dimension​Description​Label	LO	1
(0020,9450)	Patient Orientation in Frame Sequence	Patient​Orientation​In​Frame​Sequence	SQ	1
(0020,9453)	Frame Label	Frame​Label	LO	1
(0020,9518)	Acquisition Index	Acquisition​Index	US	1-n
(0020,9529)	Contributing SOP Instances Reference Sequence	Contributing​SOP​Instances​Reference​Sequence	SQ	1
(0020,9536)	Reconstruction Index	Reconstruction​Index	US	1
(0022,0001)	Light Path Filter Pass-Through Wavelength	Light​Path​Filter​Pass​Through​Wavelength	US	1
(0022,0002)	Light Path Filter Pass Band	Light​Path​Filter​Pass​Band	US	2
(0022,0003)	Image Path Filter Pass-Through Wavelength	Image​Path​Filter​Pass​Through​Wavelength	US	1
(0022,0004)	Image Path Filter Pass Band	Image​Path​Filter​Pass​Band	US	2
(0022,0005)	Patient Eye Movement Commanded	Patient​Eye​Movement​Commanded	CS	1
(0022,0006)	Patient Eye Movement Command Code Sequence	Patient​Eye​Movement​Command​Code​Sequence	SQ	1
(0022,0007)	Spherical Lens Power	Spherical​Lens​Power	FL	1
(0022,0008)	Cylinder Lens Power	Cylinder​Lens​Power	FL	1
(0022,0009)	Cylinder Axis	Cylinder​Axis	FL	1
(0022,000A)	Emmetropic Magnification	Emmetropic​Magnification	FL	1
(0022,000B)	Intra Ocular Pressure	Intra​Ocular​Pressure	FL	1
(0022,000C)	Horizontal Field of View	Horizontal​Field​Of​View	FL	1
(0022,000D)	Pupil Dilated	Pupil​Dilated	CS	1
(0022,000E)	Degree of Dilation	Degree​Of​Dilation	FL	1
(0022,0010)	Stereo Baseline Angle	Stereo​Baseline​Angle	FL	1
(0022,0011)	Stereo Baseline Displacement	Stereo​Baseline​Displacement	FL	1
(0022,0012)	Stereo Horizontal Pixel Offset	Stereo​Horizontal​Pixel​Offset	FL	1
(0022,0013)	Stereo Vertical Pixel Offset	Stereo​Vertical​Pixel​Offset	FL	1
(0022,0014)	Stereo Rotation	Stereo​Rotation	FL	1
(0022,0015)	Acquisition Device Type Code Sequence	Acquisition​Device​Type​Code​Sequence	SQ	1
(0022,0016)	Illumination Type Code Sequence	Illumination​Type​Code​Sequence	SQ	1
(0022,0017)	Light Path Filter Type Stack Code Sequence	Light​Path​Filter​Type​Stack​Code​Sequence	SQ	1
(0022,0018)	Image Path Filter Type Stack Code Sequence	Image​Path​Filter​Type​Stack​Code​Sequence	SQ	1
(0022,0019)	Lenses Code Sequence	Lenses​Code​Sequence	SQ	1
(0022,001A)	Channel Description Code Sequence	Channel​Description​Code​Sequence	SQ	1
(0022,001B)	Refractive State Sequence	Refractive​State​Sequence	SQ	1
(0022,001C)	Mydriatic Agent Code Sequence	Mydriatic​Agent​Code​Sequence	SQ	1
(0022,001D)	Relative Image Position Code Sequence	Relative​Image​Position​Code​Sequence	SQ	1
(0022,001E)	Camera Angle of View	Camera​Angle​Of​View	FL	1
(0022,0020)	Stereo Pairs Sequence	Stereo​Pairs​Sequence	SQ	1
(0022,0021)	Left Image Sequence	Left​Image​Sequence	SQ	1
(0022,0022)	Right Image Sequence	Right​Image​Sequence	SQ	1
(0022,0028)	Stereo Pairs Present	Stereo​Pairs​Present	CS	1
(0022,0030)	Axial Length of the Eye	Axial​Length​Of​The​Eye	FL	1
(0022,0031)	Ophthalmic Frame Location Sequence	Ophthalmic​Frame​Location​Sequence	SQ	1
(0022,0032)	Reference Coordinates	Reference​Coordinates	FL	2-2n
(0022,0035)	Depth Spatial Resolution	Depth​Spatial​Resolution	FL	1
(0022,0036)	Maximum Depth Distortion	Maximum​Depth​Distortion	FL	1
(0022,0037)	Along-scan Spatial Resolution	Along​Scan​Spatial​Resolution	FL	1
(0022,0038)	Maximum Along-scan Distortion	Maximum​Along​Scan​Distortion	FL	1
(0022,0039)	Ophthalmic Image Orientation	Ophthalmic​Image​Orientation	CS	1
(0022,0041)	Depth of Transverse Image	Depth​Of​Transverse​Image	FL	1
(0022,0042)	Mydriatic Agent Concentration Units Sequence	Mydriatic​Agent​Concentration​Units​Sequence	SQ	1
(0022,0048)	Across-scan Spatial Resolution	Across​Scan​Spatial​Resolution	FL	1
(0022,0049)	Maximum Across-scan Distortion	Maximum​Across​Scan​Distortion	FL	1
(0022,004E)	Mydriatic Agent Concentration	Mydriatic​Agent​Concentration	DS	1
(0022,0055)	Illumination Wave Length	Illumination​Wave​Length	FL	1
(0022,0056)	Illumination Power	Illumination​Power	FL	1
(0022,0057)	Illumination Bandwidth	Illumination​Bandwidth	FL	1
(0022,0058)	Mydriatic Agent Sequence	Mydriatic​Agent​Sequence	SQ	1
(0022,1007)	Ophthalmic Axial Measurements Right Eye Sequence	Ophthalmic​Axial​Measurements​Right​Eye​Sequence	SQ	1
(0022,1008)	Ophthalmic Axial Measurements Left Eye Sequence	Ophthalmic​Axial​Measurements​Left​Eye​Sequence	SQ	1
(0022,1009)	Ophthalmic Axial Measurements Device Type	Ophthalmic​Axial​Measurements​Device​Type	CS	1
(0022,1010)	Ophthalmic Axial Length Measurements Type	Ophthalmic​Axial​Length​Measurements​Type	CS	1
(0022,1012)	Ophthalmic Axial Length Sequence	Ophthalmic​Axial​Length​Sequence	SQ	1
(0022,1019)	Ophthalmic Axial Length	Ophthalmic​Axial​Length	FL	1
(0022,1024)	Lens Status Code Sequence	Lens​Status​Code​Sequence	SQ	1
(0022,1025)	Vitreous Status Code Sequence	Vitreous​Status​Code​Sequence	SQ	1
(0022,1028)	IOL Formula Code Sequence	IOL​Formula​Code​Sequence	SQ	1
(0022,1029)	IOL Formula Detail	IOL​Formula​Detail	LO	1
(0022,1033)	Keratometer Index	Keratometer​Index	FL	1
(0022,1035)	Source of Ophthalmic Axial Length Code Sequence	Source​Of​Ophthalmic​Axial​Length​Code​Sequence	SQ	1
(0022,1036)	Source of Corneal Size Data Code Sequence	Source​Of​Corneal​Size​Data​Code​Sequence	SQ	1
(0022,1037)	Target Refraction	Target​Refraction	FL	1
(0022,1039)	Refractive Procedure Occurred	Refractive​Procedure​Occurred	CS	1
(0022,1040)	Refractive Surgery Type Code Sequence	Refractive​Surgery​Type​Code​Sequence	SQ	1
(0022,1044)	Ophthalmic Ultrasound Method Code Sequence	Ophthalmic​Ultrasound​Method​Code​Sequence	SQ	1
(0022,1045)	Surgically Induced Astigmatism Sequence	Surgically​Induced​Astigmatism​Sequence	SQ	1
(0022,1046)	Type of Optical Correction	Type​Of​Optical​Correction	CS	1
(0022,1047)	Toric IOL Power Sequence	Toric​IOL​Power​Sequence	SQ	1
(0022,1048)	Predicted Toric Error Sequence	PredictedT​oric​Error​Sequence	SQ	1
(0022,1049)	Pre-Selected for Implantation	Pre​Selected​For​Implantation	CS	1
(0022,104A)	Toric IOL Power for Exact Emmetropia Sequence	Toric​IOL​Power​For​Exact​Emmetropia​Sequence	SQ	1
(0022,104B)	Toric IOL Power for Exact Target Refraction Sequence	Toric​IOL​Power​For​Exact​Target​Refraction​Sequence	SQ	1
(0022,1050)	Ophthalmic Axial Length Measurements Sequence	Ophthalmic​Axial​Length​Measurements​Sequence	SQ	1
(0022,1053)	IOL Power	IOL​Power	FL	1
(0022,1054)	Predicted Refractive Error	Predicted​Refractive​Error	FL	1
(0022,1059)	Ophthalmic Axial Length Velocity	Ophthalmic​Axial​Length​Velocity	FL	1
(0022,1065)	Lens Status Description	Lens​Status​Description	LO	1
(0022,1066)	Vitreous Status Description	Vitreous​Status​Description	LO	1
(0022,1090)	IOL Power Sequence	IOL​Power​Sequence	SQ	1
(0022,1092)	Lens Constant Sequence	Lens​Constant​Sequence	SQ	1
(0022,1093)	IOL Manufacturer	IOL​Manufacturer	LO	1
(0022,1094)	Lens Constant Description	Lens​Constant​Description	LO	1	RET (2011)
(0022,1095)	Implant Name	Implant​Name	LO	1
(0022,1096)	Keratometry Measurement Type Code Sequence	Keratometry​Measurement​Type​Code​Sequence	SQ	1
(0022,1097)	Implant Part Number	Implant​Part​Number	LO	1
(0022,1100)	Referenced Ophthalmic Axial Measurements Sequence	Referenced​Ophthalmic​Axial​Measurements​Sequence	SQ	1
(0022,1101)	Ophthalmic Axial Length Measurements Segment Name Code Sequence	Ophthalmic​Axial​Length​Measurements​Segment​Name​Code​Sequence	SQ	1
(0022,1103)	Refractive Error Before Refractive Surgery Code Sequence	Refractive​Error​Before​Refractive​Surgery​Code​Sequence	SQ	1
(0022,1121)	IOL Power For Exact Emmetropia	IOL​Power​For​Exact​Emmetropia	FL	1
(0022,1122)	IOL Power For Exact Target Refraction	IOL​Power​For​Exact​Target​Refraction	FL	1
(0022,1125)	Anterior Chamber Depth Definition Code Sequence	Anterior​Chamber​Depth​Definition​Code​Sequence	SQ	1
(0022,1127)	Lens Thickness Sequence	Lens​Thickness​Sequence	SQ	1
(0022,1128)	Anterior Chamber Depth Sequence	Anterior​Chamber​Depth​Sequence	SQ	1
(0022,112A)	Calculation Comment Sequence	Calculation​Comment​Sequence	SQ	1
(0022,112B)	Calculation Comment Type	Calculation​Comment​Type	CS	1
(0022,112C)	Calculation Comment	Calculation​Comment	LT	1
(0022,1130)	Lens Thickness	Lens​Thickness	FL	1
(0022,1131)	Anterior Chamber Depth	Anterior​Chamber​Depth	FL	1
(0022,1132)	Source of Lens Thickness Data Code Sequence	Source​Of​Lens​Thickness​Data​Code​Sequence	SQ	1
(0022,1133)	Source of Anterior Chamber Depth Data Code Sequence	Source​Of​Anterior​Chamber​Depth​Data​Code​Sequence	SQ	1
(0022,1134)	Source of Refractive Measurements Sequence	Source​Of​Refractive​Measurements​Sequence	SQ	1
(0022,1135)	Source of Refractive Measurements Code Sequence	Source​Of​Refractive​Measurements​Code​Sequence	SQ	1
(0022,1140)	Ophthalmic Axial Length Measurement Modified	Ophthalmic​Axial​Length​Measurement​Modified	CS	1
(0022,1150)	Ophthalmic Axial Length Data Source Code Sequence	Ophthalmic​Axial​Length​Data​Source​Code​Sequence	SQ	1
(0022,1153)	Ophthalmic Axial Length Acquisition Method Code Sequence	Ophthalmic​Axial​Length​Acquisition​Method​Code​Sequence	SQ	1	RET (2011)
(0022,1155)	Signal to Noise Ratio	Signal​To​Noise​Ratio	FL	1
(0022,1159)	Ophthalmic Axial Length Data Source Description	Ophthalmic​Axial​Length​Data​Source​Description	LO	1
(0022,1210)	Ophthalmic Axial Length Measurements Total Length Sequence	Ophthalmic​Axial​Length​Measurements​Total​Length​Sequence	SQ	1
(0022,1211)	Ophthalmic Axial Length Measurements Segmental Length Sequence	Ophthalmic​Axial​Length​Measurements​Segmental​Length​Sequence	SQ	1
(0022,1212)	Ophthalmic Axial Length Measurements Length Summation Sequence	Ophthalmic​Axial​Length​Measurements​Length​Summation​Sequence	SQ	1
(0022,1220)	Ultrasound Ophthalmic Axial Length Measurements Sequence	Ultrasound​Ophthalmic​Axial​Length​Measurements​Sequence	SQ	1
(0022,1225)	Optical Ophthalmic Axial Length Measurements Sequence	Optical​Ophthalmic​Axial​Length​Measurements​Sequence	SQ	1
(0022,1230)	Ultrasound Selected Ophthalmic Axial Length Sequence	Ultrasound​Selected​Ophthalmic​Axial​Length​Sequence	SQ	1
(0022,1250)	Ophthalmic Axial Length Selection Method Code Sequence	Ophthalmic​Axial​Length​Selection​Method​Code​Sequence	SQ	1
(0022,1255)	Optical Selected Ophthalmic Axial Length Sequence	Optical​Selected​Ophthalmic​Axial​Length​Sequence	SQ	1
(0022,1257)	Selected Segmental Ophthalmic Axial Length Sequence	Selected​Segmental​Ophthalmic​Axial​Length​Sequence	SQ	1
(0022,1260)	Selected Total Ophthalmic Axial Length Sequence	Selected​Total​Ophthalmic​Axial​Length​Sequence	SQ	1
(0022,1262)	Ophthalmic Axial Length Quality Metric Sequence	Ophthalmic​Axial​Length​Quality​Metric​Sequence	SQ	1
(0022,1265)	Ophthalmic Axial Length Quality Metric Type Code Sequence	Ophthalmic​Axial​Length​Quality​Metric​Type​Code​Sequence	SQ	1	RET
(0022,1273)	Ophthalmic Axial Length Quality Metric Type Description	Ophthalmic​Axial​Length​Quality​Metric​Type​Description	LO	1	RET (2011)
(0022,1300)	Intraocular Lens Calculations Right Eye Sequence	Intraocular​Lens​Calculations​Right​Eye​Sequence	SQ	1
(0022,1310)	Intraocular Lens Calculations Left Eye Sequence	Intraocular​Lens​Calculations​Left​Eye​Sequence	SQ	1
(0022,1330)	Referenced Ophthalmic Axial Length Measurement QC Image Sequence	Referenced​Ophthalmic​Axial​Length​Measurement​QC​Image​Sequence	SQ	1
(0022,1415)	Ophthalmic Mapping Device Type	Ophthalmic​Mapping​Device​Type	CS	1
(0022,1420)	Acquisition Method Code Sequence	Acquisition​Method​Code​Sequence	SQ	1
(0022,1423)	Acquisition Method Algorithm Sequence	Acquisition​Method​Algorithm​Sequence	SQ	1
(0022,1436)	Ophthalmic Thickness Map Type Code Sequence	Ophthalmic​Thickness​Map​Type​Code​Sequence	SQ	1
(0022,1443)	Ophthalmic Thickness Mapping Normals Sequence	Ophthalmic​Thickness​Mapping​Normals​Sequence	SQ	1
(0022,1445)	Retinal Thickness Definition Code Sequence	Retinal​Thickness​Definition​Code​Sequence	SQ	1
(0022,1450)	Pixel Value Mapping to Coded Concept Sequence	Pixel​Value​Mapping​To​Coded​Concept​Sequence	SQ	1
(0022,1452)	Mapped Pixel Value	Mapped​Pixel​Value	US or SS	1
(0022,1454)	Pixel Value Mapping Explanation	Pixel​Value​Mapping​Explanation	LO	1
(0022,1458)	Ophthalmic Thickness Map Quality Threshold Sequence	Ophthalmic​Thickness​Map​Quality​Threshold​Sequence	SQ	1
(0022,1460)	Ophthalmic Thickness Map Threshold Quality Rating	Ophthalmic​Thickness​Map​Threshold​Quality​Rating	FL	1
(0022,1463)	Anatomic Structure Reference Point	Anatomic​Structure​Reference​Point	FL	2
(0022,1465)	Registration to Localizer Sequence	Registration​To​Localizer​Sequence	SQ	1
(0022,1466)	Registered Localizer Units	Registered​Localizer​Units	CS	1
(0022,1467)	Registered Localizer Top Left Hand Corner	Registered​Localizer​Top​Left​Hand​Corner	FL	2
(0022,1468)	Registered Localizer Bottom Right Hand Corner	Registered​Localizer​Bottom​Right​Hand​Corner	FL	2
(0022,1470)	Ophthalmic Thickness Map Quality Rating Sequence	Ophthalmic​Thickness​Map​Quality​Rating​Sequence	SQ	1
(0022,1472)	Relevant OPT Attributes Sequence	Relevant​OPT​Attributes​Sequence	SQ	1
(0022,1512)	Transformation Method Code Sequence	Transformation​Method​Code​Sequence	SQ	1
(0022,1513)	Transformation Algorithm Sequence	Transformation​Algorithm​Sequence	SQ	1
(0022,1515)	Ophthalmic Axial Length Method	Ophthalmic​Axial​Length​Method	CS	1
(0022,1517)	Ophthalmic FOV	Ophthalmic​FOV	FL	1
(0022,1518)	Two Dimensional to Three Dimensional Map Sequence	Two​Dimensional​To​Three​Dimensional​Map​Sequence	SQ	1
(0022,1525)	Wide Field Ophthalmic Photography Quality Rating Sequence	Wide​Field​Ophthalmic​Photography​Quality​Rating​Sequence	SQ	1
(0022,1526)	Wide Field Ophthalmic Photography Quality Threshold Sequence	Wide​Field​Ophthalmic​Photography​Quality​Threshold​Sequence	SQ	1
(0022,1527)	Wide Field Ophthalmic Photography Threshold Quality Rating	Wide​Field​Ophthalmic​Photography​Threshold​Quality​Rating	FL	1
(0022,1528)	X Coordinates Center Pixel View Angle	X​Coordinates​Center​Pixel​View​Angle	FL	1
(0022,1529)	Y Coordinates Center Pixel View Angle	Y​Coordinates​Center​Pixel​View​Angle	FL	1
(0022,1530)	Number of Map Points	Number​Of​Map​Points	UL	1
(0022,1531)	Two Dimensional to Three Dimensional Map Data	Two​Dimensional​To​Three​Dimensional​Map​Data	OF	1
(0022,1612)	Derivation Algorithm Sequence	Derivation​Algorithm​Sequence	SQ	1
(0022,1615)	Ophthalmic Image Type Code Sequence	Ophthalmic​Image​Type​Code​Sequence	SQ	1
(0022,1616)	Ophthalmic Image Type Description	Ophthalmic​Image​Type​Description	LO	1
(0022,1618)	Scan Pattern Type Code Sequence	Scan​Pattern​Type​Code​Sequence	SQ	1
(0022,1620)	Referenced Surface Mesh Identification Sequence	Referenced​Surface​Mesh​Identification​Sequence	SQ	1
(0022,1622)	Ophthalmic Volumetric Properties Flag	Ophthalmic​Volumetric​Properties​Flag	CS	1
(0022,1624)	Ophthalmic Anatomic Reference Point X-Coordinate	Ophthalmic​Anatomic​Reference​Point​XCoordinate	FL	1
(0022,1626)	Ophthalmic Anatomic Reference Point Y-Coordinate	Ophthalmic​Anatomic​Reference​Point​YCoordinate	FL	1
(0022,1628)	Ophthalmic En Face Image Quality Rating Sequence	Ophthalmic​En​Face​Image​Quality​Rating​Sequence	SQ	1
(0022,1630)	Quality Threshold	Quality​Threshold	DS	1
(0022,1640)	OCT B-scan Analysis Acquisition Parameters Sequence	OCTBscan​Analysis​Acquisition​Parameters​Sequence	SQ	1
(0022,1642)	Number of B-scans Per Frame	Numberof​Bscans​Per​Frame	UL	1
(0022,1643)	B-scan Slab Thickness	Bscan​Slab​Thickness	FL	1
(0022,1644)	Distance Between B-scan Slabs	Distance​Between​Bscan​Slabs	FL	1
(0022,1645)	B-scan Cycle Time	Bscan​Cycle​Time	FL	1
(0022,1646)	B-scan Cycle Time Vector	Bscan​Cycle​Time​Vector	FL	1-n
(0022,1649)	A-scan Rate	Ascan​Rate	FL	1
(0022,1650)	B-scan Rate	Bscan​Rate	FL	1
(0022,1658)	Surface Mesh Z-Pixel Offset	Surface​Mesh​ZPixel​Offset	UL	1
(0024,0010)	Visual Field Horizontal Extent	Visual​Field​Horizontal​Extent	FL	1
(0024,0011)	Visual Field Vertical Extent	Visual​Field​Vertical​Extent	FL	1
(0024,0012)	Visual Field Shape	Visual​Field​Shape	CS	1
(0024,0016)	Screening Test Mode Code Sequence	Screening​Test​Mode​Code​Sequence	SQ	1
(0024,0018)	Maximum Stimulus Luminance	Maximum​Stimulus​Luminance	FL	1
(0024,0020)	Background Luminance	Background​Luminance	FL	1
(0024,0021)	Stimulus Color Code Sequence	Stimulus​Color​Code​Sequence	SQ	1
(0024,0024)	Background Illumination Color Code Sequence	Background​Illumination​Color​Code​Sequence	SQ	1
(0024,0025)	Stimulus Area	Stimulus​Area	FL	1
(0024,0028)	Stimulus Presentation Time	Stimulus​Presentation​Time	FL	1
(0024,0032)	Fixation Sequence	Fixation​Sequence	SQ	1
(0024,0033)	Fixation Monitoring Code Sequence	Fixation​Monitoring​Code​Sequence	SQ	1
(0024,0034)	Visual Field Catch Trial Sequence	Visual​Field​Catch​Trial​Sequence	SQ	1
(0024,0035)	Fixation Checked Quantity	Fixation​Checked​Quantity	US	1
(0024,0036)	Patient Not Properly Fixated Quantity	Patient​Not​Properly​Fixated​Quantity	US	1
(0024,0037)	Presented Visual Stimuli Data Flag	Presented​Visual​Stimuli​Data​Flag	CS	1
(0024,0038)	Number of Visual Stimuli	Number​Of​Visual​Stimuli	US	1
(0024,0039)	Excessive Fixation Losses Data Flag	Excessive​Fixation​Losses​Data​Flag	CS	1
(0024,0040)	Excessive Fixation Losses	Excessive​Fixation​Losses	CS	1
(0024,0042)	Stimuli Retesting Quantity	Stimuli​Retesting​Quantity	US	1
(0024,0044)	Comments on Patient's Performance of Visual Field	Comments​On​Patient​Performance​Of​Visual​Field	LT	1
(0024,0045)	False Negatives Estimate Flag	False​Negatives​Estimate​Flag	CS	1
(0024,0046)	False Negatives Estimate	False​Negatives​Estimate	FL	1
(0024,0048)	Negative Catch Trials Quantity	Negative​Catch​Trials​Quantity	US	1
(0024,0050)	False Negatives Quantity	False​Negatives​Quantity	US	1
(0024,0051)	Excessive False Negatives Data Flag	Excessive​False​Negatives​Data​Flag	CS	1
(0024,0052)	Excessive False Negatives	Excessive​False​Negatives	CS	1
(0024,0053)	False Positives Estimate Flag	False​Positives​Estimate​Flag	CS	1
(0024,0054)	False Positives Estimate	False​Positives​Estimate	FL	1
(0024,0055)	Catch Trials Data Flag	Catch​Trials​Data​Flag	CS	1
(0024,0056)	Positive Catch Trials Quantity	Positive​Catch​Trials​Quantity	US	1
(0024,0057)	Test Point Normals Data Flag	Test​Point​Normals​Data​Flag	CS	1
(0024,0058)	Test Point Normals Sequence	Test​Point​Normals​Sequence	SQ	1
(0024,0059)	Global Deviation Probability Normals Flag	Global​Deviation​Probability​Normals​Flag	CS	1
(0024,0060)	False Positives Quantity	False​Positives​Quantity	US	1
(0024,0061)	Excessive False Positives Data Flag	Excessive​False​Positives​Data​Flag	CS	1
(0024,0062)	Excessive False Positives	Excessive​False​Positives	CS	1
(0024,0063)	Visual Field Test Normals Flag	Visual​Field​Test​Normals​Flag	CS	1
(0024,0064)	Results Normals Sequence	Results​Normals​Sequence	SQ	1
(0024,0065)	Age Corrected Sensitivity Deviation Algorithm Sequence	Age​Corrected​Sensitivity​Deviation​Algorithm​Sequence	SQ	1
(0024,0066)	Global Deviation From Normal	Global​Deviation​From​Normal	FL	1
(0024,0067)	Generalized Defect Sensitivity Deviation Algorithm Sequence	Generalized​Defect​Sensitivity​Deviation​Algorithm​Sequence	SQ	1
(0024,0068)	Localized Deviation From Normal	Localized​Deviation​From​Normal	FL	1
(0024,0069)	Patient Reliability Indicator	Patient​Reliability​Indicator	LO	1
(0024,0070)	Visual Field Mean Sensitivity	Visual​Field​Mean​Sensitivity	FL	1
(0024,0071)	Global Deviation Probability	Global​Deviation​Probability	FL	1
(0024,0072)	Local Deviation Probability Normals Flag	Local​Deviation​Probability​Normals​Flag	CS	1
(0024,0073)	Localized Deviation Probability	Localized​Deviation​Probability	FL	1
(0024,0074)	Short Term Fluctuation Calculated	Short​Term​Fluctuation​Calculated	CS	1
(0024,0075)	Short Term Fluctuation	Short​Term​Fluctuation	FL	1
(0024,0076)	Short Term Fluctuation Probability Calculated	Short​Term​Fluctuation​Probability​Calculated	CS	1
(0024,0077)	Short Term Fluctuation Probability	Short​Term​Fluctuation​Probability	FL	1
(0024,0078)	Corrected Localized Deviation From Normal Calculated	Corrected​Localized​Deviation​From​Normal​Calculated	CS	1
(0024,0079)	Corrected Localized Deviation From Normal	Corrected​Localized​Deviation​From​Normal	FL	1
(0024,0080)	Corrected Localized Deviation From Normal Probability Calculated	Corrected​Localized​Deviation​From​Normal​Probability​Calculated	CS	1
(0024,0081)	Corrected Localized Deviation From Normal Probability	Corrected​Localized​Deviation​From​Normal​Probability	FL	1
(0024,0083)	Global Deviation Probability Sequence	Global​Deviation​Probability​Sequence	SQ	1
(0024,0085)	Localized Deviation Probability Sequence	Localized​Deviation​Probability​Sequence	SQ	1
(0024,0086)	Foveal Sensitivity Measured	Foveal​Sensitivity​Measured	CS	1
(0024,0087)	Foveal Sensitivity	Foveal​Sensitivity	FL	1
(0024,0088)	Visual Field Test Duration	Visual​Field​Test​Duration	FL	1
(0024,0089)	Visual Field Test Point Sequence	Visual​Field​Test​Point​Sequence	SQ	1
(0024,0090)	Visual Field Test Point X-Coordinate	Visual​Field​Test​Point​X​Coordinate	FL	1
(0024,0091)	Visual Field Test Point Y-Coordinate	Visual​Field​Test​Point​Y​Coordinate	FL	1
(0024,0092)	Age Corrected Sensitivity Deviation Value	Age​Corrected​Sensitivity​Deviation​Value	FL	1
(0024,0093)	Stimulus Results	Stimulus​Results	CS	1
(0024,0094)	Sensitivity Value	Sensitivity​Value	FL	1
(0024,0095)	Retest Stimulus Seen	Retest​Stimulus​Seen	CS	1
(0024,0096)	Retest Sensitivity Value	Retest​Sensitivity​Value	FL	1
(0024,0097)	Visual Field Test Point Normals Sequence	Visual​Field​Test​Point​Normals​Sequence	SQ	1
(0024,0098)	Quantified Defect	Quantified​Defect	FL	1
(0024,0100)	Age Corrected Sensitivity Deviation Probability Value	Age​Corrected​Sensitivity​Deviation​Probability​Value	FL	1
(0024,0102)	Generalized Defect Corrected Sensitivity Deviation Flag	Generalized​Defect​Corrected​Sensitivity​Deviation​Flag	CS	1
(0024,0103)	Generalized Defect Corrected Sensitivity Deviation Value	Generalized​Defect​Corrected​Sensitivity​Deviation​Value	FL	1
(0024,0104)	Generalized Defect Corrected Sensitivity Deviation Probability Value	Generalized​Defect​Corrected​Sensitivity​Deviation​Probability​Value	FL	1
(0024,0105)	Minimum Sensitivity Value	Minimum​Sensitivity​Value	FL	1
(0024,0106)	Blind Spot Localized	Blind​Spot​Localized	CS	1
(0024,0107)	Blind Spot X-Coordinate	Blind​Spot​X​Coordinate	FL	1
(0024,0108)	Blind Spot Y-Coordinate	Blind​Spot​Y​Coordinate	FL	1
(0024,0110)	Visual Acuity Measurement Sequence	Visual​Acuity​Measurement​Sequence	SQ	1
(0024,0112)	Refractive Parameters Used on Patient Sequence	Refractive​Parameters​Used​On​Patient​Sequence	SQ	1
(0024,0113)	Measurement Laterality	Measurement​Laterality	CS	1
(0024,0114)	Ophthalmic Patient Clinical Information Left Eye Sequence	Ophthalmic​Patient​Clinical​Information​Left​Eye​Sequence	SQ	1
(0024,0115)	Ophthalmic Patient Clinical Information Right Eye Sequence	Ophthalmic​Patient​Clinical​Information​Right​Eye​Sequence	SQ	1
(0024,0117)	Foveal Point Normative Data Flag	Foveal​Point​Normative​Data​Flag	CS	1
(0024,0118)	Foveal Point Probability Value	Foveal​Point​Probability​Value	FL	1
(0024,0120)	Screening Baseline Measured	Screening​Baseline​Measured	CS	1
(0024,0122)	Screening Baseline Measured Sequence	Screening​Baseline​Measured​Sequence	SQ	1
(0024,0124)	Screening Baseline Type	Screening​Baseline​Type	CS	1
(0024,0126)	Screening Baseline Value	Screening​Baseline​Value	FL	1
(0024,0202)	Algorithm Source	Algorithm​Source	LO	1
(0024,0306)	Data Set Name	Data​Set​Name	LO	1
(0024,0307)	Data Set Version	Data​Set​Version	LO	1
(0024,0308)	Data Set Source	Data​Set​Source	LO	1
(0024,0309)	Data Set Description	Data​Set​Description	LO	1
(0024,0317)	Visual Field Test Reliability Global Index Sequence	Visual​Field​Test​Reliability​Global​Index​Sequence	SQ	1
(0024,0320)	Visual Field Global Results Index Sequence	Visual​Field​Global​Results​Index​Sequence	SQ	1
(0024,0325)	Data Observation Sequence	Data​Observation​Sequence	SQ	1
(0024,0338)	Index Normals Flag	Index​Normals​Flag	CS	1
(0024,0341)	Index Probability	Index​Probability	FL	1
(0024,0344)	Index Probability Sequence	Index​Probability​Sequence	SQ	1
(0028,0002)	Samples per Pixel	Samples​Per​Pixel	US	1
(0028,0003)	Samples per Pixel Used	Samples​Per​Pixel​Used	US	1
(0028,0004)	Photometric Interpretation	Photometric​Interpretation	CS	1
(0028,0005)	Image Dimensions	Image​Dimensions	US	1	RET
(0028,0006)	Planar Configuration	Planar​Configuration	US	1
(0028,0008)	Number of Frames	Number​Of​Frames	IS	1
(0028,0009)	Frame Increment Pointer	Frame​Increment​Pointer	AT	1-n
(0028,000A)	Frame Dimension Pointer	Frame​Dimension​Pointer	AT	1-n
(0028,0010)	Rows	Rows	US	1
(0028,0011)	Columns	Columns	US	1
(0028,0012)	Planes	Planes	US	1	RET (2007)
(0028,0014)	Ultrasound Color Data Present	Ultrasound​Color​Data​Present	US	1
(0028,0020)					RET (2007) - See Note 3
(0028,0030)	Pixel Spacing	Pixel​Spacing	DS	2
(0028,0031)	Zoom Factor	Zoom​Factor	DS	2
(0028,0032)	Zoom Center	Zoom​Center	DS	2
(0028,0034)	Pixel Aspect Ratio	Pixel​Aspect​Ratio	IS	2
(0028,0040)	Image Format	Image​Format	CS	1	RET
(0028,0050)	Manipulated Image	Manipulated​Image	LO	1-n	RET
(0028,0051)	Corrected Image	Corrected​Image	CS	1-n
(0028,005F)	Compression Recognition Code	Compression​Recognition​Code	LO	1	RET (2007)
(0028,0060)	Compression Code	Compression​Code	CS	1	RET
(0028,0061)	Compression Originator	Compression​Originator	SH	1	RET (2007)
(0028,0062)	Compression Label	Compression​Label	LO	1	RET (2007)
(0028,0063)	Compression Description	Compression​Description	SH	1	RET (2007)
(0028,0065)	Compression Sequence	Compression​Sequence	CS	1-n	RET (2007)
(0028,0066)	Compression Step Pointers	Compression​Step​Pointers	AT	1-n	RET (2007)
(0028,0068)	Repeat Interval	Repeat​Interval	US	1	RET (2007)
(0028,0069)	Bits Grouped	Bits​Grouped	US	1	RET (2007)
(0028,0070)	Perimeter Table	Perimeter​Table	US	1-n	RET (2007)
(0028,0071)	Perimeter Value	Perimeter​Value	US or SS	1	RET (2007)
(0028,0080)	Predictor Rows	Predictor​Rows	US	1	RET (2007)
(0028,0081)	Predictor Columns	Predictor​Columns	US	1	RET (2007)
(0028,0082)	Predictor Constants	Predictor​Constants	US	1-n	RET (2007)
(0028,0090)	Blocked Pixels	Blocked​Pixels	CS	1	RET (2007)
(0028,0091)	Block Rows	Block​Rows	US	1	RET (2007)
(0028,0092)	Block Columns	Block​Columns	US	1	RET (2007)
(0028,0093)	Row Overlap	Row​Overlap	US	1	RET (2007)
(0028,0094)	Column Overlap	Column​Overlap	US	1	RET (2007)
(0028,0100)	Bits Allocated	Bits​Allocated	US	1
(0028,0101)	Bits Stored	Bits​Stored	US	1
(0028,0102)	High Bit	High​Bit	US	1
(0028,0103)	Pixel Representation	Pixel​Representation	US	1
(0028,0104)	Smallest Valid Pixel Value	Smallest​Valid​Pixel​Value	US or SS	1	RET
(0028,0105)	Largest Valid Pixel Value	Largest​Valid​Pixel​Value	US or SS	1	RET
(0028,0106)	Smallest Image Pixel Value	Smallest​Image​Pixel​Value	US or SS	1
(0028,0107)	Largest Image Pixel Value	Largest​Image​Pixel​Value	US or SS	1
(0028,0108)	Smallest Pixel Value in Series	Smallest​Pixel​Value​In​Series	US or SS	1
(0028,0109)	Largest Pixel Value in Series	Largest​Pixel​Value​In​Series	US or SS	1
(0028,0110)	Smallest Image Pixel Value in Plane	Smallest​Image​Pixel​Value​In​Plane	US or SS	1	RET (2007)
(0028,0111)	Largest Image Pixel Value in Plane	Largest​Image​Pixel​Value​In​Plane	US or SS	1	RET (2007)
(0028,0120)	Pixel Padding Value	Pixel​Padding​Value	US or SS	1
(0028,0121)	Pixel Padding Range Limit	Pixel​Padding​Range​Limit	US or SS	1
(0028,0122)	Float Pixel Padding Value	Float​Pixel​Padding​Value	FL	1
(0028,0123)	Double Float Pixel Padding Value	Double​Float​Pixel​Padding​Value	FD	1
(0028,0124)	Float Pixel Padding Range Limit	Float​Pixel​Padding​Range​Limit	FL	1
(0028,0125)	Double Float Pixel Padding Range Limit	Double​Float​Pixel​Padding​Range​Limit	FD	1
(0028,0200)	Image Location	Image​Location	US	1	RET
(0028,0300)	Quality Control Image	Quality​Control​Image	CS	1
(0028,0301)	Burned In Annotation	Burned​In​Annotation	CS	1
(0028,0302)	Recognizable Visual Features	Recognizable​Visual​Features	CS	1
(0028,0303)	Longitudinal Temporal Information Modified	Longitudinal​Temporal​Information​Modified	CS	1
(0028,0304)	Referenced Color Palette Instance UID	Referenced​Color​Palette​Instance​UID	UI	1
(0028,0400)	Transform Label	Transform​Label	LO	1	RET (2007)
(0028,0401)	Transform Version Number	Transform​Version​Number	LO	1	RET (2007)
(0028,0402)	Number of Transform Steps	Number​Of​Transform​Steps	US	1	RET (2007)
(0028,0403)	Sequence of Compressed Data	Sequence​Of​Compressed​Data	LO	1-n	RET (2007)
(0028,0404)	Details of Coefficients	Details​Of​Coefficients	AT	1-n	RET (2007)
(0028,04x0)	Rows For Nth Order Coefficients	Rows​For​Nth​Order​Coefficients	US	1	RET (2007)
(0028,04x1)	Columns For Nth Order Coefficients	Columns​For​Nth​Order​Coefficients	US	1	RET (2007)
(0028,04x2)	Coefficient Coding	Coefficient​Coding	LO	1-n	RET (2007)
(0028,04x3)	Coefficient Coding Pointers	Coefficient​Coding​Pointers	AT	1-n	RET (2007)
(0028,0700)	DCT Label	DCT​Label	LO	1	RET (2007)
(0028,0701)	Data Block Description	Data​Block​Description	CS	1-n	RET (2007)
(0028,0702)	Data Block	Data​Block	AT	1-n	RET (2007)
(0028,0710)	Normalization Factor Format	Normalization​Factor​Format	US	1	RET (2007)
(0028,0720)	Zonal Map Number Format	Zonal​Map​Number​Format	US	1	RET (2007)
(0028,0721)	Zonal Map Location	Zonal​Map​Location	AT	1-n	RET (2007)
(0028,0722)	Zonal Map Format	Zonal​Map​Format	US	1	RET (2007)
(0028,0730)	Adaptive Map Format	Adaptive​Map​Format	US	1	RET (2007)
(0028,0740)	Code Number Format	Code​Number​Format	US	1	RET (2007)
(0028,08x0)	Code Label	Code​Label	CS	1-n	RET (2007)
(0028,08x2)	Number of Tables	Number​Of​Tables	US	1	RET (2007)
(0028,08x3)	Code Table Location	Code​Table​Location	AT	1-n	RET (2007)
(0028,08x4)	Bits For Code Word	Bits​For​Code​Word	US	1	RET (2007)
(0028,08x8)	Image Data Location	Image​Data​Location	AT	1-n	RET (2007)
(0028,0A02)	Pixel Spacing Calibration Type	Pixel​Spacing​Calibration​Type	CS	1
(0028,0A04)	Pixel Spacing Calibration Description	Pixel​Spacing​Calibration​Description	LO	1
(0028,1040)	Pixel Intensity Relationship	Pixel​Intensity​Relationship	CS	1
(0028,1041)	Pixel Intensity Relationship Sign	Pixel​Intensity​Relationship​Sign	SS	1
(0028,1050)	Window Center	Window​Center	DS	1-n
(0028,1051)	Window Width	Window​Width	DS	1-n
(0028,1052)	Rescale Intercept	Rescale​Intercept	DS	1
(0028,1053)	Rescale Slope	Rescale​Slope	DS	1
(0028,1054)	Rescale Type	Rescale​Type	LO	1
(0028,1055)	Window Center & Width Explanation	Window​Center​Width​Explanation	LO	1-n
(0028,1056)	VOI LUT Function	VOILUT​Function	CS	1
(0028,1080)	Gray Scale	Gray​Scale	CS	1	RET
(0028,1090)	Recommended Viewing Mode	Recommended​Viewing​Mode	CS	1
(0028,1100)	Gray Lookup Table Descriptor	Gray​Lookup​Table​Descriptor	US or SS	3	RET
(0028,1101)	Red Palette Color Lookup Table Descriptor	Red​Palette​Color​Lookup​Table​Descriptor	US or SS	3
(0028,1102)	Green Palette Color Lookup Table Descriptor	Green​Palette​Color​Lookup​Table​Descriptor	US or SS	3
(0028,1103)	Blue Palette Color Lookup Table Descriptor	Blue​Palette​Color​Lookup​Table​Descriptor	US or SS	3
(0028,1104)	Alpha Palette Color Lookup Table Descriptor	Alpha​Palette​Color​Lookup​Table​Descriptor	US	3
(0028,1111)	Large Red Palette Color Lookup Table Descriptor	Large​Red​Palette​Color​Lookup​Table​Descriptor	US or SS	4	RET (2007)
(0028,1112)	Large Green Palette Color Lookup Table Descriptor	Large​Green​Palette​Color​Lookup​Table​Descriptor	US or SS	4	RET (2007)
(0028,1113)	Large Blue Palette Color Lookup Table Descriptor	Large​Blue​Palette​Color​Lookup​Table​Descriptor	US or SS	4	RET (2007)
(0028,1199)	Palette Color Lookup Table UID	Palette​Color​Lookup​Table​UID	UI	1
(0028,1200)	Gray Lookup Table Data	Gray​Lookup​Table​Data	US or SS or OW	1-n or 1	RET
(0028,1201)	Red Palette Color Lookup Table Data	Red​Palette​Color​Lookup​Table​Data	OW	1
(0028,1202)	Green Palette Color Lookup Table Data	Green​Palette​Color​Lookup​Table​Data	OW	1
(0028,1203)	Blue Palette Color Lookup Table Data	Blue​Palette​Color​Lookup​Table​Data	OW	1
(0028,1204)	Alpha Palette Color Lookup Table Data	Alpha​Palette​Color​Lookup​Table​Data	OW	1
(0028,1211)	Large Red Palette Color Lookup Table Data	Large​Red​Palette​Color​Lookup​Table​Data	OW	1	RET (2007)
(0028,1212)	Large Green Palette Color Lookup Table Data	Large​Green​Palette​Color​Lookup​Table​Data	OW	1	RET (2007)
(0028,1213)	Large Blue Palette Color Lookup Table Data	Large​Blue​Palette​Color​Lookup​Table​Data	OW	1	RET (2007)
(0028,1214)	Large Palette Color Lookup Table UID	Large​Palette​Color​Lookup​Table​UID	UI	1	RET (2007)
(0028,1221)	Segmented Red Palette Color Lookup Table Data	Segmented​Red​Palette​Color​Lookup​Table​Data	OW	1
(0028,1222)	Segmented Green Palette Color Lookup Table Data	Segmented​Green​Palette​Color​Lookup​Table​Data	OW	1
(0028,1223)	Segmented Blue Palette Color Lookup Table Data	Segmented​Blue​Palette​Color​Lookup​Table​Data	OW	1
(0028,1224)	Segmented Alpha Palette Color Lookup Table Data	Segmented​Alpha​Palette​Color​Lookup​Table​Data	OW	1
(0028,1230)	Stored Value Color Range Sequence	Stored​Value​Color​Range​Sequence	SQ	1
(0028,1231)	Minimum Stored Value Mapped	Minimum​Stored​Value​Mapped	FD	1
(0028,1232)	Maximum Stored Value Mapped	Maximum​Stored​Value​Mapped	FD	1
(0028,1300)	Breast Implant Present	Breast​Implant​Present	CS	1
(0028,1350)	Partial View	Partial​View	CS	1
(0028,1351)	Partial View Description	Partial​View​Description	ST	1
(0028,1352)	Partial View Code Sequence	Partial​View​Code​Sequence	SQ	1
(0028,135A)	Spatial Locations Preserved	Spatial​Locations​Preserved	CS	1
(0028,1401)	Data Frame Assignment Sequence	Data​Frame​Assignment​Sequence	SQ	1
(0028,1402)	Data Path Assignment	Data​Path​Assignment	CS	1
(0028,1403)	Bits Mapped to Color Lookup Table	Bits​Mapped​To​Color​Lookup​Table	US	1
(0028,1404)	Blending LUT 1 Sequence	Blending​LUT1Sequence	SQ	1
(0028,1405)	Blending LUT 1 Transfer Function	Blending​LUT1Transfer​Function	CS	1
(0028,1406)	Blending Weight Constant	Blending​Weight​Constant	FD	1
(0028,1407)	Blending Lookup Table Descriptor	Blending​Lookup​Table​Descriptor	US	3
(0028,1408)	Blending Lookup Table Data	Blending​Lookup​Table​Data	OW	1
(0028,140B)	Enhanced Palette Color Lookup Table Sequence	Enhanced​Palette​Color​Lookup​Table​Sequence	SQ	1
(0028,140C)	Blending LUT 2 Sequence	Blending​LUT2Sequence	SQ	1
(0028,140D)	Blending LUT 2 Transfer Function	Blending​LUT2Transfer​Function	CS	1
(0028,140E)	Data Path ID	Data​Path​ID	CS	1
(0028,140F)	RGB LUT Transfer Function	RGBLUT​Transfer​Function	CS	1
(0028,1410)	Alpha LUT Transfer Function	Alpha​LUT​Transfer​Function	CS	1
(0028,2000)	ICC Profile	ICC​Profile	OB	1
(0028,2002)	Color Space	Color​Space	CS	1
(0028,2110)	Lossy Image Compression	Lossy​Image​Compression	CS	1
(0028,2112)	Lossy Image Compression Ratio	Lossy​Image​Compression​Ratio	DS	1-n
(0028,2114)	Lossy Image Compression Method	Lossy​Image​Compression​Method	CS	1-n
(0028,3000)	Modality LUT Sequence	Modality​LUT​Sequence	SQ	1
(0028,3002)	LUT Descriptor	LUT​Descriptor	US or SS	3
(0028,3003)	LUT Explanation	LUT​Explanation	LO	1
(0028,3004)	Modality LUT Type	Modality​LUT​Type	LO	1
(0028,3006)	LUT Data	LUT​Data	US or OW	1-n or 1
(0028,3010)	VOI LUT Sequence	VOILUT​Sequence	SQ	1
(0028,3110)	Softcopy VOI LUT Sequence	Softcopy​VOILUT​Sequence	SQ	1
(0028,4000)	Image Presentation Comments	Image​Presentation​Comments	LT	1	RET
(0028,5000)	Bi-Plane Acquisition Sequence	Bi​Plane​Acquisition​Sequence	SQ	1	RET (2007)
(0028,6010)	Representative Frame Number	Representative​Frame​Number	US	1
(0028,6020)	Frame Numbers of Interest (FOI)	Frame​Numbers​Of​Interest	US	1-n
(0028,6022)	Frame of Interest Description	Frame​Of​Interest​Description	LO	1-n
(0028,6023)	Frame of Interest Type	Frame​Of​Interest​Type	CS	1-n
(0028,6030)	Mask Pointer(s)	Mask​Pointers	US	1-n	RET (2001)
(0028,6040)	R Wave Pointer	R​Wave​Pointer	US	1-n
(0028,6100)	Mask Subtraction Sequence	Mask​Subtraction​Sequence	SQ	1
(0028,6101)	Mask Operation	Mask​Operation	CS	1
(0028,6102)	Applicable Frame Range	Applicable​Frame​Range	US	2-2n
(0028,6110)	Mask Frame Numbers	Mask​Frame​Numbers	US	1-n
(0028,6112)	Contrast Frame Averaging	Contrast​Frame​Averaging	US	1
(0028,6114)	Mask Sub-pixel Shift	Mask​Sub​Pixel​Shift	FL	2
(0028,6120)	TID Offset	TID​Offset	SS	1
(0028,6190)	Mask Operation Explanation	Mask​Operation​Explanation	ST	1
(0028,7000)	Equipment Administrator Sequence	Equipment​Administrator​Sequence	SQ	1
(0028,7001)	Number of Display Subsystems	Number​Of​Display​Subsystems	US	1
(0028,7002)	Current Configuration ID	Current​Configuration​ID	US	1
(0028,7003)	Display Subsystem ID	Display​Subsystem​ID	US	1
(0028,7004)	Display Subsystem Name	Display​Subsystem​Name	SH	1
(0028,7005)	Display Subsystem Description	Display​Subsystem​Description	LO	1
(0028,7006)	System Status	System​Status	CS	1
(0028,7007)	System Status Comment	System​Status​Comment	LO	1
(0028,7008)	Target Luminance Characteristics Sequence	Target​Luminance​Characteristics​Sequence	SQ	1
(0028,7009)	Luminance Characteristics ID	Luminance​Characteristics​ID	US	1
(0028,700A)	Display Subsystem Configuration Sequence	Display​Subsystem​Configuration​Sequence	SQ	1
(0028,700B)	Configuration ID	Configuration​ID	US	1
(0028,700C)	Configuration Name	Configuration​Name	SH	1
(0028,700D)	Configuration Description	Configuration​Description	LO	1
(0028,700E)	Referenced Target Luminance Characteristics ID	Referenced​Target​Luminance​Characteristics​ID	US	1
(0028,700F)	QA Results Sequence	QAResults​Sequence	SQ	1
(0028,7010)	Display Subsystem QA Results Sequence	Display​Subsystem​QAResults​Sequence	SQ	1
(0028,7011)	Configuration QA Results Sequence	Configuration​QAResults​Sequence	SQ	1
(0028,7012)	Measurement Equipment Sequence	Measurement​Equipment​Sequence	SQ	1
(0028,7013)	Measurement Functions	Measurement​Functions	CS	1-n
(0028,7014)	Measurement Equipment Type	Measurement​Equipment​Type	CS	1
(0028,7015)	Visual Evaluation Result Sequence	Visual​Evaluation​Result​Sequence	SQ	1
(0028,7016)	Display Calibration Result Sequence	Display​Calibration​Result​Sequence	SQ	1
(0028,7017)	DDL Value	DDLValue	US	1
(0028,7018)	CIExy White Point	CIExy​White​Point	FL	2
(0028,7019)	Display Function Type	Display​Function​Type	CS	1
(0028,701A)	Gamma Value	Gamma​Value	FL	1
(0028,701B)	Number of Luminance Points	Number​Of​Luminance​Points	US	1
(0028,701C)	Luminance Response Sequence	Luminance​Response​Sequence	SQ	1
(0028,701D)	Target Minimum Luminance	Target​Minimum​Luminance	FL	1
(0028,701E)	Target Maximum Luminance	Target​Maximum​Luminance	FL	1
(0028,701F)	Luminance Value	Luminance​Value	FL	1
(0028,7020)	Luminance Response Description	Luminance​Response​Description	LO	1
(0028,7021)	White Point Flag	White​Point​Flag	CS	1
(0028,7022)	Display Device Type Code Sequence	Display​Device​Type​Code​Sequence	SQ	1
(0028,7023)	Display Subsystem Sequence	Display​Subsystem​Sequence	SQ	1
(0028,7024)	Luminance Result Sequence	Luminance​Result​Sequence	SQ	1
(0028,7025)	Ambient Light Value Source	Ambient​Light​Value​Source	CS	1
(0028,7026)	Measured Characteristics	Measured​Characteristics	CS	1-n
(0028,7027)	Luminance Uniformity Result Sequence	Luminance​Uniformity​Result​Sequence	SQ	1
(0028,7028)	Visual Evaluation Test Sequence	Visual​Evaluation​Test​Sequence	SQ	1
(0028,7029)	Test Result	Test​Result	CS	1
(0028,702A)	Test Result Comment	Test​Result​Comment	LO	1
(0028,702B)	Test Image Validation	Test​Image​Validation	CS	1
(0028,702C)	Test Pattern Code Sequence	Test​Pattern​Code​Sequence	SQ	1
(0028,702D)	Measurement Pattern Code Sequence	Measurement​Pattern​Code​Sequence	SQ	1
(0028,702E)	Visual Evaluation Method Code Sequence	Visual​Evaluation​Method​Code​Sequence	SQ	1
(0028,7FE0)	Pixel Data Provider URL	Pixel​Data​Provider​URL	UR	1
(0028,9001)	Data Point Rows	Data​Point​Rows	UL	1
(0028,9002)	Data Point Columns	Data​Point​Columns	UL	1
(0028,9003)	Signal Domain Columns	Signal​Domain​Columns	CS	1
(0028,9099)	Largest Monochrome Pixel Value	Largest​Monochrome​Pixel​Value	US	1	RET (2003)
(0028,9108)	Data Representation	Data​Representation	CS	1
(0028,9110)	Pixel Measures Sequence	Pixel​Measures​Sequence	SQ	1
(0028,9132)	Frame VOI LUT Sequence	Frame​VOILUT​Sequence	SQ	1
(0028,9145)	Pixel Value Transformation Sequence	Pixel​Value​Transformation​Sequence	SQ	1
(0028,9235)	Signal Domain Rows	Signal​Domain​Rows	CS	1
(0028,9411)	Display Filter Percentage	Display​Filter​Percentage	FL	1
(0028,9415)	Frame Pixel Shift Sequence	Frame​Pixel​Shift​Sequence	SQ	1
(0028,9416)	Subtraction Item ID	Subtraction​Item​ID	US	1
(0028,9422)	Pixel Intensity Relationship LUT Sequence	Pixel​Intensity​Relationship​LUT​Sequence	SQ	1
(0028,9443)	Frame Pixel Data Properties Sequence	Frame​Pixel​Data​Properties​Sequence	SQ	1
(0028,9444)	Geometrical Properties	Geometrical​Properties	CS	1
(0028,9445)	Geometric Maximum Distortion	Geometric​Maximum​Distortion	FL	1
(0028,9446)	Image Processing Applied	Image​Processing​Applied	CS	1-n
(0028,9454)	Mask Selection Mode	Mask​Selection​Mode	CS	1
(0028,9474)	LUT Function	LUT​Function	CS	1
(0028,9478)	Mask Visibility Percentage	Mask​Visibility​Percentage	FL	1
(0028,9501)	Pixel Shift Sequence	Pixel​Shift​Sequence	SQ	1
(0028,9502)	Region Pixel Shift Sequence	Region​Pixel​Shift​Sequence	SQ	1
(0028,9503)	Vertices of the Region	Vertices​Of​The​Region	SS	2-2n
(0028,9505)	Multi-frame Presentation Sequence	Multi​Frame​Presentation​Sequence	SQ	1
(0028,9506)	Pixel Shift Frame Range	Pixel​Shift​Frame​Range	US	2-2n
(0028,9507)	LUT Frame Range	LUT​Frame​Range	US	2-2n
(0028,9520)	Image to Equipment Mapping Matrix	Image​To​Equipment​Mapping​Matrix	DS	16
(0028,9537)	Equipment Coordinate System Identification	Equipment​Coordinate​System​Identification	CS	1
(0032,000A)	Study Status ID	Study​Status​ID	CS	1	RET (2004)
(0032,000C)	Study Priority ID	Study​Priority​ID	CS	1	RET (2004)
(0032,0012)	Study ID Issuer	Study​ID​Issuer	LO	1	RET (2004)
(0032,0032)	Study Verified Date	Study​Verified​Date	DA	1	RET (2004)
(0032,0033)	Study Verified Time	Study​Verified​Time	TM	1	RET (2004)
(0032,0034)	Study Read Date	Study​Read​Date	DA	1	RET (2004)
(0032,0035)	Study Read Time	Study​Read​Time	TM	1	RET (2004)
(0032,1000)	Scheduled Study Start Date	Scheduled​Study​Start​Date	DA	1	RET (2004)
(0032,1001)	Scheduled Study Start Time	Scheduled​Study​Start​Time	TM	1	RET (2004)
(0032,1010)	Scheduled Study Stop Date	Scheduled​Study​Stop​Date	DA	1	RET (2004)
(0032,1011)	Scheduled Study Stop Time	Scheduled​Study​Stop​Time	TM	1	RET (2004)
(0032,1020)	Scheduled Study Location	Scheduled​Study​Location	LO	1	RET (2004)
(0032,1021)	Scheduled Study Location AE Title	Scheduled​Study​Location​AE​Title	AE	1-n	RET (2004)
(0032,1030)	Reason for Study	Reason​For​Study	LO	1	RET (2004)
(0032,1031)	Requesting Physician Identification Sequence	Requesting​Physician​Identification​Sequence	SQ	1
(0032,1032)	Requesting Physician	Requesting​Physician	PN	1
(0032,1033)	Requesting Service	Requesting​Service	LO	1
(0032,1034)	Requesting Service Code Sequence	Requesting​Service​Code​Sequence	SQ	1
(0032,1040)	Study Arrival Date	Study​Arrival​Date	DA	1	RET (2004)
(0032,1041)	Study Arrival Time	Study​Arrival​Time	TM	1	RET (2004)
(0032,1050)	Study Completion Date	Study​Completion​Date	DA	1	RET (2004)
(0032,1051)	Study Completion Time	Study​Completion​Time	TM	1	RET (2004)
(0032,1055)	Study Component Status ID	Study​Component​Status​ID	CS	1	RET (2004)
(0032,1060)	Requested Procedure Description	Requested​Procedure​Description	LO	1
(0032,1064)	Requested Procedure Code Sequence	Requested​Procedure​Code​Sequence	SQ	1
(0032,1066)	Reason for Visit	Reason​For​Visit	UT	1
(0032,1067)	Reason for Visit Code Sequence	Reason​For​Visit​Code​Sequence	SQ	1
(0032,1070)	Requested Contrast Agent	Requested​Contrast​Agent	LO	1
(0032,4000)	Study Comments	Study​Comments	LT	1	RET (2007)
(0038,0004)	Referenced Patient Alias Sequence	Referenced​Patient​Alias​Sequence	SQ	1
(0038,0008)	Visit Status ID	Visit​Status​ID	CS	1
(0038,0010)	Admission ID	Admission​ID	LO	1
(0038,0011)	Issuer of Admission ID	Issuer​Of​Admission​ID	LO	1	RET (2008)
(0038,0014)	Issuer of Admission ID Sequence	Issuer​Of​Admission​ID​Sequence	SQ	1
(0038,0016)	Route of Admissions	Route​Of​Admissions	LO	1
(0038,001A)	Scheduled Admission Date	Scheduled​Admission​Date	DA	1	RET (2004)
(0038,001B)	Scheduled Admission Time	Scheduled​Admission​Time	TM	1	RET (2004)
(0038,001C)	Scheduled Discharge Date	Scheduled​Discharge​Date	DA	1	RET (2004)
(0038,001D)	Scheduled Discharge Time	Scheduled​Discharge​Time	TM	1	RET (2004)
(0038,001E)	Scheduled Patient Institution Residence	Scheduled​Patient​Institution​Residence	LO	1	RET (2004)
(0038,0020)	Admitting Date	Admitting​Date	DA	1
(0038,0021)	Admitting Time	Admitting​Time	TM	1
(0038,0030)	Discharge Date	Discharge​Date	DA	1	RET (2004)
(0038,0032)	Discharge Time	Discharge​Time	TM	1	RET (2004)
(0038,0040)	Discharge Diagnosis Description	Discharge​Diagnosis​Description	LO	1	RET (2004)
(0038,0044)	Discharge Diagnosis Code Sequence	Discharge​Diagnosis​Code​Sequence	SQ	1	RET (2004)
(0038,0050)	Special Needs	Special​Needs	LO	1
(0038,0060)	Service Episode ID	Service​Episode​ID	LO	1
(0038,0061)	Issuer of Service Episode ID	Issuer​Of​Service​Episode​ID	LO	1	RET (2008)
(0038,0062)	Service Episode Description	Service​Episode​Description	LO	1
(0038,0064)	Issuer of Service Episode ID Sequence	Issuer​Of​Service​Episode​ID​Sequence	SQ	1
(0038,0100)	Pertinent Documents Sequence	Pertinent​Documents​Sequence	SQ	1
(0038,0101)	Pertinent Resources Sequence	Pertinent​Resources​Sequence	SQ	1
(0038,0102)	Resource Description	Resource​Description	LO	1
(0038,0300)	Current Patient Location	Current​Patient​Location	LO	1
(0038,0400)	Patient's Institution Residence	Patient​Institution​Residence	LO	1
(0038,0500)	Patient State	Patient​State	LO	1
(0038,0502)	Patient Clinical Trial Participation Sequence	Patient​Clinical​Trial​Participation​Sequence	SQ	1
(0038,4000)	Visit Comments	Visit​Comments	LT	1
(003A,0004)	Waveform Originality	Waveform​Originality	CS	1
(003A,0005)	Number of Waveform Channels	Number​Of​Waveform​Channels	US	1
(003A,0010)	Number of Waveform Samples	Number​Of​Waveform​Samples	UL	1
(003A,001A)	Sampling Frequency	Sampling​Frequency	DS	1
(003A,0020)	Multiplex Group Label	Multiplex​Group​Label	SH	1
(003A,0200)	Channel Definition Sequence	Channel​Definition​Sequence	SQ	1
(003A,0202)	Waveform Channel Number	Waveform​Channel​Number	IS	1
(003A,0203)	Channel Label	Channel​Label	SH	1
(003A,0205)	Channel Status	Channel​Status	CS	1-n
(003A,0208)	Channel Source Sequence	Channel​Source​Sequence	SQ	1
(003A,0209)	Channel Source Modifiers Sequence	Channel​Source​Modifiers​Sequence	SQ	1
(003A,020A)	Source Waveform Sequence	Source​Waveform​Sequence	SQ	1
(003A,020C)	Channel Derivation Description	Channel​Derivation​Description	LO	1
(003A,0210)	Channel Sensitivity	Channel​Sensitivity	DS	1
(003A,0211)	Channel Sensitivity Units Sequence	Channel​Sensitivity​Units​Sequence	SQ	1
(003A,0212)	Channel Sensitivity Correction Factor	Channel​Sensitivity​Correction​Factor	DS	1
(003A,0213)	Channel Baseline	Channel​Baseline	DS	1
(003A,0214)	Channel Time Skew	Channel​Time​Skew	DS	1
(003A,0215)	Channel Sample Skew	Channel​Sample​Skew	DS	1
(003A,0218)	Channel Offset	Channel​Offset	DS	1
(003A,021A)	Waveform Bits Stored	Waveform​Bits​Stored	US	1
(003A,0220)	Filter Low Frequency	Filter​Low​Frequency	DS	1
(003A,0221)	Filter High Frequency	Filter​High​Frequency	DS	1
(003A,0222)	Notch Filter Frequency	Notch​Filter​Frequency	DS	1
(003A,0223)	Notch Filter Bandwidth	Notch​Filter​Bandwidth	DS	1
(003A,0230)	Waveform Data Display Scale	Waveform​Data​Display​Scale	FL	1
(003A,0231)	Waveform Display Background CIELab Value	Waveform​Display​Background​CIE​Lab​Value	US	3
(003A,0240)	Waveform Presentation Group Sequence	Waveform​Presentation​Group​Sequence	SQ	1
(003A,0241)	Presentation Group Number	Presentation​Group​Number	US	1
(003A,0242)	Channel Display Sequence	Channel​Display​Sequence	SQ	1
(003A,0244)	Channel Recommended Display CIELab Value	Channel​Recommended​Display​CIE​Lab​Value	US	3
(003A,0245)	Channel Position	Channel​Position	FL	1
(003A,0246)	Display Shading Flag	Display​Shading​Flag	CS	1
(003A,0247)	Fractional Channel Display Scale	Fractional​Channel​Display​Scale	FL	1
(003A,0248)	Absolute Channel Display Scale	Absolute​Channel​Display​Scale	FL	1
(003A,0300)	Multiplexed Audio Channels Description Code Sequence	Multiplexed​Audio​Channels​Description​Code​Sequence	SQ	1
(003A,0301)	Channel Identification Code	Channel​Identification​Code	IS	1
(003A,0302)	Channel Mode	Channel​Mode	CS	1
(0040,0001)	Scheduled Station AE Title	Scheduled​Station​AE​Title	AE	1-n
(0040,0002)	Scheduled Procedure Step Start Date	Scheduled​Procedure​Step​Start​Date	DA	1
(0040,0003)	Scheduled Procedure Step Start Time	Scheduled​Procedure​Step​Start​Time	TM	1
(0040,0004)	Scheduled Procedure Step End Date	Scheduled​Procedure​Step​End​Date	DA	1
(0040,0005)	Scheduled Procedure Step End Time	Scheduled​Procedure​Step​End​Time	TM	1
(0040,0006)	Scheduled Performing Physician's Name	Scheduled​Performing​Physician​Name	PN	1
(0040,0007)	Scheduled Procedure Step Description	Scheduled​Procedure​Step​Description	LO	1
(0040,0008)	Scheduled Protocol Code Sequence	Scheduled​Protocol​Code​Sequence	SQ	1
(0040,0009)	Scheduled Procedure Step ID	Scheduled​Procedure​Step​ID	SH	1
(0040,000A)	Stage Code Sequence	Stage​Code​Sequence	SQ	1
(0040,000B)	Scheduled Performing Physician Identification Sequence	Scheduled​Performing​Physician​Identification​Sequence	SQ	1
(0040,0010)	Scheduled Station Name	Scheduled​Station​Name	SH	1-n
(0040,0011)	Scheduled Procedure Step Location	Scheduled​Procedure​Step​Location	SH	1
(0040,0012)	Pre-Medication	Pre​Medication	LO	1
(0040,0020)	Scheduled Procedure Step Status	Scheduled​Procedure​Step​Status	CS	1
(0040,0026)	Order Placer Identifier Sequence	Order​Placer​Identifier​Sequence	SQ	1
(0040,0027)	Order Filler Identifier Sequence	Order​Filler​Identifier​Sequence	SQ	1
(0040,0031)	Local Namespace Entity ID	Local​Namespace​Entity​ID	UT	1
(0040,0032)	Universal Entity ID	Universal​Entity​ID	UT	1
(0040,0033)	Universal Entity ID Type	Universal​Entity​ID​Type	CS	1
(0040,0035)	Identifier Type Code	Identifier​Type​Code	CS	1
(0040,0036)	Assigning Facility Sequence	Assigning​Facility​Sequence	SQ	1
(0040,0039)	Assigning Jurisdiction Code Sequence	Assigning​Jurisdiction​Code​Sequence	SQ	1
(0040,003A)	Assigning Agency or Department Code Sequence	Assigning​Agency​Or​Department​Code​Sequence	SQ	1
(0040,0100)	Scheduled Procedure Step Sequence	Scheduled​Procedure​Step​Sequence	SQ	1
(0040,0220)	Referenced Non-Image Composite SOP Instance Sequence	Referenced​Non​Image​Composite​SOP​Instance​Sequence	SQ	1
(0040,0241)	Performed Station AE Title	Performed​Station​AE​Title	AE	1
(0040,0242)	Performed Station Name	Performed​Station​Name	SH	1
(0040,0243)	Performed Location	Performed​Location	SH	1
(0040,0244)	Performed Procedure Step Start Date	Performed​Procedure​Step​Start​Date	DA	1
(0040,0245)	Performed Procedure Step Start Time	Performed​Procedure​Step​Start​Time	TM	1
(0040,0250)	Performed Procedure Step End Date	Performed​Procedure​Step​End​Date	DA	1
(0040,0251)	Performed Procedure Step End Time	Performed​Procedure​Step​End​Time	TM	1
(0040,0252)	Performed Procedure Step Status	Performed​Procedure​Step​Status	CS	1
(0040,0253)	Performed Procedure Step ID	Performed​Procedure​Step​ID	SH	1
(0040,0254)	Performed Procedure Step Description	Performed​Procedure​Step​Description	LO	1
(0040,0255)	Performed Procedure Type Description	Performed​Procedure​Type​Description	LO	1
(0040,0260)	Performed Protocol Code Sequence	Performed​Protocol​Code​Sequence	SQ	1
(0040,0261)	Performed Protocol Type	Performed​Protocol​Type	CS	1
(0040,0270)	Scheduled Step Attributes Sequence	Scheduled​Step​Attributes​Sequence	SQ	1
(0040,0275)	Request Attributes Sequence	Request​Attributes​Sequence	SQ	1
(0040,0280)	Comments on the Performed Procedure Step	Comments​On​The​Performed​Procedure​Step	ST	1
(0040,0281)	Performed Procedure Step Discontinuation Reason Code Sequence	Performed​Procedure​Step​Discontinuation​Reason​Code​Sequence	SQ	1
(0040,0293)	Quantity Sequence	Quantity​Sequence	SQ	1
(0040,0294)	Quantity	Quantity	DS	1
(0040,0295)	Measuring Units Sequence	Measuring​Units​Sequence	SQ	1
(0040,0296)	Billing Item Sequence	Billing​Item​Sequence	SQ	1
(0040,0300)	Total Time of Fluoroscopy	Total​Time​Of​Fluoroscopy	US	1	RET (2017c)
(0040,0301)	Total Number of Exposures	Total​Number​Of​Exposures	US	1	RET (2017c)
(0040,0302)	Entrance Dose	Entrance​Dose	US	1
(0040,0303)	Exposed Area	Exposed​Area	US	1-2
(0040,0306)	Distance Source to Entrance	Distance​Source​To​Entrance	DS	1
(0040,0307)	Distance Source to Support	Distance​Source​To​Support	DS	1	RET (2001)
(0040,030E)	Exposure Dose Sequence	Exposure​Dose​Sequence	SQ	1	RET (2017c)
(0040,0310)	Comments on Radiation Dose	Comments​On​Radiation​Dose	ST	1
(0040,0312)	X-Ray Output	X​Ray​Output	DS	1
(0040,0314)	Half Value Layer	Half​Value​Layer	DS	1
(0040,0316)	Organ Dose	Organ​Dose	DS	1
(0040,0318)	Organ Exposed	Organ​Exposed	CS	1
(0040,0320)	Billing Procedure Step Sequence	Billing​Procedure​Step​Sequence	SQ	1
(0040,0321)	Film Consumption Sequence	Film​Consumption​Sequence	SQ	1
(0040,0324)	Billing Supplies and Devices Sequence	Billing​Supplies​And​Devices​Sequence	SQ	1
(0040,0330)	Referenced Procedure Step Sequence	Referenced​Procedure​Step​Sequence	SQ	1	RET (2001)
(0040,0340)	Performed Series Sequence	Performed​Series​Sequence	SQ	1
(0040,0400)	Comments on the Scheduled Procedure Step	Comments​On​The​Scheduled​Procedure​Step	LT	1
(0040,0440)	Protocol Context Sequence	Protocol​Context​Sequence	SQ	1
(0040,0441)	Content Item Modifier Sequence	Content​Item​Modifier​Sequence	SQ	1
(0040,0500)	Scheduled Specimen Sequence	Scheduled​Specimen​Sequence	SQ	1
(0040,050A)	Specimen Accession Number	Specimen​Accession​Number	LO	1	RET (2008)
(0040,0512)	Container Identifier	Container​Identifier	LO	1
(0040,0513)	Issuer of the Container Identifier Sequence	Issuer​Of​The​Container​Identifier​Sequence	SQ	1
(0040,0515)	Alternate Container Identifier Sequence	Alternate​Container​Identifier​Sequence	SQ	1
(0040,0518)	Container Type Code Sequence	Container​Type​Code​Sequence	SQ	1
(0040,051A)	Container Description	Container​Description	LO	1
(0040,0520)	Container Component Sequence	Container​Component​Sequence	SQ	1
(0040,0550)	Specimen Sequence	Specimen​Sequence	SQ	1	RET (2008)
(0040,0551)	Specimen Identifier	Specimen​Identifier	LO	1
(0040,0552)	Specimen Description Sequence (Trial)	Specimen​Description​Sequence​Trial	SQ	1	RET (2007)
(0040,0553)	Specimen Description (Trial)	Specimen​Description​Trial	ST	1	RET (2007)
(0040,0554)	Specimen UID	Specimen​UID	UI	1
(0040,0555)	Acquisition Context Sequence	Acquisition​Context​Sequence	SQ	1
(0040,0556)	Acquisition Context Description	Acquisition​Context​Description	ST	1
(0040,059A)	Specimen Type Code Sequence	Specimen​Type​Code​Sequence	SQ	1
(0040,0560)	Specimen Description Sequence	Specimen​Description​Sequence	SQ	1
(0040,0562)	Issuer of the Specimen Identifier Sequence	Issuer​Of​The​Specimen​Identifier​Sequence	SQ	1
(0040,0600)	Specimen Short Description	Specimen​Short​Description	LO	1
(0040,0602)	Specimen Detailed Description	Specimen​Detailed​Description	UT	1
(0040,0610)	Specimen Preparation Sequence	Specimen​Preparation​Sequence	SQ	1
(0040,0612)	Specimen Preparation Step Content Item Sequence	Specimen​Preparation​Step​Content​Item​Sequence	SQ	1
(0040,0620)	Specimen Localization Content Item Sequence	Specimen​Localization​Content​Item​Sequence	SQ	1
(0040,06FA)	Slide Identifier	Slide​Identifier	LO	1	RET (2008)
(0040,0710)	Whole Slide Microscopy Image Frame Type Sequence	Whole​Slide​Microscopy​Image​Frame​Type​Sequence	SQ	1
(0040,071A)	Image Center Point Coordinates Sequence	Image​Center​Point​Coordinates​Sequence	SQ	1
(0040,072A)	X Offset in Slide Coordinate System	X​Offset​In​Slide​Coordinate​System	DS	1
(0040,073A)	Y Offset in Slide Coordinate System	Y​Offset​In​Slide​Coordinate​System	DS	1
(0040,074A)	Z Offset in Slide Coordinate System	Z​Offset​In​Slide​Coordinate​System	DS	1
(0040,08D8)	Pixel Spacing Sequence	Pixel​Spacing​Sequence	SQ	1	RET (2009)
(0040,08DA)	Coordinate System Axis Code Sequence	Coordinate​System​Axis​Code​Sequence	SQ	1	RET (2009)
(0040,08EA)	Measurement Units Code Sequence	Measurement​Units​Code​Sequence	SQ	1
(0040,09F8)	Vital Stain Code Sequence (Trial)	Vital​Stain​Code​Sequence​Trial	SQ	1	RET (2007)
(0040,1001)	Requested Procedure ID	Requested​Procedure​ID	SH	1
(0040,1002)	Reason for the Requested Procedure	Reason​For​The​Requested​Procedure	LO	1
(0040,1003)	Requested Procedure Priority	Requested​Procedure​Priority	SH	1
(0040,1004)	Patient Transport Arrangements	Patient​Transport​Arrangements	LO	1
(0040,1005)	Requested Procedure Location	Requested​Procedure​Location	LO	1
(0040,1006)	Placer Order Number / Procedure	Placer​Order​Number​Procedure	SH	1	RET (1998)
(0040,1007)	Filler Order Number / Procedure	Filler​Order​Number​Procedure	SH	1	RET (1998)
(0040,1008)	Confidentiality Code	Confidentiality​Code	LO	1
(0040,1009)	Reporting Priority	Reporting​Priority	SH	1
(0040,100A)	Reason for Requested Procedure Code Sequence	Reason​For​Requested​Procedure​Code​Sequence	SQ	1
(0040,1010)	Names of Intended Recipients of Results	Names​Of​Intended​Recipients​Of​Results	PN	1-n
(0040,1011)	Intended Recipients of Results Identification Sequence	Intended​Recipients​Of​Results​Identification​Sequence	SQ	1
(0040,1012)	Reason For Performed Procedure Code Sequence	Reason​For​Performed​Procedure​Code​Sequence	SQ	1
(0040,1060)	Requested Procedure Description (Trial)	Requested​Procedure​Description​Trial	LO	1	RET (2009)
(0040,1101)	Person Identification Code Sequence	Person​Identification​Code​Sequence	SQ	1
(0040,1102)	Person's Address	Person​Address	ST	1
(0040,1103)	Person's Telephone Numbers	Person​Telephone​Numbers	LO	1-n
(0040,1104)	Person's Telecom Information	Person​Telecom​Information	LT	1
(0040,1400)	Requested Procedure Comments	Requested​Procedure​Comments	LT	1
(0040,2001)	Reason for the Imaging Service Request	Reason​For​The​Imaging​Service​Request	LO	1	RET (2003)
(0040,2004)	Issue Date of Imaging Service Request	Issue​Date​Of​Imaging​Service​Request	DA	1
(0040,2005)	Issue Time of Imaging Service Request	Issue​Time​Of​Imaging​Service​Request	TM	1
(0040,2006)	Placer Order Number / Imaging Service Request (Retired)	Placer​Order​Number​Imaging​Service​Request​Retired	SH	1	RET (1998)
(0040,2007)	Filler Order Number / Imaging Service Request (Retired)	Filler​Order​Number​Imaging​Service​Request​Retired	SH	1	RET (1998)
(0040,2008)	Order Entered By	Order​Entered​By	PN	1
(0040,2009)	Order Enterer's Location	Order​Enterer​Location	SH	1
(0040,2010)	Order Callback Phone Number	Order​Callback​Phone​Number	SH	1
(0040,2011)	Order Callback Telecom Information	Order​Callback​Telecom​Information	LT	1
(0040,2016)	Placer Order Number / Imaging Service Request	Placer​Order​Number​Imaging​Service​Request	LO	1
(0040,2017)	Filler Order Number / Imaging Service Request	Filler​Order​Number​Imaging​Service​Request	LO	1
(0040,2400)	Imaging Service Request Comments	Imaging​Service​Request​Comments	LT	1
(0040,3001)	Confidentiality Constraint on Patient Data Description	Confidentiality​Constraint​On​Patient​Data​Description	LO	1
(0040,4001)	General Purpose Scheduled Procedure Step Status	General​Purpose​Scheduled​Procedure​Step​Status	CS	1	RET (2011)
(0040,4002)	General Purpose Performed Procedure Step Status	General​Purpose​Performed​Procedure​Step​Status	CS	1	RET (2011)
(0040,4003)	General Purpose Scheduled Procedure Step Priority	General​Purpose​Scheduled​Procedure​Step​Priority	CS	1	RET (2011)
(0040,4004)	Scheduled Processing Applications Code Sequence	Scheduled​Processing​Applications​Code​Sequence	SQ	1	RET (2011)
(0040,4005)	Scheduled Procedure Step Start DateTime	Scheduled​Procedure​Step​Start​Date​Time	DT	1
(0040,4006)	Multiple Copies Flag	Multiple​Copies​Flag	CS	1	RET (2011)
(0040,4007)	Performed Processing Applications Code Sequence	Performed​Processing​Applications​Code​Sequence	SQ	1	RET (2017e)
(0040,4008)	Scheduled Procedure Step Expiration DateTime	Scheduled​Procedure​Step​Expiration​Date​Time	DT	1
(0040,4009)	Human Performer Code Sequence	Human​Performer​Code​Sequence	SQ	1
(0040,4010)	Scheduled Procedure Step Modification DateTime	Scheduled​Procedure​Step​Modification​Date​Time	DT	1
(0040,4011)	Expected Completion DateTime	Expected​Completion​Date​Time	DT	1
(0040,4015)	Resulting General Purpose Performed Procedure Steps Sequence	Resulting​General​Purpose​Performed​Procedure​Steps​Sequence	SQ	1	RET (2011)
(0040,4016)	Referenced General Purpose Scheduled Procedure Step Sequence	Referenced​General​Purpose​Scheduled​Procedure​Step​Sequence	SQ	1	RET (2011)
(0040,4018)	Scheduled Workitem Code Sequence	Scheduled​Workitem​Code​Sequence	SQ	1
(0040,4019)	Performed Workitem Code Sequence	Performed​Workitem​Code​Sequence	SQ	1
(0040,4020)	Input Availability Flag	Input​Availability​Flag	CS	1	RET (2017b)
(0040,4021)	Input Information Sequence	Input​Information​Sequence	SQ	1
(0040,4022)	Relevant Information Sequence	Relevant​Information​Sequence	SQ	1	RET (2011)
(0040,4023)	Referenced General Purpose Scheduled Procedure Step Transaction UID	Referenced​General​Purpose​Scheduled​Procedure​Step​Transaction​UID	UI	1	RET (2011)
(0040,4025)	Scheduled Station Name Code Sequence	Scheduled​Station​Name​Code​Sequence	SQ	1
(0040,4026)	Scheduled Station Class Code Sequence	Scheduled​Station​Class​Code​Sequence	SQ	1
(0040,4027)	Scheduled Station Geographic Location Code Sequence	Scheduled​Station​Geographic​Location​Code​Sequence	SQ	1
(0040,4028)	Performed Station Name Code Sequence	Performed​Station​Name​Code​Sequence	SQ	1
(0040,4029)	Performed Station Class Code Sequence	Performed​Station​Class​Code​Sequence	SQ	1
(0040,4030)	Performed Station Geographic Location Code Sequence	Performed​Station​Geographic​Location​Code​Sequence	SQ	1
(0040,4031)	Requested Subsequent Workitem Code Sequence	Requested​Subsequent​Workitem​Code​Sequence	SQ	1	RET (2011)
(0040,4032)	Non-DICOM Output Code Sequence	Non​DICOM​Output​Code​Sequence	SQ	1	RET (2011)
(0040,4033)	Output Information Sequence	Output​Information​Sequence	SQ	1
(0040,4034)	Scheduled Human Performers Sequence	Scheduled​Human​Performers​Sequence	SQ	1
(0040,4035)	Actual Human Performers Sequence	Actual​Human​Performers​Sequence	SQ	1
(0040,4036)	Human Performer's Organization	Human​Performer​Organization	LO	1
(0040,4037)	Human Performer's Name	Human​Performer​Name	PN	1
(0040,4040)	Raw Data Handling	Raw​Data​Handling	CS	1
(0040,4041)	Input Readiness State	Input​Readiness​State	CS	1
(0040,4050)	Performed Procedure Step Start DateTime	Performed​Procedure​Step​Start​Date​Time	DT	1
(0040,4051)	Performed Procedure Step End DateTime	Performed​Procedure​Step​End​Date​Time	DT	1
(0040,4052)	Procedure Step Cancellation DateTime	Procedure​Step​Cancellation​Date​Time	DT	1
(0040,4070)	Output Destination Sequence	Output​Destination​Sequence	SQ	1
(0040,4071)	DICOM Storage Sequence	DICOM​Storage​Sequence	SQ	1
(0040,4072)	STOW-RS Storage Sequence	STOWRS​Storage​Sequence	SQ	1
(0040,4073)	Storage URL	Storage​URL	UR	1
(0040,4074)	XDS Storage Sequence	XDS​Storage​Sequence	SQ	1
(0040,8302)	Entrance Dose in mGy	Entrance​Dose​In​mGy	DS	1
(0040,8303)	Entrance Dose Derivation	Entrance​Dose​Derivation	CS	1
(0040,9092)	Parametric Map Frame Type Sequence	Parametric​​Map​​Frame​Type​Sequence	SQ	1
(0040,9094)	Referenced Image Real World Value Mapping Sequence	Referenced​Image​Real​World​Value​Mapping​Sequence	SQ	1
(0040,9096)	Real World Value Mapping Sequence	Real​World​Value​Mapping​Sequence	SQ	1
(0040,9098)	Pixel Value Mapping Code Sequence	Pixel​Value​Mapping​Code​Sequence	SQ	1
(0040,9210)	LUT Label	LUT​Label	SH	1
(0040,9211)	Real World Value Last Value Mapped	Real​World​Value​Last​Value​Mapped	US or SS	1
(0040,9212)	Real World Value LUT Data	Real​World​Value​LUT​Data	FD	1-n
(0040,9213)	Double Float Real World Value Last Value Mapped	Double​Float​Real​World​Value​Last​Value​Mapped	FD	1
(0040,9214)	Double Float Real World Value First Value Mapped	Double​Float​Real​World​Value​First​Value​Mapped	FD	1
(0040,9216)	Real World Value First Value Mapped	Real​World​Value​First​Value​Mapped	US or SS	1
(0040,9220)	Quantity Definition Sequence	Quantity​​Definition​​Sequence	SQ	1
(0040,9224)	Real World Value Intercept	Real​World​Value​Intercept	FD	1
(0040,9225)	Real World Value Slope	Real​World​Value​Slope	FD	1
(0040,A007)	Findings Flag (Trial)	Findings​Flag​Trial	CS	1	RET (2009)
(0040,A010)	Relationship Type	Relationship​Type	CS	1
(0040,A020)	Findings Sequence (Trial)	Findings​Sequence​Trial	SQ	1	RET (2009)
(0040,A021)	Findings Group UID (Trial)	Findings​Group​UID​Trial	UI	1	RET (2009)
(0040,A022)	Referenced Findings Group UID (Trial)	Referenced​Findings​Group​UID​Trial	UI	1	RET (2009)
(0040,A023)	Findings Group Recording Date (Trial)	Findings​Group​Recording​Date​Trial	DA	1	RET (2009)
(0040,A024)	Findings Group Recording Time (Trial)	Findings​Group​Recording​Time​Trial	TM	1	RET (2009)
(0040,A026)	Findings Source Category Code Sequence (Trial)	Findings​Source​Category​Code​Sequence​Trial	SQ	1	RET (2009)
(0040,A027)	Verifying Organization	Verifying​Organization	LO	1
(0040,A028)	Documenting Organization Identifier Code Sequence (Trial)	Documenting​Organization​Identifier​Code​Sequence​Trial	SQ	1	RET (2009)
(0040,A030)	Verification DateTime	Verification​Date​Time	DT	1
(0040,A032)	Observation DateTime	Observation​Date​Time	DT	1
(0040,A040)	Value Type	Value​Type	CS	1
(0040,A043)	Concept Name Code Sequence	Concept​Name​Code​Sequence	SQ	1
(0040,A047)	Measurement Precision Description (Trial)	Measurement​Precision​Description​Trial	LO	1	RET (2009)
(0040,A050)	Continuity Of Content	Continuity​Of​Content	CS	1
(0040,A057)	Urgency or Priority Alerts (Trial)	Urgency​Or​Priority​Alerts​Trial	CS	1-n	RET (2009)
(0040,A060)	Sequencing Indicator (Trial)	Sequencing​Indicator​Trial	LO	1	RET (2009)
(0040,A066)	Document Identifier Code Sequence (Trial)	Document​Identifier​Code​Sequence​Trial	SQ	1	RET (2009)
(0040,A067)	Document Author (Trial)	Document​Author​Trial	PN	1	RET (2009)
(0040,A068)	Document Author Identifier Code Sequence (Trial)	Document​Author​Identifier​Code​Sequence​Trial	SQ	1	RET (2009)
(0040,A070)	Identifier Code Sequence (Trial)	Identifier​Code​Sequence​Trial	SQ	1	RET (2009)
(0040,A073)	Verifying Observer Sequence	Verifying​Observer​Sequence	SQ	1
(0040,A074)	Object Binary Identifier (Trial)	Object​Binary​Identifier​Trial	OB	1	RET (2009)
(0040,A075)	Verifying Observer Name	Verifying​Observer​Name	PN	1
(0040,A076)	Documenting Observer Identifier Code Sequence (Trial)	Documenting​Observer​Identifier​Code​Sequence​Trial	SQ	1	RET (2009)
(0040,A078)	Author Observer Sequence	Author​Observer​Sequence	SQ	1
(0040,A07A)	Participant Sequence	Participant​Sequence	SQ	1
(0040,A07C)	Custodial Organization Sequence	Custodial​Organization​Sequence	SQ	1
(0040,A080)	Participation Type	Participation​Type	CS	1
(0040,A082)	Participation DateTime	Participation​Date​Time	DT	1
(0040,A084)	Observer Type	Observer​Type	CS	1
(0040,A085)	Procedure Identifier Code Sequence (Trial)	Procedure​Identifier​Code​Sequence​Trial	SQ	1	RET (2009)
(0040,A088)	Verifying Observer Identification Code Sequence	Verifying​Observer​Identification​Code​Sequence	SQ	1
(0040,A089)	Object Directory Binary Identifier (Trial)	Object​Directory​Binary​Identifier​Trial	OB	1	RET (2009)
(0040,A090)	Equivalent CDA Document Sequence	Equivalent​CDA​Document​Sequence	SQ	1	RET (2006)
(0040,A0B0)	Referenced Waveform Channels	Referenced​Waveform​Channels	US	2-2n
(0040,A110)	Date of Document or Verbal Transaction (Trial)	Date​Of​Document​Or​Verbal​Transaction​Trial	DA	1	RET (2009)
(0040,A112)	Time of Document Creation or Verbal Transaction (Trial)	Time​Of​Document​Creation​Or​Verbal​Transaction​Trial	TM	1	RET (2009)
(0040,A120)	DateTime	Date​Time	DT	1
(0040,A121)	Date	Date	DA	1
(0040,A122)	Time	Time	TM	1
(0040,A123)	Person Name	Person​Name	PN	1
(0040,A124)	UID	UID	UI	1
(0040,A125)	Report Status ID (Trial)	Report​Status​ID​Trial	CS	2	RET (2009)
(0040,A130)	Temporal Range Type	Temporal​Range​Type	CS	1
(0040,A132)	Referenced Sample Positions	Referenced​Sample​Positions	UL	1-n
(0040,A136)	Referenced Frame Numbers	Referenced​Frame​Numbers	US	1-n
(0040,A138)	Referenced Time Offsets	Referenced​Time​Offsets	DS	1-n
(0040,A13A)	Referenced DateTime	Referenced​Date​Time	DT	1-n
(0040,A160)	Text Value	Text​Value	UT	1
(0040,A161)	Floating Point Value	Floating​Point​Value	FD	1-n
(0040,A162)	Rational Numerator Value	Rational​Numerator​Value	SL	1-n
(0040,A163)	Rational Denominator Value	Rational​Denominator​Value	UL	1-n
(0040,A167)	Observation Category Code Sequence (Trial)	Observation​Category​Code​Sequence​Trial	SQ	1	RET (2009)
(0040,A168)	Concept Code Sequence	Concept​Code​Sequence	SQ	1
(0040,A16A)	Bibliographic Citation (Trial)	Bibliographic​Citation​Trial	ST	1	RET (2009)
(0040,A170)	Purpose of Reference Code Sequence	Purpose​Of​Reference​Code​Sequence	SQ	1	See Note 1
(0040,A171)	Observation UID	Observation​UID	UI	1
(0040,A172)	Referenced Observation UID (Trial)	Referenced​Observation​UID​Trial	UI	1	RET (2009)
(0040,A173)	Referenced Observation Class (Trial)	Referenced​Observation​Class​Trial	CS	1	RET (2009)
(0040,A174)	Referenced Object Observation Class (Trial)	Referenced​Object​Observation​Class​Trial	CS	1	RET (2009)
(0040,A180)	Annotation Group Number	Annotation​Group​Number	US	1
(0040,A192)	Observation Date (Trial)	Observation​Date​Trial	DA	1	RET (2009)
(0040,A193)	Observation Time (Trial)	Observation​Time​Trial	TM	1	RET (2009)
(0040,A194)	Measurement Automation (Trial)	Measurement​Automation​Trial	CS	1	RET (2009)
(0040,A195)	Modifier Code Sequence	Modifier​Code​Sequence	SQ	1
(0040,A224)	Identification Description (Trial)	Identification​Description​Trial	ST	1	RET (2009)
(0040,A290)	Coordinates Set Geometric Type (Trial)	Coordinates​Set​Geometric​Type​Trial	CS	1	RET (2009)
(0040,A296)	Algorithm Code Sequence (Trial)	Algorithm​Code​Sequence​Trial	SQ	1	RET (2009)
(0040,A297)	Algorithm Description (Trial)	Algorithm​Description​Trial	ST	1	RET (2009)
(0040,A29A)	Pixel Coordinates Set (Trial)	Pixel​Coordinates​Set​Trial	SL	2-2n	RET (2009)
(0040,A300)	Measured Value Sequence	Measured​Value​Sequence	SQ	1
(0040,A301)	Numeric Value Qualifier Code Sequence	Numeric​Value​Qualifier​Code​Sequence	SQ	1
(0040,A307)	Current Observer (Trial)	Current​Observer​Trial	PN	1	RET (2009)
(0040,A30A)	Numeric Value	Numeric​Value	DS	1-n
(0040,A313)	Referenced Accession Sequence (Trial)	Referenced​Accession​Sequence​Trial	SQ	1	RET (2009)
(0040,A33A)	Report Status Comment (Trial)	Report​Status​Comment​Trial	ST	1	RET (2009)
(0040,A340)	Procedure Context Sequence (Trial)	Procedure​Context​Sequence​Trial	SQ	1	RET (2009)
(0040,A352)	Verbal Source (Trial)	Verbal​Source​Trial	PN	1	RET (2009)
(0040,A353)	Address (Trial)	Address​Trial	ST	1	RET (2009)
(0040,A354)	Telephone Number (Trial)	Telephone​Number​Trial	LO	1	RET (2009)
(0040,A358)	Verbal Source Identifier Code Sequence (Trial)	Verbal​Source​Identifier​Code​Sequence​Trial	SQ	1	RET (2009)
(0040,A360)	Predecessor Documents Sequence	Predecessor​Documents​Sequence	SQ	1
(0040,A370)	Referenced Request Sequence	Referenced​Request​Sequence	SQ	1
(0040,A372)	Performed Procedure Code Sequence	Performed​Procedure​Code​Sequence	SQ	1
(0040,A375)	Current Requested Procedure Evidence Sequence	Current​Requested​Procedure​Evidence​Sequence	SQ	1
(0040,A380)	Report Detail Sequence (Trial)	Report​Detail​Sequence​Trial	SQ	1	RET (2009)
(0040,A385)	Pertinent Other Evidence Sequence	Pertinent​Other​Evidence​Sequence	SQ	1
(0040,A390)	HL7 Structured Document Reference Sequence	HL7Structured​Document​Reference​Sequence	SQ	1
(0040,A402)	Observation Subject UID (Trial)	Observation​Subject​UID​Trial	UI	1	RET (2009)
(0040,A403)	Observation Subject Class (Trial)	Observation​Subject​Class​Trial	CS	1	RET (2009)
(0040,A404)	Observation Subject Type Code Sequence (Trial)	Observation​Subject​Type​Code​Sequence​Trial	SQ	1	RET (2009)
(0040,A491)	Completion Flag	Completion​Flag	CS	1
(0040,A492)	Completion Flag Description	Completion​Flag​Description	LO	1
(0040,A493)	Verification Flag	Verification​Flag	CS	1
(0040,A494)	Archive Requested	Archive​Requested	CS	1
(0040,A496)	Preliminary Flag	Preliminary​Flag	CS	1
(0040,A504)	Content Template Sequence	Content​Template​Sequence	SQ	1
(0040,A525)	Identical Documents Sequence	Identical​Documents​Sequence	SQ	1
(0040,A600)	Observation Subject Context Flag (Trial)	Observation​Subject​Context​Flag​Trial	CS	1	RET (2009)
(0040,A601)	Observer Context Flag (Trial)	Observer​Context​Flag​Trial	CS	1	RET (2009)
(0040,A603)	Procedure Context Flag (Trial)	Procedure​Context​Flag​Trial	CS	1	RET (2009)
(0040,A730)	Content Sequence	Content​Sequence	SQ	1
(0040,A731)	Relationship Sequence (Trial)	Relationship​Sequence​Trial	SQ	1	RET (2009)
(0040,A732)	Relationship Type Code Sequence (Trial)	Relationship​Type​Code​Sequence​Trial	SQ	1	RET (2009)
(0040,A744)	Language Code Sequence (Trial)	Language​Code​Sequence​Trial	SQ	1	RET (2009)
(0040,A992)	Uniform Resource Locator (Trial)	Uniform​Resource​Locator​Trial	ST	1	RET (2009)
(0040,B020)	Waveform Annotation Sequence	Waveform​Annotation​Sequence	SQ	1
(0040,DB00)	Template Identifier	Template​Identifier	CS	1
(0040,DB06)	Template Version	Template​Version	DT	1	RET (2001)
(0040,DB07)	Template Local Version	Template​Local​Version	DT	1	RET (2001)
(0040,DB0B)	Template Extension Flag	Template​Extension​Flag	CS	1	RET (2001)
(0040,DB0C)	Template Extension Organization UID	Template​Extension​Organization​UID	UI	1	RET (2001)
(0040,DB0D)	Template Extension Creator UID	Template​Extension​Creator​UID	UI	1	RET (2001)
(0040,DB73)	Referenced Content Item Identifier	Referenced​Content​Item​Identifier	UL	1-n
(0040,E001)	HL7 Instance Identifier	HL7Instance​Identifier	ST	1
(0040,E004)	HL7 Document Effective Time	HL7Document​Effective​Time	DT	1
(0040,E006)	HL7 Document Type Code Sequence	HL7Document​Type​Code​Sequence	SQ	1
(0040,E008)	Document Class Code Sequence	Document​Class​Code​Sequence	SQ	1
(0040,E010)	Retrieve URI	Retrieve​URI	UR	1
(0040,E011)	Retrieve Location UID	Retrieve​Location​UID	UI	1
(0040,E020)	Type of Instances	Type​Of​Instances	CS	1
(0040,E021)	DICOM Retrieval Sequence	DICOM​Retrieval​Sequence	SQ	1
(0040,E022)	DICOM Media Retrieval Sequence	DICOM​Media​Retrieval​Sequence	SQ	1
(0040,E023)	WADO Retrieval Sequence	WADO​Retrieval​Sequence	SQ	1
(0040,E024)	XDS Retrieval Sequence	XDS​Retrieval​Sequence	SQ	1
(0040,E025)	WADO-RS Retrieval Sequence	WADORS​Retrieval​Sequence	SQ	1
(0040,E030)	Repository Unique ID	Repository​Unique​ID	UI	1
(0040,E031)	Home Community ID	Home​Community​ID	UI	1
(0042,0010)	Document Title	Document​Title	ST	1
(0042,0011)	Encapsulated Document	Encapsulated​Document	OB	1
(0042,0012)	MIME Type of Encapsulated Document	MIME​Type​Of​Encapsulated​Document	LO	1
(0042,0013)	Source Instance Sequence	Source​Instance​Sequence	SQ	1
(0042,0014)	List of MIME Types	List​Of​MIME​Types	LO	1-n
(0044,0001)	Product Package Identifier	Product​Package​Identifier	ST	1
(0044,0002)	Substance Administration Approval	Substance​Administration​Approval	CS	1
(0044,0003)	Approval Status Further Description	Approval​Status​Further​Description	LT	1
(0044,0004)	Approval Status DateTime	Approval​Status​Date​Time	DT	1
(0044,0007)	Product Type Code Sequence	Product​Type​Code​Sequence	SQ	1
(0044,0008)	Product Name	Product​Name	LO	1-n
(0044,0009)	Product Description	Product​Description	LT	1
(0044,000A)	Product Lot Identifier	Product​Lot​Identifier	LO	1
(0044,000B)	Product Expiration DateTime	Product​Expiration​Date​Time	DT	1
(0044,0010)	Substance Administration DateTime	Substance​Administration​Date​Time	DT	1
(0044,0011)	Substance Administration Notes	Substance​Administration​Notes	LO	1
(0044,0012)	Substance Administration Device ID	Substance​Administration​Device​ID	LO	1
(0044,0013)	Product Parameter Sequence	Product​Parameter​Sequence	SQ	1
(0044,0019)	Substance Administration Parameter Sequence	Substance​Administration​Parameter​Sequence	SQ	1
(0044,0100)	Approval Sequence	Approval​Sequence	SQ	1
(0044,0101)	Assertion Code Sequence	Assertion​Code​Sequence	SQ	1
(0044,0102)	Assertion UID	Assertion​UID	UI	1
(0044,0103)	Asserter Identification Sequence	Asserter​Identification​Sequence	SQ	1
(0044,0104)	Assertion DateTime	Assertion​DateTime	DT	1
(0044,0105)	Assertion Expiration DateTime	Assertion​Expiration​DateTime	DT	1
(0044,0106)	Assertion Comments	Assertion​Comments	UT	1
(0044,0107)	Related Assertion Sequence	Related​Assertion​Sequence	SQ	1
(0044,0108)	Referenced Assertion UID	Referenced​Assertion​UID	UI	1
(0044,0109)	Approval Subject Sequence	Approval​Subject​Sequence	SQ	1
(0044,010A)	Organizational Role Code Sequence	Organizational​Role​Code​Sequence	SQ	1
(0046,0012)	Lens Description	Lens​Description	LO	1
(0046,0014)	Right Lens Sequence	Right​Lens​Sequence	SQ	1
(0046,0015)	Left Lens Sequence	Left​Lens​Sequence	SQ	1
(0046,0016)	Unspecified Laterality Lens Sequence	Unspecified​Laterality​Lens​Sequence	SQ	1
(0046,0018)	Cylinder Sequence	Cylinder​Sequence	SQ	1
(0046,0028)	Prism Sequence	Prism​Sequence	SQ	1
(0046,0030)	Horizontal Prism Power	Horizontal​Prism​Power	FD	1
(0046,0032)	Horizontal Prism Base	Horizontal​Prism​Base	CS	1
(0046,0034)	Vertical Prism Power	Vertical​Prism​Power	FD	1
(0046,0036)	Vertical Prism Base	Vertical​Prism​Base	CS	1
(0046,0038)	Lens Segment Type	Lens​Segment​Type	CS	1
(0046,0040)	Optical Transmittance	Optical​Transmittance	FD	1
(0046,0042)	Channel Width	Channel​Width	FD	1
(0046,0044)	Pupil Size	Pupil​Size	FD	1
(0046,0046)	Corneal Size	Corneal​Size	FD	1
(0046,0047)	Corneal Size Sequence	Corneal​Size​Sequence	SQ	1
(0046,0050)	Autorefraction Right Eye Sequence	Autorefraction​Right​Eye​Sequence	SQ	1
(0046,0052)	Autorefraction Left Eye Sequence	Autorefraction​Left​Eye​Sequence	SQ	1
(0046,0060)	Distance Pupillary Distance	Distance​Pupillary​Distance	FD	1
(0046,0062)	Near Pupillary Distance	Near​Pupillary​Distance	FD	1
(0046,0063)	Intermediate Pupillary Distance	Intermediate​Pupillary​Distance	FD	1
(0046,0064)	Other Pupillary Distance	Other​Pupillary​Distance	FD	1
(0046,0070)	Keratometry Right Eye Sequence	Keratometry​Right​Eye​Sequence	SQ	1
(0046,0071)	Keratometry Left Eye Sequence	Keratometry​Left​Eye​Sequence	SQ	1
(0046,0074)	Steep Keratometric Axis Sequence	Steep​Keratometric​Axis​Sequence	SQ	1
(0046,0075)	Radius of Curvature	Radius​Of​Curvature	FD	1
(0046,0076)	Keratometric Power	Keratometric​Power	FD	1
(0046,0077)	Keratometric Axis	Keratometric​Axis	FD	1
(0046,0080)	Flat Keratometric Axis Sequence	Flat​Keratometric​Axis​Sequence	SQ	1
(0046,0092)	Background Color	Background​Color	CS	1
(0046,0094)	Optotype	Optotype	CS	1
(0046,0095)	Optotype Presentation	Optotype​Presentation	CS	1
(0046,0097)	Subjective Refraction Right Eye Sequence	Subjective​Refraction​Right​Eye​Sequence	SQ	1
(0046,0098)	Subjective Refraction Left Eye Sequence	Subjective​Refraction​Left​Eye​Sequence	SQ	1
(0046,0100)	Add Near Sequence	Add​Near​Sequence	SQ	1
(0046,0101)	Add Intermediate Sequence	Add​Intermediate​Sequence	SQ	1
(0046,0102)	Add Other Sequence	Add​Other​Sequence	SQ	1
(0046,0104)	Add Power	Add​Power	FD	1
(0046,0106)	Viewing Distance	Viewing​Distance	FD	1
(0046,0110)	Cornea Measurements Sequence	Cornea​Measurements​Sequence	SQ	1
(0046,0111)	Source of Cornea Measurement Data Code Sequence	Sourc​eOf​Cornea​Measurement​Data​Code​Sequence	SQ	1
(0046,0112)	Steep Corneal Axis Sequence	Steep​Corneal​Axis​Sequence	SQ	1
(0046,0113)	Flat Corneal Axis Sequence	Flat​Corneal​Axis​Sequence	SQ	1
(0046,0114)	Corneal Power	Corneal​Power	FD	1
(0046,0115)	Corneal Axis	Corneal​Axis	FD	1
(0046,0116)	Cornea Measurement Method Code Sequence	Cornea​Measurement​Method​CodeS​equence	SQ	1
(0046,0117)	Refractive Index of Cornea	Refractive​Index​Of​Cornea	FL	1
(0046,0118)	Refractive Index of Aqueous Humor	Refractive​Index​Of​Aqueous​Humor	FL	1
(0046,0121)	Visual Acuity Type Code Sequence	Visual​Acuity​Type​Code​Sequence	SQ	1
(0046,0122)	Visual Acuity Right Eye Sequence	Visual​Acuity​Right​Eye​Sequence	SQ	1
(0046,0123)	Visual Acuity Left Eye Sequence	Visual​Acuity​Left​Eye​Sequence	SQ	1
(0046,0124)	Visual Acuity Both Eyes Open Sequence	Visual​Acuity​Both​Eyes​Open​Sequence	SQ	1
(0046,0125)	Viewing Distance Type	Viewing​Distance​Type	CS	1
(0046,0135)	Visual Acuity Modifiers	Visual​Acuity​Modifiers	SS	2
(0046,0137)	Decimal Visual Acuity	Decimal​Visual​Acuity	FD	1
(0046,0139)	Optotype Detailed Definition	Optotype​Detailed​Definition	LO	1
(0046,0145)	Referenced Refractive Measurements Sequence	Referenced​Refractive​Measurements​Sequence	SQ	1
(0046,0146)	Sphere Power	Sphere​Power	FD	1
(0046,0147)	Cylinder Power	Cylinder​Power	FD	1
(0046,0201)	Corneal Topography Surface	Corneal​Topography​Surface	CS	1
(0046,0202)	Corneal Vertex Location	Corneal​Vertex​Location	FL	2
(0046,0203)	Pupil Centroid X-Coordinate	Pupil​Centroid​X​Coordinate	FL	1
(0046,0204)	Pupil Centroid Y-Coordinate	Pupil​Centroid​Y​Coordinate	FL	1
(0046,0205)	Equivalent Pupil Radius	Equivalent​Pupil​Radius	FL	1
(0046,0207)	Corneal Topography Map Type Code Sequence	Corneal​Topography​Map​Type​Code​Sequence	SQ	1
(0046,0208)	Vertices of the Outline of Pupil	Vertices​Of​The​Outline​Of​Pupil	IS	2-2n
(0046,0210)	Corneal Topography Mapping Normals Sequence	Corneal​Topography​Mapping​Normals​Sequence	SQ	1
(0046,0211)	Maximum Corneal Curvature Sequence	Maximum​Corneal​Curvature​Sequence	SQ	1
(0046,0212)	Maximum Corneal Curvature	Maximum​Corneal​Curvature	FL	1
(0046,0213)	Maximum Corneal Curvature Location	Maximum​Corneal​Curvature​Location	FL	2
(0046,0215)	Minimum Keratometric Sequence	Minimum​Keratometric​Sequence	SQ	1
(0046,0218)	Simulated Keratometric Cylinder Sequence	Simulated​Keratometric​Cylinder​Sequence	SQ	1
(0046,0220)	Average Corneal Power	Average​Corneal​Power	FL	1
(0046,0224)	Corneal I-S Value	Corneal​IS​Value	FL	1
(0046,0227)	Analyzed Area	Analyzed​Area	FL	1
(0046,0230)	Surface Regularity Index	Surface​Regularity​Index	FL	1
(0046,0232)	Surface Asymmetry Index	Surface​Asymmetry​Index	FL	1
(0046,0234)	Corneal Eccentricity Index	Corneal​Eccentricity​Index	FL	1
(0046,0236)	Keratoconus Prediction Index	Keratoconus​Prediction​Index	FL	1
(0046,0238)	Decimal Potential Visual Acuity	Decimal​Potential​Visual​Acuity	FL	1
(0046,0242)	Corneal Topography Map Quality Evaluation	Corneal​Topography​Map​Quality​Evaluation	CS	1
(0046,0244)	Source Image Corneal Processed Data Sequence	Source​Image​Corneal​Processed​Data​Sequence	SQ	1
(0046,0247)	Corneal Point Location	Corneal​Point​Location	FL	3
(0046,0248)	Corneal Point Estimated	Corneal​Point​Estimated	CS	1
(0046,0249)	Axial Power	Axial​Power	FL	1
(0046,0250)	Tangential Power	Tangential​Power	FL	1
(0046,0251)	Refractive Power	Refractive​Power	FL	1
(0046,0252)	Relative Elevation	Relative​Elevation	FL	1
(0046,0253)	Corneal Wavefront	Corneal​Wavefront	FL	1
(0048,0001)	Imaged Volume Width	Imaged​Volume​Width	FL	1
(0048,0002)	Imaged Volume Height	Imaged​Volume​Height	FL	1
(0048,0003)	Imaged Volume Depth	Imaged​Volume​Depth	FL	1
(0048,0006)	Total Pixel Matrix Columns	Total​Pixel​Matrix​Columns	UL	1
(0048,0007)	Total Pixel Matrix Rows	Total​Pixel​Matrix​Rows	UL	1
(0048,0008)	Total Pixel Matrix Origin Sequence	Total​Pixel​Matrix​Origin​Sequence	SQ	1
(0048,0010)	Specimen Label in Image	Specimen​Label​In​Image	CS	1
(0048,0011)	Focus Method	Focus​Method	CS	1
(0048,0012)	Extended Depth of Field	Extended​Depth​Of​Field	CS	1
(0048,0013)	Number of Focal Planes	Number​Of​Focal​Planes	US	1
(0048,0014)	Distance Between Focal Planes	Distance​Between​Focal​Planes	FL	1
(0048,0015)	Recommended Absent Pixel CIELab Value	Recommended​Absent​Pixel​CIE​Lab​Value	US	3
(0048,0100)	Illuminator Type Code Sequence	Illuminator​Type​Code​Sequence	SQ	1
(0048,0102)	Image Orientation (Slide)	Image​Orientation​Slide	DS	6
(0048,0105)	Optical Path Sequence	Optical​Path​Sequence	SQ	1
(0048,0106)	Optical Path Identifier	Optical​Path​Identifier	SH	1
(0048,0107)	Optical Path Description	Optical​Path​Description	ST	1
(0048,0108)	Illumination Color Code Sequence	Illumination​Color​Code​Sequence	SQ	1
(0048,0110)	Specimen Reference Sequence	Specimen​Reference​Sequence	SQ	1
(0048,0111)	Condenser Lens Power	Condenser​Lens​Power	DS	1
(0048,0112)	Objective Lens Power	Objective​Lens​Power	DS	1
(0048,0113)	Objective Lens Numerical Aperture	Objective​Lens​Numerical​Aperture	DS	1
(0048,0120)	Palette Color Lookup Table Sequence	Palette​Color​Lookup​Table​Sequence	SQ	1
(0048,0200)	Referenced Image Navigation Sequence	Referenced​Image​Navigation​Sequence	SQ	1
(0048,0201)	Top Left Hand Corner of Localizer Area	Top​Left​Hand​Corner​Of​Localizer​Area	US	2
(0048,0202)	Bottom Right Hand Corner of Localizer Area	Bottom​Right​Hand​Corner​Of​Localizer​Area	US	2
(0048,0207)	Optical Path Identification Sequence	Optical​Path​Identification​Sequence	SQ	1
(0048,021A)	Plane Position (Slide) Sequence	Plane​Position​Slide​Sequence	SQ	1
(0048,021E)	Column Position In Total Image Pixel Matrix	Column​Position​In​Total​Image​Pixel​Matrix	SL	1
(0048,021F)	Row Position In Total Image Pixel Matrix	Row​Position​In​Total​Image​Pixel​Matrix	SL	1
(0048,0301)	Pixel Origin Interpretation	Pixel​Origin​Interpretation	CS	1
(0048,0302)	Number of Optical Paths	Number​Of​Optical​Paths	UL	1
(0048,0303)	Total Pixel Matrix Focal Planes	Total​Pixel​Matrix​Focal​Planes	UL	1
(0050,0004)	Calibration Image	Calibration​Image	CS	1
(0050,0010)	Device Sequence	Device​Sequence	SQ	1
(0050,0012)	Container Component Type Code Sequence	Container​Component​Type​Code​Sequence	SQ	1
(0050,0013)	Container Component Thickness	Container​Component​Thickness	FD	1
(0050,0014)	Device Length	Device​Length	DS	1
(0050,0015)	Container Component Width	Container​Component​Width	FD	1
(0050,0016)	Device Diameter	Device​Diameter	DS	1
(0050,0017)	Device Diameter Units	Device​Diameter​Units	CS	1
(0050,0018)	Device Volume	Device​Volume	DS	1
(0050,0019)	Inter-Marker Distance	Inter​Marker​Distance	DS	1
(0050,001A)	Container Component Material	Container​Component​Material	CS	1
(0050,001B)	Container Component ID	Container​Component​ID	LO	1
(0050,001C)	Container Component Length	Container​Component​Length	FD	1
(0050,001D)	Container Component Diameter	Container​Component​Diameter	FD	1
(0050,001E)	Container Component Description	Container​Component​Description	LO	1
(0050,0020)	Device Description	Device​Description	LO	1
(0050,0021)	Long Device Description	Long​Device​Description	ST	1
(0052,0001)	Contrast/Bolus Ingredient Percent by Volume	Contrast​Bolus​Ingredient​Percent​By​Volume	FL	1
(0052,0002)	OCT Focal Distance	OCT​Focal​Distance	FD	1
(0052,0003)	Beam Spot Size	Beam​Spot​Size	FD	1
(0052,0004)	Effective Refractive Index	Effective​Refractive​Index	FD	1
(0052,0006)	OCT Acquisition Domain	OCT​Acquisition​Domain	CS	1
(0052,0007)	OCT Optical Center Wavelength	OCT​Optical​Center​Wavelength	FD	1
(0052,0008)	Axial Resolution	Axial​Resolution	FD	1
(0052,0009)	Ranging Depth	Ranging​Depth	FD	1
(0052,0011)	A-line Rate	A​Line​Rate	FD	1
(0052,0012)	A-lines Per Frame	A​Lines​Per​Frame	US	1
(0052,0013)	Catheter Rotational Rate	Catheter​Rotational​Rate	FD	1
(0052,0014)	A-line Pixel Spacing	A​Line​Pixel​Spacing	FD	1
(0052,0016)	Mode of Percutaneous Access Sequence	Mode​Of​Percutaneous​Access​Sequence	SQ	1
(0052,0025)	Intravascular OCT Frame Type Sequence	Intravascular​OCT​Frame​Type​Sequence	SQ	1
(0052,0026)	OCT Z Offset Applied	OCTZ​Offset​Applied	CS	1
(0052,0027)	Intravascular Frame Content Sequence	Intravascular​Frame​Content​Sequence	SQ	1
(0052,0028)	Intravascular Longitudinal Distance	Intravascular​Longitudinal​Distance	FD	1
(0052,0029)	Intravascular OCT Frame Content Sequence	Intravascular​OCT​Frame​Content​Sequence	SQ	1
(0052,0030)	OCT Z Offset Correction	OCTZ​Offset​Correction	SS	1
(0052,0031)	Catheter Direction of Rotation	Catheter​Direction​Of​Rotation	CS	1
(0052,0033)	Seam Line Location	Seam​Line​Location	FD	1
(0052,0034)	First A-line Location	First​A​Line​Location	FD	1
(0052,0036)	Seam Line Index	Seam​Line​Index	US	1
(0052,0038)	Number of Padded A-lines	Number​Of​Padded​ALines	US	1
(0052,0039)	Interpolation Type	Interpolation​Type	CS	1
(0052,003A)	Refractive Index Applied	Refractive​Index​Applied	CS	1
(0054,0010)	Energy Window Vector	Energy​Window​Vector	US	1-n
(0054,0011)	Number of Energy Windows	Number​Of​Energy​Windows	US	1
(0054,0012)	Energy Window Information Sequence	Energy​Window​Information​Sequence	SQ	1
(0054,0013)	Energy Window Range Sequence	Energy​Window​Range​Sequence	SQ	1
(0054,0014)	Energy Window Lower Limit	Energy​Window​Lower​Limit	DS	1
(0054,0015)	Energy Window Upper Limit	Energy​Window​Upper​Limit	DS	1
(0054,0016)	Radiopharmaceutical Information Sequence	Radio​pharmaceutical​Information​Sequence	SQ	1
(0054,0017)	Residual Syringe Counts	Residual​Syringe​Counts	IS	1
(0054,0018)	Energy Window Name	Energy​Window​Name	SH	1
(0054,0020)	Detector Vector	Detector​Vector	US	1-n
(0054,0021)	Number of Detectors	Number​Of​Detectors	US	1
(0054,0022)	Detector Information Sequence	Detector​Information​Sequence	SQ	1
(0054,0030)	Phase Vector	Phase​Vector	US	1-n
(0054,0031)	Number of Phases	Number​Of​Phases	US	1
(0054,0032)	Phase Information Sequence	Phase​Information​Sequence	SQ	1
(0054,0033)	Number of Frames in Phase	Number​Of​Frames​In​Phase	US	1
(0054,0036)	Phase Delay	Phase​Delay	IS	1
(0054,0038)	Pause Between Frames	Pause​Between​Frames	IS	1
(0054,0039)	Phase Description	Phase​Description	CS	1
(0054,0050)	Rotation Vector	Rotation​Vector	US	1-n
(0054,0051)	Number of Rotations	Number​Of​Rotations	US	1
(0054,0052)	Rotation Information Sequence	Rotation​Information​Sequence	SQ	1
(0054,0053)	Number of Frames in Rotation	Number​Of​Frames​In​Rotation	US	1
(0054,0060)	R-R Interval Vector	RR​Interval​Vector	US	1-n
(0054,0061)	Number of R-R Intervals	Number​Of​RR​Intervals	US	1
(0054,0062)	Gated Information Sequence	Gated​Information​Sequence	SQ	1
(0054,0063)	Data Information Sequence	Data​Information​Sequence	SQ	1
(0054,0070)	Time Slot Vector	Time​Slot​Vector	US	1-n
(0054,0071)	Number of Time Slots	Number​Of​Time​Slots	US	1
(0054,0072)	Time Slot Information Sequence	Time​Slot​Information​Sequence	SQ	1
(0054,0073)	Time Slot Time	Time​Slot​Time	DS	1
(0054,0080)	Slice Vector	Slice​Vector	US	1-n
(0054,0081)	Number of Slices	Number​Of​Slices	US	1
(0054,0090)	Angular View Vector	Angular​View​Vector	US	1-n
(0054,0100)	Time Slice Vector	Time​Slice​Vector	US	1-n
(0054,0101)	Number of Time Slices	Number​Of​Time​Slices	US	1
(0054,0200)	Start Angle	Start​Angle	DS	1
(0054,0202)	Type of Detector Motion	Type​Of​Detector​Motion	CS	1
(0054,0210)	Trigger Vector	Trigger​Vector	IS	1-n
(0054,0211)	Number of Triggers in Phase	Number​Of​Triggers​In​Phase	US	1
(0054,0220)	View Code Sequence	View​Code​Sequence	SQ	1
(0054,0222)	View Modifier Code Sequence	View​Modifier​Code​Sequence	SQ	1
(0054,0300)	Radionuclide Code Sequence	Radionuclide​Code​Sequence	SQ	1
(0054,0302)	Administration Route Code Sequence	Administration​Route​Code​Sequence	SQ	1
(0054,0304)	Radiopharmaceutical Code Sequence	Radio​pharmaceutical​Code​Sequence	SQ	1
(0054,0306)	Calibration Data Sequence	Calibration​Data​Sequence	SQ	1
(0054,0308)	Energy Window Number	Energy​Window​Number	US	1
(0054,0400)	Image ID	Image​ID	SH	1
(0054,0410)	Patient Orientation Code Sequence	Patient​Orientation​Code​Sequence	SQ	1
(0054,0412)	Patient Orientation Modifier Code Sequence	Patient​Orientation​Modifier​Code​Sequence	SQ	1
(0054,0414)	Patient Gantry Relationship Code Sequence	Patient​Gantry​Relationship​Code​Sequence	SQ	1
(0054,0500)	Slice Progression Direction	Slice​Progression​Direction	CS	1
(0054,0501)	Scan Progression Direction	Scan​Progression​Direction	CS	1
(0054,1000)	Series Type	Series​Type	CS	2
(0054,1001)	Units	Units	CS	1
(0054,1002)	Counts Source	Counts​Source	CS	1
(0054,1004)	Reprojection Method	Reprojection​Method	CS	1
(0054,1006)	SUV Type	SUV​Type	CS	1
(0054,1100)	Randoms Correction Method	Randoms​Correction​Method	CS	1
(0054,1101)	Attenuation Correction Method	Attenuation​Correction​Method	LO	1
(0054,1102)	Decay Correction	Decay​Correction	CS	1
(0054,1103)	Reconstruction Method	Reconstruction​Method	LO	1
(0054,1104)	Detector Lines of Response Used	Detector​Lines​Of​Response​Used	LO	1
(0054,1105)	Scatter Correction Method	Scatter​Correction​Method	LO	1
(0054,1200)	Axial Acceptance	Axial​Acceptance	DS	1
(0054,1201)	Axial Mash	Axial​Mash	IS	2
(0054,1202)	Transverse Mash	Transverse​Mash	IS	1
(0054,1203)	Detector Element Size	Detector​Element​Size	DS	2
(0054,1210)	Coincidence Window Width	Coincidence​Window​Width	DS	1
(0054,1220)	Secondary Counts Type	Secondary​Counts​Type	CS	1-n
(0054,1300)	Frame Reference Time	Frame​Reference​Time	DS	1
(0054,1310)	Primary (Prompts) Counts Accumulated	Primary​Prompts​Counts​Accumulated	IS	1
(0054,1311)	Secondary Counts Accumulated	Secondary​Counts​Accumulated	IS	1-n
(0054,1320)	Slice Sensitivity Factor	Slice​Sensitivity​Factor	DS	1
(0054,1321)	Decay Factor	Decay​Factor	DS	1
(0054,1322)	Dose Calibration Factor	Dose​Calibration​Factor	DS	1
(0054,1323)	Scatter Fraction Factor	Scatter​Fraction​Factor	DS	1
(0054,1324)	Dead Time Factor	Dead​Time​Factor	DS	1
(0054,1330)	Image Index	Image​Index	US	1
(0054,1400)	Counts Included	Counts​Included	CS	1-n	RET (2007)
(0054,1401)	Dead Time Correction Flag	Dead​Time​Correction​Flag	CS	1	RET (2007)
(0060,3000)	Histogram Sequence	Histogram​Sequence	SQ	1
(0060,3002)	Histogram Number of Bins	Histogram​Number​Of​Bins	US	1
(0060,3004)	Histogram First Bin Value	Histogram​First​Bin​Value	US or SS	1
(0060,3006)	Histogram Last Bin Value	Histogram​Last​Bin​Value	US or SS	1
(0060,3008)	Histogram Bin Width	Histogram​Bin​Width	US	1
(0060,3010)	Histogram Explanation	Histogram​Explanation	LO	1
(0060,3020)	Histogram Data	Histogram​Data	UL	1-n
(0062,0001)	Segmentation Type	Segmentation​Type	CS	1
(0062,0002)	Segment Sequence	Segment​Sequence	SQ	1
(0062,0003)	Segmented Property Category Code Sequence	Segmented​Property​Category​Code​Sequence	SQ	1
(0062,0004)	Segment Number	Segment​Number	US	1
(0062,0005)	Segment Label	Segment​Label	LO	1
(0062,0006)	Segment Description	Segment​Description	ST	1
(0062,0007)	Segmentation Algorithm Identification Sequence	Segmentation​Algorithm​Identification​Sequence	SQ	1
(0062,0008)	Segment Algorithm Type	Segment​Algorithm​Type	CS	1
(0062,0009)	Segment Algorithm Name	Segment​Algorithm​Name	LO	1
(0062,000A)	Segment Identification Sequence	Segment​Identification​Sequence	SQ	1
(0062,000B)	Referenced Segment Number	Referenced​Segment​Number	US	1-n
(0062,000C)	Recommended Display Grayscale Value	Recommended​Display​Grayscale​Value	US	1
(0062,000D)	Recommended Display CIELab Value	Recommended​Display​CIE​Lab​Value	US	3
(0062,000E)	Maximum Fractional Value	Maximum​Fractional​Value	US	1
(0062,000F)	Segmented Property Type Code Sequence	Segmented​Property​Type​Code​Sequence	SQ	1
(0062,0010)	Segmentation Fractional Type	Segmentation​Fractional​Type	CS	1
(0062,0011)	Segmented Property Type Modifier Code Sequence	Segmented​Property​Type​Modifier​Code​Sequence	SQ	1
(0062,0012)	Used Segments Sequence	Used​Segments​Sequence	SQ	1
(0062,0013)	Segments Overlap	Segments​Overlap	CS	1
(0062,0020)	Tracking ID	Tracking​ID	UT	1
(0062,0021)	Tracking UID	Tracking​UID	UI	1
(0064,0002)	Deformable Registration Sequence	Deformable​Registration​Sequence	SQ	1
(0064,0003)	Source Frame of Reference UID	Source​Frame​Of​Reference​UID	UI	1
(0064,0005)	Deformable Registration Grid Sequence	Deformable​Registration​Grid​Sequence	SQ	1
(0064,0007)	Grid Dimensions	Grid​Dimensions	UL	3
(0064,0008)	Grid Resolution	Grid​Resolution	FD	3
(0064,0009)	Vector Grid Data	Vector​Grid​Data	OF	1
(0064,000F)	Pre Deformation Matrix Registration Sequence	Pre​Deformation​Matrix​Registration​Sequence	SQ	1
(0064,0010)	Post Deformation Matrix Registration Sequence	Post​Deformation​Matrix​Registration​Sequence	SQ	1
(0066,0001)	Number of Surfaces	Number​Of​Surfaces	UL	1
(0066,0002)	Surface Sequence	Surface​Sequence	SQ	1
(0066,0003)	Surface Number	Surface​Number	UL	1
(0066,0004)	Surface Comments	Surface​Comments	LT	1
(0066,0009)	Surface Processing	Surface​Processing	CS	1
(0066,000A)	Surface Processing Ratio	Surface​Processing​Ratio	FL	1
(0066,000B)	Surface Processing Description	Surface​Processing​Description	LO	1
(0066,000C)	Recommended Presentation Opacity	Recommended​Presentation​Opacity	FL	1
(0066,000D)	Recommended Presentation Type	Recommended​Presentation​Type	CS	1
(0066,000E)	Finite Volume	Finite​Volume	CS	1
(0066,0010)	Manifold	Manifold	CS	1
(0066,0011)	Surface Points Sequence	Surface​Points​Sequence	SQ	1
(0066,0012)	Surface Points Normals Sequence	Surface​Points​Normals​Sequence	SQ	1
(0066,0013)	Surface Mesh Primitives Sequence	Surface​Mesh​Primitives​Sequence	SQ	1
(0066,0015)	Number of Surface Points	Number​Of​Surface​Points	UL	1
(0066,0016)	Point Coordinates Data	Point​Coordinates​Data	OF	1
(0066,0017)	Point Position Accuracy	Point​Position​Accuracy	FL	3
(0066,0018)	Mean Point Distance	Mean​Point​Distance	FL	1
(0066,0019)	Maximum Point Distance	Maximum​Point​Distance	FL	1
(0066,001A)	Points Bounding Box Coordinates	Points​Bounding​Box​Coordinates	FL	6
(0066,001B)	Axis of Rotation	Axis​Of​Rotation	FL	3
(0066,001C)	Center of Rotation	Center​Of​Rotation	FL	3
(0066,001E)	Number of Vectors	Number​Of​Vectors	UL	1
(0066,001F)	Vector Dimensionality	Vector​Dimensionality	US	1
(0066,0020)	Vector Accuracy	Vector​Accuracy	FL	1-n
(0066,0021)	Vector Coordinate Data	Vector​Coordinate​Data	OF	1
(0066,0023)	Triangle Point Index List	Triangle​Point​Index​List	OW	1	RET (2015c)
(0066,0024)	Edge Point Index List	Edge​Point​Index​List	OW	1	RET (2015c)
(0066,0025)	Vertex Point Index List	Vertex​Point​Index​List	OW	1	RET (2015c)
(0066,0026)	Triangle Strip Sequence	Triangle​Strip​Sequence	SQ	1
(0066,0027)	Triangle Fan Sequence	Triangle​Fan​Sequence	SQ	1
(0066,0028)	Line Sequence	Line​Sequence	SQ	1
(0066,0029)	Primitive Point Index List	Primitive​Point​Index​List	OW	1	RET (2015c)
(0066,002A)	Surface Count	Surface​Count	UL	1
(0066,002B)	Referenced Surface Sequence	Referenced​Surface​Sequence	SQ	1
(0066,002C)	Referenced Surface Number	Referenced​Surface​Number	UL	1
(0066,002D)	Segment Surface Generation Algorithm Identification Sequence	Segment​Surface​Generation​Algorithm​Identification​Sequence	SQ	1
(0066,002E)	Segment Surface Source Instance Sequence	Segment​Surface​Source​Instance​Sequence	SQ	1
(0066,002F)	Algorithm Family Code Sequence	Algorithm​Family​Code​Sequence	SQ	1
(0066,0030)	Algorithm Name Code Sequence	Algorithm​Name​Code​Sequence	SQ	1
(0066,0031)	Algorithm Version	Algorithm​Version	LO	1
(0066,0032)	Algorithm Parameters	Algorithm​Parameters	LT	1
(0066,0034)	Facet Sequence	Facet​Sequence	SQ	1
(0066,0035)	Surface Processing Algorithm Identification Sequence	Surface​Processing​Algorithm​Identification​Sequence	SQ	1
(0066,0036)	Algorithm Name	Algorithm​Name	LO	1
(0066,0037)	Recommended Point Radius	Recommended​Point​Radius	FL	1
(0066,0038)	Recommended Line Thickness	Recommended​Line​Thickness	FL	1
(0066,0040)	Long Primitive Point Index List	Long​Primitive​Point​Index​List	OL	1
(0066,0041)	Long Triangle Point Index List	Long​Triangle​Point​Index​List	OL	1
(0066,0042)	Long Edge Point Index List	Long​Edge​Point​Index​List	OL	1
(0066,0043)	Long Vertex Point Index List	Long​Vertex​Point​Index​List	OL	1
(0066,0101)	Track Set Sequence	Track​Set​Sequence	SQ	1
(0066,0102)	Track Sequence	Track​Sequence	SQ	1
(0066,0103)	Recommended Display CIELab Value List	Recommended​Display​CIELab​Value​List	OW	1
(0066,0104)	Tracking Algorithm Identification Sequence	Tracking​Algorithm​Identification​Sequence	SQ	1
(0066,0105)	Track Set Number	Track​Set​Number	UL	1
(0066,0106)	Track Set Label	Track​Set​Label	LO	1
(0066,0107)	Track Set Description	Track​Set​Description	UT	1
(0066,0108)	Track Set Anatomical Type Code Sequence	Track​Set​Anatomical​Type​Code​Sequence	SQ	1
(0066,0121)	Measurements Sequence	Measurements​Sequence	SQ	1
(0066,0124)	Track Set Statistics Sequence	Track​Set​Statistics​Sequence	SQ	1
(0066,0125)	Floating Point Values	Floating​Point​Values	OF	1
(0066,0129)	Track Point Index List	Track​Point​Index​List	OL	1
(0066,0130)	Track Statistics Sequence	Track​Statistics​Sequence	SQ	1
(0066,0132)	Measurement Values Sequence	Measurement​Values​Sequence	SQ	1
(0066,0133)	Diffusion Acquisition Code Sequence	Diffusion​Acquisition​Code​Sequence	SQ	1
(0066,0134)	Diffusion Model Code Sequence	Diffusion​Model​Code​Sequence	SQ	1
(0068,6210)	Implant Size	Implant​Size	LO	1
(0068,6221)	Implant Template Version	Implant​Template​Version	LO	1
(0068,6222)	Replaced Implant Template Sequence	Replaced​Implant​Template​Sequence	SQ	1
(0068,6223)	Implant Type	Implant​Type	CS	1
(0068,6224)	Derivation Implant Template Sequence	Derivation​Implant​Template​Sequence	SQ	1
(0068,6225)	Original Implant Template Sequence	Original​Implant​Template​Sequence	SQ	1
(0068,6226)	Effective DateTime	Effective​Date​Time	DT	1
(0068,6230)	Implant Target Anatomy Sequence	Implant​Target​Anatomy​Sequence	SQ	1
(0068,6260)	Information From Manufacturer Sequence	Information​From​Manufacturer​Sequence	SQ	1
(0068,6265)	Notification From Manufacturer Sequence	Notification​From​Manufacturer​Sequence	SQ	1
(0068,6270)	Information Issue DateTime	Information​Issue​Date​Time	DT	1
(0068,6280)	Information Summary	Information​Summary	ST	1
(0068,62A0)	Implant Regulatory Disapproval Code Sequence	Implant​Regulatory​Disapproval​Code​Sequence	SQ	1
(0068,62A5)	Overall Template Spatial Tolerance	Overall​Template​Spatial​Tolerance	FD	1
(0068,62C0)	HPGL Document Sequence	HPGL​Document​Sequence	SQ	1
(0068,62D0)	HPGL Document ID	HPGL​Document​ID	US	1
(0068,62D5)	HPGL Document Label	HPGL​Document​Label	LO	1
(0068,62E0)	View Orientation Code Sequence	View​Orientation​Code​Sequence	SQ	1
(0068,62F0)	View Orientation Modifier Code Sequence	View​Orientation​Modifier​Code​Sequence	SQ	1
(0068,62F2)	HPGL Document Scaling	HPGL​Document​Scaling	FD	1
(0068,6300)	HPGL Document	HPGL​Document	OB	1
(0068,6310)	HPGL Contour Pen Number	HPGL​Contour​Pen​Number	US	1
(0068,6320)	HPGL Pen Sequence	HPGL​Pen​Sequence	SQ	1
(0068,6330)	HPGL Pen Number	HPGL​Pen​Number	US	1
(0068,6340)	HPGL Pen Label	HPGL​Pen​Label	LO	1
(0068,6345)	HPGL Pen Description	HPGL​Pen​Description	ST	1
(0068,6346)	Recommended Rotation Point	Recommended​Rotation​Point	FD	2
(0068,6347)	Bounding Rectangle	Bounding​Rectangle	FD	4
(0068,6350)	Implant Template 3D Model Surface Number	Implant​Template3D​Model​Surface​Number	US	1-n
(0068,6360)	Surface Model Description Sequence	Surface​Model​Description​Sequence	SQ	1
(0068,6380)	Surface Model Label	Surface​Model​Label	LO	1
(0068,6390)	Surface Model Scaling Factor	Surface​Model​Scaling​Factor	FD	1
(0068,63A0)	Materials Code Sequence	Materials​Code​Sequence	SQ	1
(0068,63A4)	Coating Materials Code Sequence	Coating​Materials​Code​Sequence	SQ	1
(0068,63A8)	Implant Type Code Sequence	Implant​Type​Code​Sequence	SQ	1
(0068,63AC)	Fixation Method Code Sequence	Fixation​Method​Code​Sequence	SQ	1
(0068,63B0)	Mating Feature Sets Sequence	Mating​Feature​Sets​Sequence	SQ	1
(0068,63C0)	Mating Feature Set ID	Mating​Feature​Set​ID	US	1
(0068,63D0)	Mating Feature Set Label	Mating​Feature​Set​Label	LO	1
(0068,63E0)	Mating Feature Sequence	Mating​Feature​Sequence	SQ	1
(0068,63F0)	Mating Feature ID	Mating​Feature​ID	US	1
(0068,6400)	Mating Feature Degree of Freedom Sequence	Mating​Feature​Degree​Of​Freedom​Sequence	SQ	1
(0068,6410)	Degree of Freedom ID	Degree​Of​Freedom​ID	US	1
(0068,6420)	Degree of Freedom Type	Degree​Of​Freedom​Type	CS	1
(0068,6430)	2D Mating Feature Coordinates Sequence	Two​D​Mating​Feature​Coordinates​Sequence	SQ	1
(0068,6440)	Referenced HPGL Document ID	Referenced​HPGL​Document​ID	US	1
(0068,6450)	2D Mating Point	Two​D​Mating​Point	FD	2
(0068,6460)	2D Mating Axes	Two​D​Mating​Axes	FD	4
(0068,6470)	2D Degree of Freedom Sequence	Two​D​Degree​Of​Freedom​Sequence	SQ	1
(0068,6490)	3D Degree of Freedom Axis	Three​D​Degree​Of​Freedom​Axis	FD	3
(0068,64A0)	Range of Freedom	Range​Of​Freedom	FD	2
(0068,64C0)	3D Mating Point	Three​D​Mating​Point	FD	3
(0068,64D0)	3D Mating Axes	Three​D​Mating​Axes	FD	9
(0068,64F0)	2D Degree of Freedom Axis	Two​D​Degree​Of​Freedom​Axis	FD	3
(0068,6500)	Planning Landmark Point Sequence	Planning​Landmark​Point​Sequence	SQ	1
(0068,6510)	Planning Landmark Line Sequence	Planning​Landmark​Line​Sequence	SQ	1
(0068,6520)	Planning Landmark Plane Sequence	Planning​Landmark​Plane​Sequence	SQ	1
(0068,6530)	Planning Landmark ID	Planning​Landmark​ID	US	1
(0068,6540)	Planning Landmark Description	Planning​Landmark​Description	LO	1
(0068,6545)	Planning Landmark Identification Code Sequence	Planning​Landmark​Identification​Code​Sequence	SQ	1
(0068,6550)	2D Point Coordinates Sequence	Two​D​Point​Coordinates​Sequence	SQ	1
(0068,6560)	2D Point Coordinates	Two​D​Point​Coordinates	FD	2
(0068,6590)	3D Point Coordinates	Three​D​Point​Coordinates	FD	3
(0068,65A0)	2D Line Coordinates Sequence	Two​D​Line​Coordinates​Sequence	SQ	1
(0068,65B0)	2D Line Coordinates	Two​D​Line​Coordinates	FD	4
(0068,65D0)	3D Line Coordinates	Three​D​Line​Coordinates	FD	6
(0068,65E0)	2D Plane Coordinates Sequence	Two​D​Plane​Coordinates​Sequence	SQ	1
(0068,65F0)	2D Plane Intersection	Two​D​Plane​Intersection	FD	4
(0068,6610)	3D Plane Origin	Three​D​Plane​Origin	FD	3
(0068,6620)	3D Plane Normal	Three​D​Plane​Normal	FD	3
(0068,7001)	Model Modification	Model​Modification	CS	1
(0068,7002)	Model Mirroring	Model​Mirroring	CS	1
(0068,7003)	Model Usage Code Sequence	Model​Usage​Code​Sequence	SQ	1
(0070,0001)	Graphic Annotation Sequence	Graphic​Annotation​Sequence	SQ	1
(0070,0002)	Graphic Layer	Graphic​Layer	CS	1
(0070,0003)	Bounding Box Annotation Units	Bounding​Box​Annotation​Units	CS	1
(0070,0004)	Anchor Point Annotation Units	Anchor​Point​Annotation​Units	CS	1
(0070,0005)	Graphic Annotation Units	Graphic​Annotation​Units	CS	1
(0070,0006)	Unformatted Text Value	Unformatted​Text​Value	ST	1
(0070,0008)	Text Object Sequence	Text​Object​Sequence	SQ	1
(0070,0009)	Graphic Object Sequence	Graphic​Object​Sequence	SQ	1
(0070,0010)	Bounding Box Top Left Hand Corner	Bounding​Box​Top​Left​Hand​Corner	FL	2
(0070,0011)	Bounding Box Bottom Right Hand Corner	Bounding​Box​Bottom​Right​Hand​Corner	FL	2
(0070,0012)	Bounding Box Text Horizontal Justification	Bounding​Box​Text​Horizontal​Justification	CS	1
(0070,0014)	Anchor Point	Anchor​Point	FL	2
(0070,0015)	Anchor Point Visibility	Anchor​Point​Visibility	CS	1
(0070,0020)	Graphic Dimensions	Graphic​Dimensions	US	1
(0070,0021)	Number of Graphic Points	Number​Of​Graphic​Points	US	1
(0070,0022)	Graphic Data	Graphic​Data	FL	2-n
(0070,0023)	Graphic Type	Graphic​Type	CS	1
(0070,0024)	Graphic Filled	Graphic​Filled	CS	1
(0070,0040)	Image Rotation (Retired)	Image​Rotation​Retired	IS	1	RET (2007)
(0070,0041)	Image Horizontal Flip	Image​Horizontal​Flip	CS	1
(0070,0042)	Image Rotation	Image​Rotation	US	1
(0070,0050)	Displayed Area Top Left Hand Corner (Trial)	Displayed​Area​Top​Left​Hand​Corner​Trial	US	2	RET (2007)
(0070,0051)	Displayed Area Bottom Right Hand Corner (Trial)	Displayed​Area​Bottom​Right​Hand​Corner​Trial	US	2	RET (2007)
(0070,0052)	Displayed Area Top Left Hand Corner	Displayed​Area​Top​Left​Hand​Corner	SL	2
(0070,0053)	Displayed Area Bottom Right Hand Corner	Displayed​Area​Bottom​Right​Hand​Corner	SL	2
(0070,005A)	Displayed Area Selection Sequence	Displayed​Area​Selection​Sequence	SQ	1
(0070,0060)	Graphic Layer Sequence	Graphic​Layer​Sequence	SQ	1
(0070,0062)	Graphic Layer Order	Graphic​Layer​Order	IS	1
(0070,0066)	Graphic Layer Recommended Display Grayscale Value	Graphic​Layer​Recommended​Display​Grayscale​Value	US	1
(0070,0067)	Graphic Layer Recommended Display RGB Value	Graphic​Layer​Recommended​Display​RGB​Value	US	3	RET (2004)
(0070,0068)	Graphic Layer Description	Graphic​Layer​Description	LO	1
(0070,0080)	Content Label	Content​Label	CS	1
(0070,0081)	Content Description	Content​Description	LO	1
(0070,0082)	Presentation Creation Date	Presentation​Creation​Date	DA	1
(0070,0083)	Presentation Creation Time	Presentation​Creation​Time	TM	1
(0070,0084)	Content Creator's Name	Content​Creator​Name	PN	1
(0070,0086)	Content Creator's Identification Code Sequence	Content​Creator​Identification​Code​Sequence	SQ	1
(0070,0087)	Alternate Content Description Sequence	Alternate​Content​Description​Sequence	SQ	1
(0070,0100)	Presentation Size Mode	Presentation​Size​Mode	CS	1
(0070,0101)	Presentation Pixel Spacing	Presentation​Pixel​Spacing	DS	2
(0070,0102)	Presentation Pixel Aspect Ratio	Presentation​Pixel​Aspect​Ratio	IS	2
(0070,0103)	Presentation Pixel Magnification Ratio	Presentation​Pixel​Magnification​Ratio	FL	1
(0070,0207)	Graphic Group Label	Graphic​Group​Label	LO	1
(0070,0208)	Graphic Group Description	Graphic​Group​Description	ST	1
(0070,0209)	Compound Graphic Sequence	Compound​Graphic​Sequence	SQ	1
(0070,0226)	Compound Graphic Instance ID	Compound​Graphic​Instance​ID	UL	1
(0070,0227)	Font Name	Font​Name	LO	1
(0070,0228)	Font Name Type	Font​Name​Type	CS	1
(0070,0229)	CSS Font Name	CSS​Font​Name	LO	1
(0070,0230)	Rotation Angle	Rotation​Angle	FD	1
(0070,0231)	Text Style Sequence	Text​Style​Sequence	SQ	1
(0070,0232)	Line Style Sequence	Line​Style​Sequence	SQ	1
(0070,0233)	Fill Style Sequence	Fill​Style​Sequence	SQ	1
(0070,0234)	Graphic Group Sequence	Graphic​Group​Sequence	SQ	1
(0070,0241)	Text Color CIELab Value	Text​Color​CIE​Lab​Value	US	3
(0070,0242)	Horizontal Alignment	Horizontal​Alignment	CS	1
(0070,0243)	Vertical Alignment	Vertical​Alignment	CS	1
(0070,0244)	Shadow Style	Shadow​Style	CS	1
(0070,0245)	Shadow Offset X	Shadow​Offset​X	FL	1
(0070,0246)	Shadow Offset Y	Shadow​Offset​Y	FL	1
(0070,0247)	Shadow Color CIELab Value	Shadow​Color​CIE​Lab​Value	US	3
(0070,0248)	Underlined	Underlined	CS	1
(0070,0249)	Bold	Bold	CS	1
(0070,0250)	Italic	Italic	CS	1
(0070,0251)	Pattern On Color CIELab Value	Pattern​On​Color​CIE​Lab​Value	US	3
(0070,0252)	Pattern Off Color CIELab Value	Pattern​Off​Color​CIE​Lab​Value	US	3
(0070,0253)	Line Thickness	Line​Thickness	FL	1
(0070,0254)	Line Dashing Style	Line​Dashing​Style	CS	1
(0070,0255)	Line Pattern	Line​Pattern	UL	1
(0070,0256)	Fill Pattern	Fill​Pattern	OB	1
(0070,0257)	Fill Mode	Fill​Mode	CS	1
(0070,0258)	Shadow Opacity	Shadow​Opacity	FL	1
(0070,0261)	Gap Length	Gap​Length	FL	1
(0070,0262)	Diameter of Visibility	Diameter​Of​Visibility	FL	1
(0070,0273)	Rotation Point	Rotation​Point	FL	2
(0070,0274)	Tick Alignment	Tick​Alignment	CS	1
(0070,0278)	Show Tick Label	Show​Tick​Label	CS	1
(0070,0279)	Tick Label Alignment	Tick​Label​Alignment	CS	1
(0070,0282)	Compound Graphic Units	Compound​Graphic​Units	CS	1
(0070,0284)	Pattern On Opacity	Pattern​On​Opacity	FL	1
(0070,0285)	Pattern Off Opacity	Pattern​Off​Opacity	FL	1
(0070,0287)	Major Ticks Sequence	Major​Ticks​Sequence	SQ	1
(0070,0288)	Tick Position	Tick​Position	FL	1
(0070,0289)	Tick Label	Tick​Label	SH	1
(0070,0294)	Compound Graphic Type	Compound​Graphic​Type	CS	1
(0070,0295)	Graphic Group ID	Graphic​Group​ID	UL	1
(0070,0306)	Shape Type	Shape​Type	CS	1
(0070,0308)	Registration Sequence	Registration​Sequence	SQ	1
(0070,0309)	Matrix Registration Sequence	Matrix​Registration​Sequence	SQ	1
(0070,030A)	Matrix Sequence	Matrix​Sequence	SQ	1
(0070,030B)	Frame of Reference to Displayed Coordinate System Transformation Matrix	Frame​Of​Reference​To​Displayed​Coordinate​System​Transformation​Matrix	FD	16
(0070,030C)	Frame of Reference Transformation Matrix Type	Frame​Of​Reference​Transformation​Matrix​Type	CS	1
(0070,030D)	Registration Type Code Sequence	Registration​Type​Code​Sequence	SQ	1
(0070,030F)	Fiducial Description	Fiducial​Description	ST	1
(0070,0310)	Fiducial Identifier	Fiducial​Identifier	SH	1
(0070,0311)	Fiducial Identifier Code Sequence	Fiducial​Identifier​Code​Sequence	SQ	1
(0070,0312)	Contour Uncertainty Radius	Contour​Uncertainty​Radius	FD	1
(0070,0314)	Used Fiducials Sequence	Used​Fiducials​Sequence	SQ	1
(0070,0318)	Graphic Coordinates Data Sequence	Graphic​Coordinates​Data​Sequence	SQ	1
(0070,031A)	Fiducial UID	Fiducial​UID	UI	1
(0070,031B)	Referenced Fiducial UID	Referenced​Fiducial​UID	UI	1
(0070,031C)	Fiducial Set Sequence	Fiducial​Set​Sequence	SQ	1
(0070,031E)	Fiducial Sequence	Fiducial​Sequence	SQ	1
(0070,031F)	Fiducials Property Category Code Sequence	Fiducials​Property​Category​Code​Sequence	SQ	1
(0070,0401)	Graphic Layer Recommended Display CIELab Value	Graphic​Layer​Recommended​Display​CIE​Lab​Value	US	3
(0070,0402)	Blending Sequence	Blending​Sequence	SQ	1
(0070,0403)	Relative Opacity	Relative​Opacity	FL	1
(0070,0404)	Referenced Spatial Registration Sequence	Referenced​Spatial​Registration​Sequence	SQ	1
(0070,0405)	Blending Position	Blending​Position	CS	1
(0070,1101)	Presentation Display Collection UID	Presentation​Display​Collection​UID	UI	1
(0070,1102)	Presentation Sequence Collection UID	Presentation​Sequence​Collection​UID	UI	1
(0070,1103)	Presentation Sequence Position Index	Presentation​Sequence​Position​Index	US	1
(0070,1104)	Rendered Image Reference Sequence	Rendered​Image​Reference​Sequence	SQ	1
(0070,1201)	Volumetric Presentation State Input Sequence	Volumetric​Presentation​State​Input​Sequence	SQ	1
(0070,1202)	Presentation Input Type	Presentation​Input​Type	CS	1
(0070,1203)	Input Sequence Position Index	Input​Sequence​Position​Index	US	1
(0070,1204)	Crop	Crop	CS	1
(0070,1205)	Cropping Specification Index	Cropping​Specification​Index	US	1-n
(0070,1206)	Compositing Method	Compositing​Method	CS	1	RET (2017a)
(0070,1207)	Volumetric Presentation Input Number	Volumetric​Presentation​Input​Number	US	1
(0070,1208)	Image Volume Geometry	Image​Volume​Geometry	CS	1
(0070,1209)	Volumetric Presentation Input Set UID	Volumetric​Presentation​Input​Set​UID	UI	1
(0070,120A)	Volumetric Presentation Input Set Sequence	Volumetric​Presentation​Input​Set​Sequence	SQ	1
(0070,120B)	Global Crop	Global​Crop	CS	1
(0070,120C)	Global Cropping Specification Index	Global​Cropping​Specification​Index	US	1-n
(0070,120D)	Rendering Method	Rendering​Method	CS	1
(0070,1301)	Volume Cropping Sequence	Volume​Cropping​Sequence	SQ	1
(0070,1302)	Volume Cropping Method	Volume​Cropping​Method	CS	1
(0070,1303)	Bounding Box Crop	Bounding​Box​Crop	FD	6
(0070,1304)	Oblique Cropping Plane Sequence	Oblique​Cropping​Plane​Sequence	SQ	1
(0070,1305)	Plane	Plane	FD	4
(0070,1306)	Plane Normal	Plane​Normal	FD	3
(0070,1309)	Cropping Specification Number	Cropping​Specification​Number	US	1
(0070,1501)	Multi-Planar Reconstruction Style	Multi​Planar​Reconstruction​Style	CS	1
(0070,1502)	MPR Thickness Type	MPRThickness​Type	CS	1
(0070,1503)	MPR Slab Thickness	MPRSlab​Thickness	FD	1
(0070,1505)	MPR Top Left Hand Corner	MPRTop​Left​Hand​Corner	FD	3
(0070,1507)	MPR View Width Direction	MPRView​Width​Direction	FD	3
(0070,1508)	MPR View Width	MPRView​Width	FD	1
(0070,150C)	Number of Volumetric Curve Points	Number​Of​Volumetric​Curve​Points	UL	1
(0070,150D)	Volumetric Curve Points	Volumetric​Curve​Points	OD	1
(0070,1511)	MPR View Height Direction	MPRView​Height​Direction	FD	3
(0070,1512)	MPR View Height	MPRView​Height	FD	1
(0070,1602)	Render Projection	Render​Projection	CS	1
(0070,1603)	Viewpoint Position	Viewpoint​Position	FD	3
(0070,1604)	Viewpoint LookAt Point	Viewpoint​LookAt​Point	FD	3
(0070,1605)	Viewpoint Up Direction	Viewpoint​Up​Direction	FD	3
(0070,1606)	Render Field of View	Render​Field​Of​View	FD	6
(0070,1607)	Sampling Step Size	Sampling​Step​Size	FD	1
(0070,1701)	Shading Style	Shading​Style	CS	1
(0070,1702)	Ambient Reflection Intensity	Ambient​Reflection​Intensity	FD	1
(0070,1703)	Light Direction	Light​Direction	FD	3
(0070,1704)	Diffuse Reflection Intensity	Diffuse​Reflection​Intensity	FD	1
(0070,1705)	Specular Reflection Intensity	Specular​Reflection​Intensity	FD	1
(0070,1706)	Shininess	Shininess	FD	1
(0070,1801)	Presentation State Classification Component Sequence	Presentation​State​Classification​Component​Sequence	SQ	1
(0070,1802)	Component Type	Component​Type	CS	1
(0070,1803)	Component Input Sequence	Component​Input​Sequence	SQ	1
(0070,1804)	Volumetric Presentation Input Index	Volumetric​Presentation​Input​Index	US	1
(0070,1805)	Presentation State Compositor Component Sequence	Presentation​State​Compositor​Component​Sequence	SQ	1
(0070,1806)	Weighting Transfer Function Sequence	Weighting​Transfer​Function​Sequence	SQ	1
(0070,1807)	Weighting Lookup Table Descriptor	Weighting​Lookup​Table​Descriptor	US	3
(0070,1808)	Weighting Lookup Table Data	Weighting​Lookup​Table​Data	OB	1
(0070,1901)	Volumetric Annotation Sequence	Volumetric​Annotation​Sequence	SQ	1
(0070,1903)	Referenced Structured Context Sequence	Referenced​Structured​Context​Sequence	SQ	1
(0070,1904)	Referenced Content Item	Referenced​Content​Item	UI	1
(0070,1905)	Volumetric Presentation Input Annotation Sequence	Volumetric​Presentation​Input​Annotation​Sequence	SQ	1
(0070,1907)	Annotation Clipping	Annotation​Clipping	CS	1
(0070,1A01)	Presentation Animation Style	Presentation​Animation​Style	CS	1
(0070,1A03)	Recommended Animation Rate	Recommended​Animation​Rate	FD	1
(0070,1A04)	Animation Curve Sequence	Animation​Curve​Sequence	SQ	1
(0070,1A05)	Animation Step Size	Animation​Step​Size	FD	1
(0070,1A06)	Swivel Range	Swivel​Range	FD	1
(0070,1A07)	Volumetric Curve Up Directions	Volumetric​Curve​Up​Directions	OD	1
(0070,1A08)	Volume Stream Sequence	Volume​Stream​Sequence	SQ	1
(0070,1A09)	RGBA Transfer Function Description	RGBA​Transfer​Function​Description	LO	1
(0070,1B01)	Advanced Blending Sequence	Advanced​Blending​Sequence	SQ	1
(0070,1B02)	Blending Input Number	Blending​Input​Number	US	1
(0070,1B03)	Blending Display Input Sequence	Blending​Display​Input​Sequence	SQ	1
(0070,1B04)	Blending Display Sequence	Blending​Display​Sequence	SQ	1
(0070,1B06)	Blending Mode	Blending​Mode	CS	1
(0070,1B07)	Time Series Blending	Time​Series​Blending	CS	1
(0070,1B08)	Geometry for Display	Geometry​For​Display	CS	1
(0070,1B11)	Threshold Sequence	Threshold​Sequence	SQ	1
(0070,1B12)	Threshold Value Sequence	Threshold​Value​Sequence	SQ	1
(0070,1B13)	Threshold Type	Threshold​Type	CS	1
(0070,1B14)	Threshold Value	Threshold​Value	FD	1
(0072,0002)	Hanging Protocol Name	Hanging​Protocol​Name	SH	1
(0072,0004)	Hanging Protocol Description	Hanging​Protocol​Description	LO	1
(0072,0006)	Hanging Protocol Level	Hanging​Protocol​Level	CS	1
(0072,0008)	Hanging Protocol Creator	Hanging​Protocol​Creator	LO	1
(0072,000A)	Hanging Protocol Creation Date​Time	Hanging​Protocol​Creation​Date​Time	DT	1
(0072,000C)	Hanging Protocol Definition Sequence	Hanging​Protocol​Definition​Sequence	SQ	1
(0072,000E)	Hanging Protocol User Identification Code Sequence	Hanging​Protocol​User​Identification​Code​Sequence	SQ	1
(0072,0010)	Hanging Protocol User Group Name	Hanging​Protocol​User​Group​Name	LO	1
(0072,0012)	Source Hanging Protocol Sequence	Source​Hanging​Protocol​Sequence	SQ	1
(0072,0014)	Number of Priors Referenced	Number​Of​Priors​Referenced	US	1
(0072,0020)	Image Sets Sequence	Image​Sets​Sequence	SQ	1
(0072,0022)	Image Set Selector Sequence	Image​Set​Selector​Sequence	SQ	1
(0072,0024)	Image Set Selector Usage Flag	Image​Set​Selector​Usage​Flag	CS	1
(0072,0026)	Selector Attribute	Selector​Attribute	AT	1
(0072,0028)	Selector Value Number	Selector​Value​Number	US	1
(0072,0030)	Time Based Image Sets Sequence	Time​Based​Image​Sets​Sequence	SQ	1
(0072,0032)	Image Set Number	Image​Set​Number	US	1
(0072,0034)	Image Set Selector Category	Image​Set​Selector​Category	CS	1
(0072,0038)	Relative Time	Relative​Time	US	2
(0072,003A)	Relative Time Units	Relative​Time​Units	CS	1
(0072,003C)	Abstract Prior Value	Abstract​Prior​Value	SS	2
(0072,003E)	Abstract Prior Code Sequence	Abstract​Prior​Code​Sequence	SQ	1
(0072,0040)	Image Set Label	Image​Set​Label	LO	1
(0072,0050)	Selector Attribute VR	Selector​Attribute​VR	CS	1
(0072,0052)	Selector Sequence Pointer	Selector​Sequence​Pointer	AT	1-n
(0072,0054)	Selector Sequence Pointer Private Creator	Selector​Sequence​Pointer​Private​Creator	LO	1-n
(0072,0056)	Selector Attribute Private Creator	Selector​Attribute​Private​Creator	LO	1
(0072,005E)	Selector AE Value	Selector​AEValue	AE	1-n
(0072,005F)	Selector AS Value	Selector​ASValue	AS	1-n
(0072,0060)	Selector AT Value	Selector​AT​Value	AT	1-n
(0072,0061)	Selector DA Value	Selector​DAValue	DA	1-n
(0072,0062)	Selector CS Value	Selector​CS​Value	CS	1-n
(0072,0063)	Selector DT Value	Selector​DTValue	DT	1-n
(0072,0064)	Selector IS Value	Selector​IS​Value	IS	1-n
(0072,0065)	Selector OB Value	Selector​OBValue	OB	1
(0072,0066)	Selector LO Value	Selector​LO​Value	LO	1-n
(0072,0067)	Selector OF Value	Selector​OFValue	OF	1
(0072,0068)	Selector LT Value	Selector​LT​Value	LT	1
(0072,0069)	Selector OW Value	Selector​OWValue	OW	1
(0072,006A)	Selector PN Value	Selector​PN​Value	PN	1-n
(0072,006B)	Selector TM Value	Selector​TMValue	TM	1-n
(0072,006C)	Selector SH Value	Selector​SH​Value	SH	1-n
(0072,006D)	Selector UN Value	Selector​UNValue	UN	1
(0072,006E)	Selector ST Value	Selector​ST​Value	ST	1
(0072,006F)	Selector UC Value	Selector​UCValue	UC	1-n
(0072,0070)	Selector UT Value	Selector​UT​Value	UT	1
(0072,0071)	Selector UR Value	Selector​URValue	UR	1
(0072,0072)	Selector DS Value	Selector​DS​Value	DS	1-n
(0072,0073)	Selector OD Value	Selector​ODValue	OD	1
(0072,0074)	Selector FD Value	Selector​FD​Value	FD	1-n
(0072,0075)	Selector OL Value	Selector​OLValue	OL	1
(0072,0076)	Selector FL Value	Selector​FL​Value	FL	1-n
(0072,0078)	Selector UL Value	Selector​UL​Value	UL	1-n
(0072,007A)	Selector US Value	Selector​US​Value	US	1-n
(0072,007C)	Selector SL Value	Selector​SL​Value	SL	1-n
(0072,007E)	Selector SS Value	Selector​SS​Value	SS	1-n
(0072,007F)	Selector UI Value	Selector​UI​Value	UI	1-n
(0072,0080)	Selector Code Sequence Value	Selector​Code​Sequence​Value	SQ	1
(0072,0100)	Number of Screens	Number​Of​Screens	US	1
(0072,0102)	Nominal Screen Definition Sequence	Nominal​Screen​Definition​Sequence	SQ	1
(0072,0104)	Number of Vertical Pixels	Number​Of​Vertical​Pixels	US	1
(0072,0106)	Number of Horizontal Pixels	Number​Of​Horizontal​Pixels	US	1
(0072,0108)	Display Environment Spatial Position	Display​Environment​Spatial​Position	FD	4
(0072,010A)	Screen Minimum Grayscale Bit Depth	Screen​Minimum​Grayscale​Bit​Depth	US	1
(0072,010C)	Screen Minimum Color Bit Depth	Screen​Minimum​Color​Bit​Depth	US	1
(0072,010E)	Application Maximum Repaint Time	Application​Maximum​Repaint​Time	US	1
(0072,0200)	Display Sets Sequence	Display​Sets​Sequence	SQ	1
(0072,0202)	Display Set Number	Display​Set​Number	US	1
(0072,0203)	Display Set Label	Display​Set​Label	LO	1
(0072,0204)	Display Set Presentation Group	Display​Set​Presentation​Group	US	1
(0072,0206)	Display Set Presentation Group Description	Display​Set​Presentation​Group​Description	LO	1
(0072,0208)	Partial Data Display Handling	Partial​Data​Display​Handling	CS	1
(0072,0210)	Synchronized Scrolling Sequence	Synchronized​Scrolling​Sequence	SQ	1
(0072,0212)	Display Set Scrolling Group	Display​Set​Scrolling​Group	US	2-n
(0072,0214)	Navigation Indicator Sequence	Navigation​Indicator​Sequence	SQ	1
(0072,0216)	Navigation Display Set	Navigation​Display​Set	US	1
(0072,0218)	Reference Display Sets	Reference​Display​Sets	US	1-n
(0072,0300)	Image Boxes Sequence	Image​Boxes​Sequence	SQ	1
(0072,0302)	Image Box Number	Image​Box​Number	US	1
(0072,0304)	Image Box Layout Type	Image​Box​Layout​Type	CS	1
(0072,0306)	Image Box Tile Horizontal Dimension	Image​Box​Tile​Horizontal​Dimension	US	1
(0072,0308)	Image Box Tile Vertical Dimension	Image​Box​Tile​Vertical​Dimension	US	1
(0072,0310)	Image Box Scroll Direction	Image​Box​Scroll​Direction	CS	1
(0072,0312)	Image Box Small Scroll Type	Image​Box​Small​Scroll​Type	CS	1
(0072,0314)	Image Box Small Scroll Amount	Image​Box​Small​Scroll​Amount	US	1
(0072,0316)	Image Box Large Scroll Type	Image​Box​Large​Scroll​Type	CS	1
(0072,0318)	Image Box Large Scroll Amount	Image​Box​Large​Scroll​Amount	US	1
(0072,0320)	Image Box Overlap Priority	Image​Box​Overlap​Priority	US	1
(0072,0330)	Cine Relative to Real-Time	Cine​Relative​To​Real​Time	FD	1
(0072,0400)	Filter Operations Sequence	Filter​Operations​Sequence	SQ	1
(0072,0402)	Filter-by Category	Filter​By​Category	CS	1
(0072,0404)	Filter-by Attribute Presence	Filter​By​Attribute​Presence	CS	1
(0072,0406)	Filter-by Operator	Filter​By​Operator	CS	1
(0072,0420)	Structured Display Background CIELab Value	Structured​Display​Background​CIE​Lab​Value	US	3
(0072,0421)	Empty Image Box CIELab Value	Empty​Image​Box​CIE​Lab​Value	US	3
(0072,0422)	Structured Display Image Box Sequence	Structured​Display​Image​Box​Sequence	SQ	1
(0072,0424)	Structured Display Text Box Sequence	Structured​Display​Text​Box​Sequence	SQ	1
(0072,0427)	Referenced First Frame Sequence	Referenced​First​Frame​Sequence	SQ	1
(0072,0430)	Image Box Synchronization Sequence	Image​Box​Synchronization​Sequence	SQ	1
(0072,0432)	Synchronized Image Box List	Synchronized​Image​Box​List	US	2-n
(0072,0434)	Type of Synchronization	Type​Of​Synchronization	CS	1
(0072,0500)	Blending Operation Type	Blending​Operation​Type	CS	1
(0072,0510)	Reformatting Operation Type	Reformatting​Operation​Type	CS	1
(0072,0512)	Reformatting Thickness	Reformatting​Thickness	FD	1
(0072,0514)	Reformatting Interval	Reformatting​Interval	FD	1
(0072,0516)	Reformatting Operation Initial View Direction	Reformatting​Operation​Initial​View​Direction	CS	1
(0072,0520)	3D Rendering Type	Three​D​Rendering​Type	CS	1-n
(0072,0600)	Sorting Operations Sequence	Sorting​Operations​Sequence	SQ	1
(0072,0602)	Sort-by Category	Sort​By​Category	CS	1
(0072,0604)	Sorting Direction	Sorting​Direction	CS	1
(0072,0700)	Display Set Patient Orientation	Display​Set​Patient​Orientation	CS	2
(0072,0702)	VOI Type	VOI​Type	CS	1
(0072,0704)	Pseudo-Color Type	Pseudo​Color​Type	CS	1
(0072,0705)	Pseudo-Color Palette Instance Reference Sequence	Pseudo​Color​Palette​Instance​Reference​Sequence	SQ	1
(0072,0706)	Show Grayscale Inverted	Show​Grayscale​Inverted	CS	1
(0072,0710)	Show Image True Size Flag	Show​Image​True​Size​Flag	CS	1
(0072,0712)	Show Graphic Annotation Flag	Show​Graphic​Annotation​Flag	CS	1
(0072,0714)	Show Patient Demographics Flag	Show​Patient​Demographics​Flag	CS	1
(0072,0716)	Show Acquisition Techniques Flag	Show​Acquisition​Techniques​Flag	CS	1
(0072,0717)	Display Set Horizontal Justification	Display​Set​Horizontal​Justification	CS	1
(0072,0718)	Display Set Vertical Justification	Display​Set​Vertical​Justification	CS	1
(0074,0120)	Continuation Start Meterset	Continuation​Start​Meterset	FD	1
(0074,0121)	Continuation End Meterset	Continuation​End​Meterset	FD	1
(0074,1000)	Procedure Step State	Procedure​Step​State	CS	1
(0074,1002)	Procedure Step Progress Information Sequence	Procedure​Step​Progress​Information​Sequence	SQ	1
(0074,1004)	Procedure Step Progress	Procedure​Step​Progress	DS	1
(0074,1006)	Procedure Step Progress Description	Procedure​Step​Progress​Description	ST	1
(0074,1007)	Procedure Step Progress Parameters Sequence	Procedure​Step​Progress​Parameters​Sequence	SQ	1
(0074,1008)	Procedure Step Communications URI Sequence	Procedure​Step​Communications​URI​Sequence	SQ	1
(0074,100A)	Contact URI	Contact​URI	UR	1
(0074,100C)	Contact Display Name	Contact​Display​Name	LO	1
(0074,100E)	Procedure Step Discontinuation Reason Code Sequence	Procedure​Step​Discontinuation​Reason​Code​Sequence	SQ	1
(0074,1020)	Beam Task Sequence	Beam​Task​Sequence	SQ	1
(0074,1022)	Beam Task Type	Beam​Task​Type	CS	1
(0074,1024)	Beam Order Index (Trial)	Beam​Order​Index​Trial	IS	1	RET (2009)
(0074,1025)	Autosequence Flag	Autosequence​Flag	CS	1
(0074,1026)	Table Top Vertical Adjusted Position	Table​Top​Vertical​Adjusted​Position	FD	1
(0074,1027)	Table Top Longitudinal Adjusted Position	Table​Top​Longitudinal​Adjusted​Position	FD	1
(0074,1028)	Table Top Lateral Adjusted Position	Table​Top​Lateral​Adjusted​Position	FD	1
(0074,102A)	Patient Support Adjusted Angle	Patient​Support​Adjusted​Angle	FD	1
(0074,102B)	Table Top Eccentric Adjusted Angle	Table​Top​Eccentric​Adjusted​Angle	FD	1
(0074,102C)	Table Top Pitch Adjusted Angle	Table​Top​Pitch​Adjusted​Angle	FD	1
(0074,102D)	Table Top Roll Adjusted Angle	Table​Top​Roll​Adjusted​Angle	FD	1
(0074,1030)	Delivery Verification Image Sequence	Delivery​Verification​Image​Sequence	SQ	1
(0074,1032)	Verification Image Timing	Verification​Image​Timing	CS	1
(0074,1034)	Double Exposure Flag	Double​Exposure​Flag	CS	1
(0074,1036)	Double Exposure Ordering	Double​Exposure​Ordering	CS	1
(0074,1038)	Double Exposure Meterset (Trial)	Double​Exposure​Meterset​Trial	DS	1	RET (2009)
(0074,103A)	Double Exposure Field Delta (Trial)	Double​Exposure​Field​Delta​Trial	DS	4	RET (2009)
(0074,1040)	Related Reference RT Image Sequence	Related​Reference​RT​Image​Sequence	SQ	1
(0074,1042)	General Machine Verification Sequence	General​Machine​Verification​Sequence	SQ	1
(0074,1044)	Conventional Machine Verification Sequence	Conventional​Machine​Verification​Sequence	SQ	1
(0074,1046)	Ion Machine Verification Sequence	Ion​Machine​Verification​Sequence	SQ	1
(0074,1048)	Failed Attributes Sequence	Failed​Attributes​Sequence	SQ	1
(0074,104A)	Overridden Attributes Sequence	Overridden​Attributes​Sequence	SQ	1
(0074,104C)	Conventional Control Point Verification Sequence	Conventional​Control​Point​Verification​Sequence	SQ	1
(0074,104E)	Ion Control Point Verification Sequence	Ion​Control​Point​Verification​Sequence	SQ	1
(0074,1050)	Attribute Occurrence Sequence	Attribute​Occurrence​Sequence	SQ	1
(0074,1052)	Attribute Occurrence Pointer	Attribute​Occurrence​Pointer	AT	1
(0074,1054)	Attribute Item Selector	Attribute​Item​Selector	UL	1
(0074,1056)	Attribute Occurrence Private Creator	Attribute​Occurrence​Private​Creator	LO	1
(0074,1057)	Selector Sequence Pointer Items	Selector​Sequence​Pointer​Items	IS	1-n
(0074,1200)	Scheduled Procedure Step Priority	Scheduled​Procedure​Step​Priority	CS	1
(0074,1202)	Worklist Label	Worklist​Label	LO	1
(0074,1204)	Procedure Step Label	Procedure​Step​Label	LO	1
(0074,1210)	Scheduled Processing Parameters Sequence	Scheduled​Processing​Parameters​Sequence	SQ	1
(0074,1212)	Performed Processing Parameters Sequence	Performed​Processing​Parameters​Sequence	SQ	1
(0074,1216)	Unified Procedure Step Performed Procedure Sequence	Unified​Procedure​Step​Performed​Procedure​Sequence	SQ	1
(0074,1220)	Related Procedure Step Sequence	Related​Procedure​Step​Sequence	SQ	1	RET (2009)
(0074,1222)	Procedure Step Relationship Type	Procedure​Step​Relationship​Type	LO	1	RET (2009)
(0074,1224)	Replaced Procedure Step Sequence	Replaced​Procedure​Step​Sequence	SQ	1
(0074,1230)	Deletion Lock	Deletion​Lock	LO	1
(0074,1234)	Receiving AE	Receiving​AE	AE	1
(0074,1236)	Requesting AE	Requesting​AE	AE	1
(0074,1238)	Reason for Cancellation	Reason​For​Cancellation	LT	1
(0074,1242)	SCP Status	SCP​Status	CS	1
(0074,1244)	Subscription List Status	Subscription​List​Status	CS	1
(0074,1246)	Unified Procedure Step List Status	Unified​Procedure​Step​List​Status	CS	1
(0074,1324)	Beam Order Index	Beam​Order​Index	UL	1
(0074,1338)	Double Exposure Meterset	Double​Exposure​Meterset	FD	1
(0074,133A)	Double Exposure Field Delta	Double​Exposure​Field​Delta	FD	4
(0074,1401)	Brachy Task Sequence	Brachy​Task​Sequence	SQ	1
(0074,1402)	Continuation Start Total Reference Air Kerma	Continuation​Start​Total​Reference​Air​Kerma	DS	1
(0074,1403)	Continuation End Total Reference Air Kerma	Continuation​End​Total​Reference​Air​Kerma	DS	1
(0074,1404)	Continuation Pulse Number	Continuation​Pulse​Number	IS	1
(0074,1405)	Channel Delivery Order Sequence	Channel​Delivery​Order​Sequence	SQ	1
(0074,1406)	Referenced Channel Number	Referenced​Channel​Number	IS	1
(0074,1407)	Start Cumulative Time Weight	Start​Cumulative​Time​Weight	DS	1
(0074,1408)	End Cumulative Time Weight	End​Cumulative​Time​Weight	DS	1
(0074,1409)	Omitted Channel Sequence	Omitted​Channel​Sequence	SQ	1
(0074,140A)	Reason for Channel Omission	Reason​For​Channel​Omission	CS	1
(0074,140B)	Reason for Channel Omission Description	Reason​For​Channel​Omission​Description	LO	1
(0074,140C)	Channel Delivery Order Index	Channel​Delivery​Order​Index	IS	1
(0074,140D)	Channel Delivery Continuation Sequence	Channel​Delivery​Continuation​Sequence	SQ	1
(0074,140E)	Omitted Application Setup Sequence	Omitted​Application​Setup​Sequence	SQ	1
(0076,0001)	Implant Assembly Template Name	Implant​Assembly​Template​Name	LO	1
(0076,0003)	Implant Assembly Template Issuer	Implant​Assembly​Template​Issuer	LO	1
(0076,0006)	Implant Assembly Template Version	Implant​Assembly​Template​Version	LO	1
(0076,0008)	Replaced Implant Assembly Template Sequence	Replaced​Implant​Assembly​Template​Sequence	SQ	1
(0076,000A)	Implant Assembly Template Type	Implant​Assembly​Template​Type	CS	1
(0076,000C)	Original Implant Assembly Template Sequence	Original​Implant​Assembly​Template​Sequence	SQ	1
(0076,000E)	Derivation Implant Assembly Template Sequence	Derivation​Implant​Assembly​Template​Sequence	SQ	1
(0076,0010)	Implant Assembly Template Target Anatomy Sequence	Implant​Assembly​Template​Target​Anatomy​Sequence	SQ	1
(0076,0020)	Procedure Type Code Sequence	Procedure​Type​Code​Sequence	SQ	1
(0076,0030)	Surgical Technique	Surgical​Technique	LO	1
(0076,0032)	Component Types Sequence	Component​Types​Sequence	SQ	1
(0076,0034)	Component Type Code Sequence	Component​Type​Code​Sequence	SQ	1
(0076,0036)	Exclusive Component Type	Exclusive​Component​Type	CS	1
(0076,0038)	Mandatory Component Type	Mandatory​Component​Type	CS	1
(0076,0040)	Component Sequence	Component​Sequence	SQ	1
(0076,0055)	Component ID	Component​ID	US	1
(0076,0060)	Component Assembly Sequence	Component​Assembly​Sequence	SQ	1
(0076,0070)	Component 1 Referenced ID	Component1Referenced​ID	US	1
(0076,0080)	Component 1 Referenced Mating Feature Set ID	Component1Referenced​Mating​Feature​Set​ID	US	1
(0076,0090)	Component 1 Referenced Mating Feature ID	Component1Referenced​Mating​Feature​ID	US	1
(0076,00A0)	Component 2 Referenced ID	Component2Referenced​ID	US	1
(0076,00B0)	Component 2 Referenced Mating Feature Set ID	Component2Referenced​Mating​Feature​Set​ID	US	1
(0076,00C0)	Component 2 Referenced Mating Feature ID	Component2Referenced​Mating​Feature​ID	US	1
(0078,0001)	Implant Template Group Name	Implant​Template​Group​Name	LO	1
(0078,0010)	Implant Template Group Description	Implant​Template​Group​Description	ST	1
(0078,0020)	Implant Template Group Issuer	Implant​Template​Group​Issuer	LO	1
(0078,0024)	Implant Template Group Version	Implant​Template​Group​Version	LO	1
(0078,0026)	Replaced Implant Template Group Sequence	Replaced​Implant​Template​Group​Sequence	SQ	1
(0078,0028)	Implant Template Group Target Anatomy Sequence	Implant​Template​Group​Target​Anatomy​Sequence	SQ	1
(0078,002A)	Implant Template Group Members Sequence	Implant​Template​Group​Members​Sequence	SQ	1
(0078,002E)	Implant Template Group Member ID	Implant​Template​Group​Member​ID	US	1
(0078,0050)	3D Implant Template Group Member Matching Point	Three​D​Implant​Template​Group​Member​Matching​Point	FD	3
(0078,0060)	3D Implant Template Group Member Matching Axes	Three​D​Implant​Template​Group​Member​Matching​Axes	FD	9
(0078,0070)	Implant Template Group Member Matching 2D Coordinates Sequence	Implant​Template​Group​Member​Matching2D​Coordinates​Sequence	SQ	1
(0078,0090)	2D Implant Template Group Member Matching Point	Two​D​Implant​Template​Group​Member​Matching​Point	FD	2
(0078,00A0)	2D Implant Template Group Member Matching Axes	Two​D​Implant​Template​Group​Member​Matching​Axes	FD	4
(0078,00B0)	Implant Template Group Variation Dimension Sequence	Implant​Template​Group​Variation​Dimension​Sequence	SQ	1
(0078,00B2)	Implant Template Group Variation Dimension Name	Implant​Template​Group​Variation​Dimension​Name	LO	1
(0078,00B4)	Implant Template Group Variation Dimension Rank Sequence	Implant​Template​Group​Variation​Dimension​Rank​Sequence	SQ	1
(0078,00B6)	Referenced Implant Template Group Member ID	Referenced​Implant​Template​Group​Member​ID	US	1
(0078,00B8)	Implant Template Group Variation Dimension Rank	Implant​Template​Group​Variation​Dimension​Rank	US	1
(0080,0001)	Surface Scan Acquisition Type Code Sequence	Surface​Scan​Acquisition​Type​Code​Sequence	SQ	1
(0080,0002)	Surface Scan Mode Code Sequence	Surface​Scan​Mode​Code​Sequence	SQ	1
(0080,0003)	Registration Method Code Sequence	Registration​Method​Code​Sequence	SQ	1
(0080,0004)	Shot Duration Time	Shot​Duration​Time	FD	1
(0080,0005)	Shot Offset Time	Shot​Offset​Time	FD	1
(0080,0006)	Surface Point Presentation Value Data	Surface​Point​Presentation​Value​Data	US	1-n
(0080,0007)	Surface Point Color CIELab Value Data	Surface​Point​Color​CIE​Lab​Value​Data	US	3-3n
(0080,0008)	UV Mapping Sequence	UV​Mapping​Sequence	SQ	1
(0080,0009)	Texture Label	Texture​Label	SH	1
(0080,0010)	U Value Data	U​Value​Data	OF	1
(0080,0011)	V Value Data	V​Value​Data	OF	1
(0080,0012)	Referenced Texture Sequence	Referenced​Texture​Sequence	SQ	1
(0080,0013)	Referenced Surface Data Sequence	Referenced​Surface​Data​Sequence	SQ	1
(0082,0001)	Assessment Summary	Assessment​Summary	CS	1
(0082,0003)	Assessment Summary Description	Assessment​Summary​Description	UT	1
(0082,0004)	Assessed SOP Instance Sequence	Assessed​SOPInstance​Sequence	SQ	1
(0082,0005)	Referenced Comparison SOP Instance Sequence	Referenced​Comparison​SOPInstance​Sequence	SQ	1
(0082,0006)	Number of Assessment Observations	Number​Of​Assessment​Observations	UL	1
(0082,0007)	Assessment Observations Sequence	Assessment​Observations​Sequence	SQ	1
(0082,0008)	Observation Significance	Observation​Significance	CS	1
(0082,000A)	Observation Description	Observation​Description	UT	1
(0082,000C)	Structured Constraint Observation Sequence	Structured​Constraint​Observation​Sequence	SQ	1
(0082,0010)	Assessed Attribute Value Sequence	Assessed​Attribute​Value​Sequence	SQ	1
(0082,0016)	Assessment Set ID	Assessment​Set​ID	LO	1
(0082,0017)	Assessment Requester Sequence	Assessment​Requester​Sequence	SQ	1
(0082,0018)	Selector Attribute Name	Selector​Attribute​Name	LO	1
(0082,0019)	Selector Attribute Keyword	Selector​Attribute​Keyword	LO	1
(0082,0021)	Assessment Type Code Sequence	Assessment​Type​Code​Sequence	SQ	1
(0082,0022)	Observation Basis Code Sequence	Observation​Basis​Code​Sequence	SQ	1
(0082,0023)	Assessment Label	Assessment​Label	LO	1
(0082,0032)	Constraint Type	Constraint​Type	CS	1
(0082,0033)	Specification Selection Guidance	Specification​Selection​Guidance	UT	1
(0082,0034)	Constraint Value Sequence	Constraint​Value​Sequence	SQ	1
(0082,0035)	Recommended Default Value Sequence	Recommended​Default​Value​Sequence	SQ	1
(0082,0036)	Constraint Violation Significance	Constraint​Violation​Significance	CS	1
(0082,0037)	Constraint Violation Condition	Constraint​Violation​Condition	UT	1
(0082,0038)	Modifiable Constraint Flag	Modifiable​Constraint​Flag	CS	1
(0088,0130)	Storage Media File-set ID	Storage​Media​File​Set​ID	SH	1
(0088,0140)	Storage Media File-set UID	Storage​Media​File​Set​UID	UI	1
(0088,0200)	Icon Image Sequence	Icon​Image​Sequence	SQ	1
(0088,0904)	Topic Title	Topic​Title	LO	1	RET (2007)
(0088,0906)	Topic Subject	Topic​Subject	ST	1	RET (2007)
(0088,0910)	Topic Author	Topic​Author	LO	1	RET (2007)
(0088,0912)	Topic Keywords	Topic​Keywords	LO	1-32	RET (2007)
(0100,0410)	SOP Instance Status	SOP​Instance​Status	CS	1
(0100,0420)	SOP Authorization DateTime	SOP​Authorization​Date​Time	DT	1
(0100,0424)	SOP Authorization Comment	SOP​Authorization​Comment	LT	1
(0100,0426)	Authorization Equipment Certification Number	Authorization​Equipment​Certification​Number	LO	1
(0400,0005)	MAC ID Number	MACID​Number	US	1
(0400,0010)	MAC Calculation Transfer Syntax UID	MAC​Calculation​Transfer​Syntax​UID	UI	1
(0400,0015)	MAC Algorithm	MAC​Algorithm	CS	1
(0400,0020)	Data Elements Signed	Data​Elements​Signed	AT	1-n
(0400,0100)	Digital Signature UID	Digital​Signature​UID	UI	1
(0400,0105)	Digital Signature DateTime	Digital​Signature​Date​Time	DT	1
(0400,0110)	Certificate Type	Certificate​Type	CS	1
(0400,0115)	Certificate of Signer	Certificate​Of​Signer	OB	1
(0400,0120)	Signature	Signature	OB	1
(0400,0305)	Certified Timestamp Type	Certified​Timestamp​Type	CS	1
(0400,0310)	Certified Timestamp	Certified​Timestamp	OB	1
(0400,0315)			FL	1	RET (2015c)
(0400,0401)	Digital Signature Purpose Code Sequence	Digital​Signature​Purpose​Code​Sequence	SQ	1
(0400,0402)	Referenced Digital Signature Sequence	Referenced​Digital​Signature​Sequence	SQ	1
(0400,0403)	Referenced SOP Instance MAC Sequence	Referenced​SOP​Instance​MAC​Sequence	SQ	1
(0400,0404)	MAC	MAC	OB	1
(0400,0500)	Encrypted Attributes Sequence	Encrypted​Attributes​Sequence	SQ	1
(0400,0510)	Encrypted Content Transfer Syntax UID	Encrypted​Content​Transfer​Syntax​UID	UI	1
(0400,0520)	Encrypted Content	Encrypted​Content	OB	1
(0400,0550)	Modified Attributes Sequence	Modified​Attributes​Sequence	SQ	1
(0400,0561)	Original Attributes Sequence	Original​Attributes​Sequence	SQ	1
(0400,0562)	Attribute Modification DateTime	Attribute​Modification​Date​Time	DT	1
(0400,0563)	Modifying System	Modifying​System	LO	1
(0400,0564)	Source of Previous Values	Source​Of​Previous​Values	LO	1
(0400,0565)	Reason for the Attribute Modification	Reason​For​The​Attribute​Modification	CS	1
(0400,0600)	Instance Origin Status	Instance​Origin​Status	CS	1
(1000,xxx0)	Escape Triplet	Escape​Triplet	US	3	RET (2007)
(1000,xxx1)	Run Length Triplet	Run​Length​Triplet	US	3	RET (2007)
(1000,xxx2)	Huffman Table Size	Huffman​Table​Size	US	1	RET (2007)
(1000,xxx3)	Huffman Table Triplet	Huffman​Table​Triplet	US	3	RET (2007)
(1000,xxx4)	Shift Table Size	Shift​Table​Size	US	1	RET (2007)
(1000,xxx5)	Shift Table Triplet	Shift​Table​Triplet	US	3	RET (2007)
(1010,xxxx)	Zonal Map	Zonal​Map	US	1-n	RET (2007)
(2000,0010)	Number of Copies	Number​Of​Copies	IS	1
(2000,001E)	Printer Configuration Sequence	Printer​Configuration​Sequence	SQ	1
(2000,0020)	Print Priority	Print​Priority	CS	1
(2000,0030)	Medium Type	Medium​Type	CS	1
(2000,0040)	Film Destination	Film​Destination	CS	1
(2000,0050)	Film Session Label	Film​Session​Label	LO	1
(2000,0060)	Memory Allocation	Memory​Allocation	IS	1
(2000,0061)	Maximum Memory Allocation	Maximum​Memory​Allocation	IS	1
(2000,0062)	Color Image Printing Flag	Color​Image​Printing​Flag	CS	1	RET (2004)
(2000,0063)	Collation Flag	Collation​Flag	CS	1	RET (2004)
(2000,0065)	Annotation Flag	Annotation​Flag	CS	1	RET (2004)
(2000,0067)	Image Overlay Flag	Image​Overlay​Flag	CS	1	RET (2004)
(2000,0069)	Presentation LUT Flag	Presentation​LUT​Flag	CS	1	RET (2004)
(2000,006A)	Image Box Presentation LUT Flag	Image​Box​Presentation​LUT​Flag	CS	1	RET (2004)
(2000,00A0)	Memory Bit Depth	Memory​Bit​Depth	US	1
(2000,00A1)	Printing Bit Depth	Printing​Bit​Depth	US	1
(2000,00A2)	Media Installed Sequence	Media​Installed​Sequence	SQ	1
(2000,00A4)	Other Media Available Sequence	Other​Media​Available​Sequence	SQ	1
(2000,00A8)	Supported Image Display Formats Sequence	Supported​Image​Display​Formats​Sequence	SQ	1
(2000,0500)	Referenced Film Box Sequence	Referenced​Film​Box​Sequence	SQ	1
(2000,0510)	Referenced Stored Print Sequence	Referenced​Stored​Print​Sequence	SQ	1	RET (2007)
(2010,0010)	Image Display Format	Image​Display​Format	ST	1
(2010,0030)	Annotation Display Format ID	Annotation​Display​Format​ID	CS	1
(2010,0040)	Film Orientation	Film​Orientation	CS	1
(2010,0050)	Film Size ID	Film​Size​ID	CS	1
(2010,0052)	Printer Resolution ID	Printer​Resolution​ID	CS	1
(2010,0054)	Default Printer Resolution ID	Default​Printer​Resolution​ID	CS	1
(2010,0060)	Magnification Type	Magnification​Type	CS	1
(2010,0080)	Smoothing Type	Smoothing​Type	CS	1
(2010,00A6)	Default Magnification Type	Default​Magnification​Type	CS	1
(2010,00A7)	Other Magnification Types Available	Other​Magnification​Types​Available	CS	1-n
(2010,00A8)	Default Smoothing Type	Default​Smoothing​Type	CS	1
(2010,00A9)	Other Smoothing Types Available	Other​Smoothing​Types​Available	CS	1-n
(2010,0100)	Border Density	Border​Density	CS	1
(2010,0110)	Empty Image Density	Empty​Image​Density	CS	1
(2010,0120)	Min Density	Min​Density	US	1
(2010,0130)	Max Density	Max​Density	US	1
(2010,0140)	Trim	Trim	CS	1
(2010,0150)	Configuration Information	Configuration​Information	ST	1
(2010,0152)	Configuration Information Description	Configuration​Information​Description	LT	1
(2010,0154)	Maximum Collated Films	Maximum​Collated​Films	IS	1
(2010,015E)	Illumination	Illumination	US	1
(2010,0160)	Reflected Ambient Light	Reflected​Ambient​Light	US	1
(2010,0376)	Printer Pixel Spacing	Printer​Pixel​Spacing	DS	2
(2010,0500)	Referenced Film Session Sequence	Referenced​Film​Session​Sequence	SQ	1
(2010,0510)	Referenced Image Box Sequence	Referenced​Image​Box​Sequence	SQ	1
(2010,0520)	Referenced Basic Annotation Box Sequence	Referenced​Basic​Annotation​Box​Sequence	SQ	1
(2020,0010)	Image Box Position	Image​Box​Position	US	1
(2020,0020)	Polarity	Polarity	CS	1
(2020,0030)	Requested Image Size	Requested​Image​Size	DS	1
(2020,0040)	Requested Decimate/Crop Behavior	Requested​Decimate​Crop​Behavior	CS	1
(2020,0050)	Requested Resolution ID	Requested​Resolution​ID	CS	1
(2020,00A0)	Requested Image Size Flag	Requested​Image​Size​Flag	CS	1
(2020,00A2)	Decimate/Crop Result	Decimate​Crop​Result	CS	1
(2020,0110)	Basic Grayscale Image Sequence	Basic​Grayscale​Image​Sequence	SQ	1
(2020,0111)	Basic Color Image Sequence	Basic​Color​Image​Sequence	SQ	1
(2020,0130)	Referenced Image Overlay Box Sequence	Referenced​Image​Overlay​Box​Sequence	SQ	1	RET (1998)
(2020,0140)	Referenced VOI LUT Box Sequence	Referenced​VOILUT​Box​Sequence	SQ	1	RET (1998)
(2030,0010)	Annotation Position	Annotation​Position	US	1
(2030,0020)	Text String	Text​String	LO	1
(2040,0010)	Referenced Overlay Plane Sequence	Referenced​Overlay​Plane​Sequence	SQ	1	RET (2004)
(2040,0011)	Referenced Overlay Plane Groups	Referenced​Overlay​Plane​Groups	US	1-99	RET (2004)
(2040,0020)	Overlay Pixel Data Sequence	Overlay​Pixel​Data​Sequence	SQ	1	RET (2004)
(2040,0060)	Overlay Magnification Type	Overlay​Magnification​Type	CS	1	RET (2004)
(2040,0070)	Overlay Smoothing Type	Overlay​Smoothing​Type	CS	1	RET (2004)
(2040,0072)	Overlay or Image Magnification	Overlay​Or​Image​Magnification	CS	1	RET (2004)
(2040,0074)	Magnify to Number of Columns	Magnify​To​Number​Of​Columns	US	1	RET (2004)
(2040,0080)	Overlay Foreground Density	Overlay​Foreground​Density	CS	1	RET (2004)
(2040,0082)	Overlay Background Density	Overlay​Background​Density	CS	1	RET (2004)
(2040,0090)	Overlay Mode	Overlay​Mode	CS	1	RET (1998)
(2040,0100)	Threshold Density	Threshold​Density	CS	1	RET (1998)
(2040,0500)	Referenced Image Box Sequence (Retired)	Referenced​Image​Box​Sequence​Retired	SQ	1	RET (1993)
(2050,0010)	Presentation LUT Sequence	Presentation​LUT​Sequence	SQ	1
(2050,0020)	Presentation LUT Shape	Presentation​LUT​Shape	CS	1
(2050,0500)	Referenced Presentation LUT Sequence	Referenced​Presentation​LUT​Sequence	SQ	1
(2100,0010)	Print Job ID	Print​Job​ID	SH	1	RET (2007)
(2100,0020)	Execution Status	Execution​Status	CS	1
(2100,0030)	Execution Status Info	Execution​Status​Info	CS	1
(2100,0040)	Creation Date	Creation​Date	DA	1
(2100,0050)	Creation Time	Creation​Time	TM	1
(2100,0070)	Originator	Originator	AE	1
(2100,0140)	Destination AE	Destination​AE	AE	1
(2100,0160)	Owner ID	Owner​ID	SH	1
(2100,0170)	Number of Films	Number​Of​Films	IS	1
(2100,0500)	Referenced Print Job Sequence (Pull Stored Print)	Referenced​Print​Job​Sequence​Pull​Stored​Print	SQ	1	RET (2004)
(2110,0010)	Printer Status	Printer​Status	CS	1
(2110,0020)	Printer Status Info	Printer​Status​Info	CS	1
(2110,0030)	Printer Name	Printer​Name	LO	1
(2110,0099)	Print Queue ID	Print​Queue​ID	SH	1	RET (2004)
(2120,0010)	Queue Status	Queue​Status	CS	1	RET (2004)
(2120,0050)	Print Job Description Sequence	Print​Job​Description​Sequence	SQ	1	RET (2004)
(2120,0070)	Referenced Print Job Sequence	Referenced​Print​Job​Sequence	SQ	1	RET (2004)
(2130,0010)	Print Management Capabilities Sequence	Print​Management​Capabilities​Sequence	SQ	1	RET (2004)
(2130,0015)	Printer Characteristics Sequence	Printer​Characteristics​Sequence	SQ	1	RET (2004)
(2130,0030)	Film Box Content Sequence	Film​Box​Content​Sequence	SQ	1	RET (2004)
(2130,0040)	Image Box Content Sequence	Image​Box​Content​Sequence	SQ	1	RET (2004)
(2130,0050)	Annotation Content Sequence	Annotation​Content​Sequence	SQ	1	RET (2004)
(2130,0060)	Image Overlay Box Content Sequence	Image​Overlay​Box​Content​Sequence	SQ	1	RET (2004)
(2130,0080)	Presentation LUT Content Sequence	Presentation​LUT​Content​Sequence	SQ	1	RET (2004)
(2130,00A0)	Proposed Study Sequence	Proposed​Study​Sequence	SQ	1	RET (2004)
(2130,00C0)	Original Image Sequence	Original​Image​Sequence	SQ	1	RET (2004)
(2200,0001)	Label Using Information Extracted From Instances	Label​Using​Information​Extracted​From​Instances	CS	1
(2200,0002)	Label Text	Label​Text	UT	1
(2200,0003)	Label Style Selection	Label​Style​Selection	CS	1
(2200,0004)	Media Disposition	Media​Disposition	LT	1
(2200,0005)	Barcode Value	Barcode​Value	LT	1
(2200,0006)	Barcode Symbology	Barcode​Symbology	CS	1
(2200,0007)	Allow Media Splitting	Allow​Media​Splitting	CS	1
(2200,0008)	Include Non-DICOM Objects	Include​Non​DICOM​Objects	CS	1
(2200,0009)	Include Display Application	Include​Display​Application	CS	1
(2200,000A)	Preserve Composite Instances After Media Creation	Preserve​Composite​Instances​After​Media​Creation	CS	1
(2200,000B)	Total Number of Pieces of Media Created	Total​Number​Of​Pieces​Of​Media​Created	US	1
(2200,000C)	Requested Media Application Profile	Requested​Media​Application​Profile	LO	1
(2200,000D)	Referenced Storage Media Sequence	Referenced​Storage​Media​Sequence	SQ	1
(2200,000E)	Failure Attributes	Failure​Attributes	AT	1-n
(2200,000F)	Allow Lossy Compression	Allow​Lossy​Compression	CS	1
(2200,0020)	Request Priority	Request​Priority	CS	1
(3002,0002)	RT Image Label	RT​Image​Label	SH	1
(3002,0003)	RT Image Name	RT​Image​Name	LO	1
(3002,0004)	RT Image Description	RT​Image​Description	ST	1
(3002,000A)	Reported Values Origin	Reported​Values​Origin	CS	1
(3002,000C)	RT Image Plane	RT​Image​Plane	CS	1
(3002,000D)	X-Ray Image Receptor Translation	X​Ray​Image​Receptor​Translation	DS	3
(3002,000E)	X-Ray Image Receptor Angle	X​Ray​Image​Receptor​Angle	DS	1
(3002,0010)	RT Image Orientation	RT​Image​Orientation	DS	6
(3002,0011)	Image Plane Pixel Spacing	Image​Plane​Pixel​Spacing	DS	2
(3002,0012)	RT Image Position	RT​Image​Position	DS	2
(3002,0020)	Radiation Machine Name	Radiation​Machine​Name	SH	1
(3002,0022)	Radiation Machine SAD	Radiation​Machine​SAD	DS	1
(3002,0024)	Radiation Machine SSD	Radiation​Machine​SSD	DS	1
(3002,0026)	RT Image SID	RT​Image​SID	DS	1
(3002,0028)	Source to Reference Object Distance	Source​To​Reference​Object​Distance	DS	1
(3002,0029)	Fraction Number	Fraction​Number	IS	1
(3002,0030)	Exposure Sequence	Exposure​Sequence	SQ	1
(3002,0032)	Meterset Exposure	Meterset​Exposure	DS	1
(3002,0034)	Diaphragm Position	Diaphragm​Position	DS	4
(3002,0040)	Fluence Map Sequence	Fluence​Map​Sequence	SQ	1
(3002,0041)	Fluence Data Source	Fluence​Data​Source	CS	1
(3002,0042)	Fluence Data Scale	Fluence​Data​Scale	DS	1
(3002,0050)	Primary Fluence Mode Sequence	Primary​Fluence​Mode​Sequence	SQ	1
(3002,0051)	Fluence Mode	Fluence​Mode	CS	1
(3002,0052)	Fluence Mode ID	Fluence​Mode​ID	SH	1
(3004,0001)	DVH Type	DVH​Type	CS	1
(3004,0002)	Dose Units	Dose​Units	CS	1
(3004,0004)	Dose Type	Dose​Type	CS	1
(3004,0005)	Spatial Transform of Dose	Spatial​Transform​Of​Dose	CS	1
(3004,0006)	Dose Comment	Dose​Comment	LO	1
(3004,0008)	Normalization Point	Normalization​Point	DS	3
(3004,000A)	Dose Summation Type	Dose​Summation​Type	CS	1
(3004,000C)	Grid Frame Offset Vector	Grid​Frame​Offset​Vector	DS	2-n
(3004,000E)	Dose Grid Scaling	Dose​Grid​Scaling	DS	1
(3004,0010)	RT Dose ROI Sequence	RT​Dose​ROI​Sequence	SQ	1
(3004,0012)	Dose Value	Dose​Value	DS	1
(3004,0014)	Tissue Heterogeneity Correction	Tissue​Heterogeneity​Correction	CS	1-3
(3004,0040)	DVH Normalization Point	DVH​Normalization​Point	DS	3
(3004,0042)	DVH Normalization Dose Value	DVH​Normalization​Dose​Value	DS	1
(3004,0050)	DVH Sequence	DVH​Sequence	SQ	1
(3004,0052)	DVH Dose Scaling	DVH​Dose​Scaling	DS	1
(3004,0054)	DVH Volume Units	DVH​Volume​Units	CS	1
(3004,0056)	DVH Number of Bins	DVH​Number​Of​Bins	IS	1
(3004,0058)	DVH Data	DVH​Data	DS	2-2n
(3004,0060)	DVH Referenced ROI Sequence	DVH​Referenced​ROI​Sequence	SQ	1
(3004,0062)	DVH ROI Contribution Type	DVHROI​Contribution​Type	CS	1
(3004,0070)	DVH Minimum Dose	DVH​Minimum​Dose	DS	1
(3004,0072)	DVH Maximum Dose	DVH​Maximum​Dose	DS	1
(3004,0074)	DVH Mean Dose	DVH​Mean​Dose	DS	1
(3006,0002)	Structure Set Label	Structure​Set​Label	SH	1
(3006,0004)	Structure Set Name	Structure​Set​Name	LO	1
(3006,0006)	Structure Set Description	Structure​Set​Description	ST	1
(3006,0008)	Structure Set Date	Structure​Set​Date	DA	1
(3006,0009)	Structure Set Time	Structure​Set​Time	TM	1
(3006,0010)	Referenced Frame of Reference Sequence	Referenced​Frame​Of​Reference​Sequence	SQ	1
(3006,0012)	RT Referenced Study Sequence	RT​Referenced​Study​Sequence	SQ	1
(3006,0014)	RT Referenced Series Sequence	RT​Referenced​Series​Sequence	SQ	1
(3006,0016)	Contour Image Sequence	Contour​Image​Sequence	SQ	1
(3006,0018)	Predecessor Structure Set Sequence	Predecessor​Structure​Set​Sequence	SQ	1
(3006,0020)	Structure Set ROI Sequence	Structure​Set​ROI​Sequence	SQ	1
(3006,0022)	ROI Number	ROI​Number	IS	1
(3006,0024)	Referenced Frame of Reference UID	Referenced​Frame​Of​Reference​UID	UI	1
(3006,0026)	ROI Name	ROI​Name	LO	1
(3006,0028)	ROI Description	ROI​Description	ST	1
(3006,002A)	ROI Display Color	ROI​Display​Color	IS	3
(3006,002C)	ROI Volume	ROI​Volume	DS	1
(3006,0030)	RT Related ROI Sequence	RT​Related​ROI​Sequence	SQ	1
(3006,0033)	RT ROI Relationship	RTROI​Relationship	CS	1
(3006,0036)	ROI Generation Algorithm	ROI​Generation​Algorithm	CS	1
(3006,0037)	ROI Derivation Algorithm Identification Sequence	ROI​Derivation​Algorithm​Identification​Sequence	SQ	1
(3006,0038)	ROI Generation Description	ROI​Generation​Description	LO	1
(3006,0039)	ROI Contour Sequence	ROI​Contour​Sequence	SQ	1
(3006,0040)	Contour Sequence	Contour​Sequence	SQ	1
(3006,0042)	Contour Geometric Type	Contour​Geometric​Type	CS	1
(3006,0044)	Contour Slab Thickness	Contour​Slab​Thickness	DS	1
(3006,0045)	Contour Offset Vector	Contour​Offset​Vector	DS	3
(3006,0046)	Number of Contour Points	Number​Of​Contour​Points	IS	1
(3006,0048)	Contour Number	Contour​Number	IS	1
(3006,0049)	Attached Contours	Attached​Contours	IS	1-n
(3006,0050)	Contour Data	Contour​Data	DS	3-3n
(3006,0080)	RT ROI Observations Sequence	RTROI​Observations​Sequence	SQ	1
(3006,0082)	Observation Number	Observation​Number	IS	1
(3006,0084)	Referenced ROI Number	Referenced​ROI​Number	IS	1
(3006,0085)	ROI Observation Label	ROI​Observation​Label	SH	1
(3006,0086)	RT ROI Identification Code Sequence	RTROI​Identification​Code​Sequence	SQ	1
(3006,0088)	ROI Observation Description	ROI​Observation​Description	ST	1
(3006,00A0)	Related RT ROI Observations Sequence	Related​RTROI​Observations​Sequence	SQ	1
(3006,00A4)	RT ROI Interpreted Type	RTROI​Interpreted​Type	CS	1
(3006,00A6)	ROI Interpreter	ROI​Interpreter	PN	1
(3006,00B0)	ROI Physical Properties Sequence	ROI​Physical​Properties​Sequence	SQ	1
(3006,00B2)	ROI Physical Property	ROI​Physical​Property	CS	1
(3006,00B4)	ROI Physical Property Value	ROI​Physical​Property​Value	DS	1
(3006,00B6)	ROI Elemental Composition Sequence	ROI​Elemental​Composition​Sequence	SQ	1
(3006,00B7)	ROI Elemental Composition Atomic Number	ROI​Elemental​Composition​Atomic​Number	US	1
(3006,00B8)	ROI Elemental Composition Atomic Mass Fraction	ROI​Elemental​Composition​Atomic​Mass​Fraction	FL	1
(3006,00B9)	Additional RT ROI Identification Code Sequence	Additional​RT​ROI​Identification​Code​Sequence	SQ	1	RET (2016c)
(3006,00C0)	Frame of Reference Relationship Sequence	Frame​Of​Reference​Relationship​Sequence	SQ	1	RET (2011)
(3006,00C2)	Related Frame of Reference UID	Related​Frame​Of​Reference​UID	UI	1	RET (2011)
(3006,00C4)	Frame of Reference Transformation Type	Frame​Of​Reference​Transformation​Type	CS	1	RET (2011)
(3006,00C6)	Frame of Reference Transformation Matrix	Frame​Of​Reference​Transformation​Matrix	DS	16
(3006,00C8)	Frame of Reference Transformation Comment	Frame​Of​Reference​Transformation​Comment	LO	1
(3008,0010)	Measured Dose Reference Sequence	Measured​Dose​Reference​Sequence	SQ	1
(3008,0012)	Measured Dose Description	Measured​Dose​Description	ST	1
(3008,0014)	Measured Dose Type	Measured​Dose​Type	CS	1
(3008,0016)	Measured Dose Value	Measured​Dose​Value	DS	1
(3008,0020)	Treatment Session Beam Sequence	Treatment​Session​Beam​Sequence	SQ	1
(3008,0021)	Treatment Session Ion Beam Sequence	Treatment​Session​Ion​Beam​Sequence	SQ	1
(3008,0022)	Current Fraction Number	Current​Fraction​Number	IS	1
(3008,0024)	Treatment Control Point Date	Treatment​Control​Point​Date	DA	1
(3008,0025)	Treatment Control Point Time	Treatment​Control​Point​Time	TM	1
(3008,002A)	Treatment Termination Status	Treatment​Termination​Status	CS	1
(3008,002B)	Treatment Termination Code	Treatment​Termination​Code	SH	1
(3008,002C)	Treatment Verification Status	Treatment​Verification​Status	CS	1
(3008,0030)	Referenced Treatment Record Sequence	Referenced​Treatment​Record​Sequence	SQ	1
(3008,0032)	Specified Primary Meterset	Specified​Primary​Meterset	DS	1
(3008,0033)	Specified Secondary Meterset	Specified​Secondary​Meterset	DS	1
(3008,0036)	Delivered Primary Meterset	Delivered​Primary​Meterset	DS	1
(3008,0037)	Delivered Secondary Meterset	Delivered​Secondary​Meterset	DS	1
(3008,003A)	Specified Treatment Time	Specified​Treatment​Time	DS	1
(3008,003B)	Delivered Treatment Time	Delivered​Treatment​Time	DS	1
(3008,0040)	Control Point Delivery Sequence	Control​Point​Delivery​Sequence	SQ	1
(3008,0041)	Ion Control Point Delivery Sequence	Ion​Control​Point​Delivery​Sequence	SQ	1
(3008,0042)	Specified Meterset	Specified​Meterset	DS	1
(3008,0044)	Delivered Meterset	Delivered​Meterset	DS	1
(3008,0045)	Meterset Rate Set	Meterset​Rate​Set	FL	1
(3008,0046)	Meterset Rate Delivered	Meterset​Rate​Delivered	FL	1
(3008,0047)	Scan Spot Metersets Delivered	Scan​Spot​Metersets​Delivered	FL	1-n
(3008,0048)	Dose Rate Delivered	Dose​Rate​Delivered	DS	1
(3008,0050)	Treatment Summary Calculated Dose Reference Sequence	Treatment​Summary​Calculated​Dose​Reference​Sequence	SQ	1
(3008,0052)	Cumulative Dose to Dose Reference	Cumulative​Dose​To​Dose​Reference	DS	1
(3008,0054)	First Treatment Date	First​Treatment​Date	DA	1
(3008,0056)	Most Recent Treatment Date	Most​Recent​Treatment​Date	DA	1
(3008,005A)	Number of Fractions Delivered	Number​Of​Fractions​Delivered	IS	1
(3008,0060)	Override Sequence	Override​Sequence	SQ	1
(3008,0061)	Parameter Sequence Pointer	Parameter​Sequence​Pointer	AT	1
(3008,0062)	Override Parameter Pointer	Override​Parameter​Pointer	AT	1
(3008,0063)	Parameter Item Index	Parameter​Item​Index	IS	1
(3008,0064)	Measured Dose Reference Number	Measured​Dose​Reference​Number	IS	1
(3008,0065)	Parameter Pointer	Parameter​Pointer	AT	1
(3008,0066)	Override Reason	Override​Reason	ST	1
(3008,0067)	Parameter Value Number	Parameter​Value​Number	US	1
(3008,0068)	Corrected Parameter Sequence	Corrected​Parameter​Sequence	SQ	1
(3008,006A)	Correction Value	Correction​Value	FL	1
(3008,0070)	Calculated Dose Reference Sequence	Calculated​Dose​Reference​Sequence	SQ	1
(3008,0072)	Calculated Dose Reference Number	Calculated​Dose​Reference​Number	IS	1
(3008,0074)	Calculated Dose Reference Description	Calculated​Dose​Reference​Description	ST	1
(3008,0076)	Calculated Dose Reference Dose Value	Calculated​Dose​Reference​Dose​Value	DS	1
(3008,0078)	Start Meterset	Start​Meterset	DS	1
(3008,007A)	End Meterset	End​Meterset	DS	1
(3008,0080)	Referenced Measured Dose Reference Sequence	Referenced​Measured​Dose​Reference​Sequence	SQ	1
(3008,0082)	Referenced Measured Dose Reference Number	Referenced​Measured​Dose​Reference​Number	IS	1
(3008,0090)	Referenced Calculated Dose Reference Sequence	Referenced​Calculated​Dose​Reference​Sequence	SQ	1
(3008,0092)	Referenced Calculated Dose Reference Number	Referenced​Calculated​Dose​Reference​Number	IS	1
(3008,00A0)	Beam Limiting Device Leaf Pairs Sequence	Beam​Limiting​Device​Leaf​Pairs​Sequence	SQ	1
(3008,00B0)	Recorded Wedge Sequence	Recorded​Wedge​Sequence	SQ	1
(3008,00C0)	Recorded Compensator Sequence	Recorded​Compensator​Sequence	SQ	1
(3008,00D0)	Recorded Block Sequence	Recorded​Block​Sequence	SQ	1
(3008,00E0)	Treatment Summary Measured Dose Reference Sequence	Treatment​Summary​Measured​Dose​Reference​Sequence	SQ	1
(3008,00F0)	Recorded Snout Sequence	Recorded​Snout​Sequence	SQ	1
(3008,00F2)	Recorded Range Shifter Sequence	Recorded​Range​Shifter​Sequence	SQ	1
(3008,00F4)	Recorded Lateral Spreading Device Sequence	Recorded​Lateral​Spreading​Device​Sequence	SQ	1
(3008,00F6)	Recorded Range Modulator Sequence	Recorded​Range​Modulator​Sequence	SQ	1
(3008,0100)	Recorded Source Sequence	Recorded​Source​Sequence	SQ	1
(3008,0105)	Source Serial Number	Source​Serial​Number	LO	1
(3008,0110)	Treatment Session Application Setup Sequence	Treatment​Session​Application​Setup​Sequence	SQ	1
(3008,0116)	Application Setup Check	Application​Setup​Check	CS	1
(3008,0120)	Recorded Brachy Accessory Device Sequence	Recorded​Brachy​Accessory​Device​Sequence	SQ	1
(3008,0122)	Referenced Brachy Accessory Device Number	Referenced​Brachy​Accessory​Device​Number	IS	1
(3008,0130)	Recorded Channel Sequence	Recorded​Channel​Sequence	SQ	1
(3008,0132)	Specified Channel Total Time	Specified​Channel​Total​Time	DS	1
(3008,0134)	Delivered Channel Total Time	Delivered​Channel​Total​Time	DS	1
(3008,0136)	Specified Number of Pulses	Specified​Number​Of​Pulses	IS	1
(3008,0138)	Delivered Number of Pulses	Delivered​Number​Of​Pulses	IS	1
(3008,013A)	Specified Pulse Repetition Interval	Specified​Pulse​Repetition​Interval	DS	1
(3008,013C)	Delivered Pulse Repetition Interval	Delivered​Pulse​Repetition​Interval	DS	1
(3008,0140)	Recorded Source Applicator Sequence	Recorded​Source​Applicator​Sequence	SQ	1
(3008,0142)	Referenced Source Applicator Number	Referenced​Source​Applicator​Number	IS	1
(3008,0150)	Recorded Channel Shield Sequence	Recorded​Channel​Shield​Sequence	SQ	1
(3008,0152)	Referenced Channel Shield Number	Referenced​Channel​Shield​Number	IS	1
(3008,0160)	Brachy Control Point Delivered Sequence	Brachy​Control​Point​Delivered​Sequence	SQ	1
(3008,0162)	Safe Position Exit Date	Safe​Position​Exit​Date	DA	1
(3008,0164)	Safe Position Exit Time	Safe​Position​Exit​Time	TM	1
(3008,0166)	Safe Position Return Date	Safe​Position​Return​Date	DA	1
(3008,0168)	Safe Position Return Time	Safe​Position​Return​Time	TM	1
(3008,0171)	Pulse Specific Brachy Control Point Delivered Sequence	Pulse​Specific​Brachy​Control​Point​Delivered​Sequence	SQ	1
(3008,0172)	Pulse Number	Pulse​Number	US	1
(3008,0173)	Brachy Pulse Control Point Delivered Sequence	Brachy​Pulse​Control​Point​Delivered​Sequence	SQ	1
(3008,0200)	Current Treatment Status	Current​Treatment​Status	CS	1
(3008,0202)	Treatment Status Comment	Treatment​Status​Comment	ST	1
(3008,0220)	Fraction Group Summary Sequence	Fraction​Group​Summary​Sequence	SQ	1
(3008,0223)	Referenced Fraction Number	Referenced​Fraction​Number	IS	1
(3008,0224)	Fraction Group Type	Fraction​Group​Type	CS	1
(3008,0230)	Beam Stopper Position	Beam​Stopper​Position	CS	1
(3008,0240)	Fraction Status Summary Sequence	Fraction​Status​Summary​Sequence	SQ	1
(3008,0250)	Treatment Date	Treatment​Date	DA	1
(3008,0251)	Treatment Time	Treatment​Time	TM	1
(300A,0002)	RT Plan Label	RT​Plan​Label	SH	1
(300A,0003)	RT Plan Name	RT​Plan​Name	LO	1
(300A,0004)	RT Plan Description	RT​Plan​Description	ST	1
(300A,0006)	RT Plan Date	RT​Plan​Date	DA	1
(300A,0007)	RT Plan Time	RT​Plan​Time	TM	1
(300A,0009)	Treatment Protocols	Treatment​Protocols	LO	1-n
(300A,000A)	Plan Intent	Plan​Intent	CS	1
(300A,000B)	Treatment Sites	Treatment​Sites	LO	1-n
(300A,000C)	RT Plan Geometry	RT​Plan​Geometry	CS	1
(300A,000E)	Prescription Description	Prescription​Description	ST	1
(300A,0010)	Dose Reference Sequence	Dose​Reference​Sequence	SQ	1
(300A,0012)	Dose Reference Number	Dose​Reference​Number	IS	1
(300A,0013)	Dose Reference UID	Dose​Reference​UID	UI	1
(300A,0014)	Dose Reference Structure Type	Dose​Reference​Structure​Type	CS	1
(300A,0015)	Nominal Beam Energy Unit	Nominal​Beam​Energy​Unit	CS	1
(300A,0016)	Dose Reference Description	Dose​Reference​Description	LO	1
(300A,0018)	Dose Reference Point Coordinates	Dose​Reference​Point​Coordinates	DS	3
(300A,001A)	Nominal Prior Dose	Nominal​Prior​Dose	DS	1
(300A,0020)	Dose Reference Type	Dose​Reference​Type	CS	1
(300A,0021)	Constraint Weight	Constraint​Weight	DS	1
(300A,0022)	Delivery Warning Dose	Delivery​Warning​Dose	DS	1
(300A,0023)	Delivery Maximum Dose	Delivery​Maximum​Dose	DS	1
(300A,0025)	Target Minimum Dose	Target​Minimum​Dose	DS	1
(300A,0026)	Target Prescription Dose	Target​Prescription​Dose	DS	1
(300A,0027)	Target Maximum Dose	Target​Maximum​Dose	DS	1
(300A,0028)	Target Underdose Volume Fraction	Target​Underdose​Volume​Fraction	DS	1
(300A,002A)	Organ at Risk Full-volume Dose	Organ​At​Risk​Full​Volume​Dose	DS	1
(300A,002B)	Organ at Risk Limit Dose	Organ​At​Risk​Limit​Dose	DS	1
(300A,002C)	Organ at Risk Maximum Dose	Organ​At​Risk​Maximum​Dose	DS	1
(300A,002D)	Organ at Risk Overdose Volume Fraction	Organ​At​Risk​Overdose​Volume​Fraction	DS	1
(300A,0040)	Tolerance Table Sequence	Tolerance​Table​Sequence	SQ	1
(300A,0042)	Tolerance Table Number	Tolerance​Table​Number	IS	1
(300A,0043)	Tolerance Table Label	Tolerance​Table​Label	SH	1
(300A,0044)	Gantry Angle Tolerance	Gantry​Angle​Tolerance	DS	1
(300A,0046)	Beam Limiting Device Angle Tolerance	Beam​Limiting​Device​Angle​Tolerance	DS	1
(300A,0048)	Beam Limiting Device Tolerance Sequence	Beam​Limiting​Device​Tolerance​Sequence	SQ	1
(300A,004A)	Beam Limiting Device Position Tolerance	Beam​Limiting​Device​Position​Tolerance	DS	1
(300A,004B)	Snout Position Tolerance	Snout​Position​Tolerance	FL	1
(300A,004C)	Patient Support Angle Tolerance	Patient​Support​Angle​Tolerance	DS	1
(300A,004E)	Table Top Eccentric Angle Tolerance	Table​Top​Eccentric​Angle​Tolerance	DS	1
(300A,004F)	Table Top Pitch Angle Tolerance	Table​Top​Pitch​Angle​Tolerance	FL	1
(300A,0050)	Table Top Roll Angle Tolerance	Table​Top​Roll​Angle​Tolerance	FL	1
(300A,0051)	Table Top Vertical Position Tolerance	Table​Top​Vertical​Position​Tolerance	DS	1
(300A,0052)	Table Top Longitudinal Position Tolerance	Table​Top​Longitudinal​Position​Tolerance	DS	1
(300A,0053)	Table Top Lateral Position Tolerance	Table​Top​Lateral​Position​Tolerance	DS	1
(300A,0055)	RT Plan Relationship	RT​Plan​Relationship	CS	1
(300A,0070)	Fraction Group Sequence	Fraction​Group​Sequence	SQ	1
(300A,0071)	Fraction Group Number	Fraction​Group​Number	IS	1
(300A,0072)	Fraction Group Description	Fraction​Group​Description	LO	1
(300A,0078)	Number of Fractions Planned	Number​Of​Fractions​Planned	IS	1
(300A,0079)	Number of Fraction Pattern Digits Per Day	Number​Of​Fraction​Pattern​Digits​Per​Day	IS	1
(300A,007A)	Repeat Fraction Cycle Length	Repeat​Fraction​Cycle​Length	IS	1
(300A,007B)	Fraction Pattern	Fraction​Pattern	LT	1
(300A,0080)	Number of Beams	Number​Of​Beams	IS	1
(300A,0082)	Beam Dose Specification Point	Beam​Dose​Specification​Point	DS	3
(300A,0083)	Referenced Dose Reference UID	Referenced​Dose​Reference​UID	UI	1
(300A,0084)	Beam Dose	Beam​Dose	DS	1
(300A,0086)	Beam Meterset	Beam​Meterset	DS	1
(300A,0088)	Beam Dose Point Depth	Beam​Dose​Point​Depth	FL	1
(300A,0089)	Beam Dose Point Equivalent Depth	Beam​Dose​Point​Equivalent​Depth	FL	1
(300A,008A)	Beam Dose Point SSD	Beam​Dose​Point​SSD	FL	1
(300A,008B)	Beam Dose Meaning	Beam​Dose​Meaning	CS	1
(300A,008C)	Beam Dose Verification Control Point Sequence	Beam​Dose​Verification​Control​Point​Sequence	SQ	1
(300A,008D)	Average Beam Dose Point Depth	Average​Beam​Dose​Point​Depth	FL	1	RET (2017c)
(300A,008E)	Average Beam Dose Point Equivalent Depth	Average​Beam​Dose​Point​Equivalent​Depth	FL	1	RET (2017c)
(300A,008F)	Average Beam Dose Point SSD	Average​Beam​Dose​Point​SSD	FL	1	RET (2017c)
(300A,0090)	Beam Dose Type	Beam​Dose​Type	CS	1
(300A,0091)	Alternate Beam Dose	Alternate​Beam​Dose	DS	1
(300A,0092)	Alternate Beam Dose Type	Alternate​Beam​Dose​Type	CS	1
(300A,0093)	Depth Value Averaging Flag	Depth​Value​Averaging​Flag	CS	1
(300A,0094)	Beam Dose Point Source to External Contour Distance	Beam​Dose​Point​Source​To​External​Contour​Distance	DS	1
(300A,00A0)	Number of Brachy Application Setups	Number​Of​Brachy​Application​Setups	IS	1
(300A,00A2)	Brachy Application Setup Dose Specification Point	Brachy​Application​Setup​Dose​Specification​Point	DS	3
(300A,00A4)	Brachy Application Setup Dose	Brachy​Application​Setup​Dose	DS	1
(300A,00B0)	Beam Sequence	Beam​Sequence	SQ	1
(300A,00B2)	Treatment Machine Name	Treatment​Machine​Name	SH	1
(300A,00B3)	Primary Dosimeter Unit	Primary​Dosimeter​Unit	CS	1
(300A,00B4)	Source-Axis Distance	Source​Axis​Distance	DS	1
(300A,00B6)	Beam Limiting Device Sequence	Beam​Limiting​Device​Sequence	SQ	1
(300A,00B8)	RT Beam Limiting Device Type	RT​Beam​Limiting​Device​Type	CS	1
(300A,00BA)	Source to Beam Limiting Device Distance	Source​To​Beam​Limiting​Device​Distance	DS	1
(300A,00BB)	Isocenter to Beam Limiting Device Distance	Isocenter​To​Beam​Limiting​Device​Distance	FL	1
(300A,00BC)	Number of Leaf/Jaw Pairs	Number​Of​Leaf​Jaw​Pairs	IS	1
(300A,00BE)	Leaf Position Boundaries	Leaf​Position​Boundaries	DS	3-n
(300A,00C0)	Beam Number	Beam​Number	IS	1
(300A,00C2)	Beam Name	Beam​Name	LO	1
(300A,00C3)	Beam Description	Beam​Description	ST	1
(300A,00C4)	Beam Type	Beam​Type	CS	1
(300A,00C5)	Beam Delivery Duration Limit	Beam​Delivery​Duration​Limit	FD	1
(300A,00C6)	Radiation Type	Radiation​Type	CS	1
(300A,00C7)	High-Dose Technique Type	High​Dose​Technique​Type	CS	1
(300A,00C8)	Reference Image Number	Reference​Image​Number	IS	1
(300A,00CA)	Planned Verification Image Sequence	Planned​Verification​Image​Sequence	SQ	1
(300A,00CC)	Imaging Device-Specific Acquisition Parameters	Imaging​Device​Specific​Acquisition​Parameters	LO	1-n
(300A,00CE)	Treatment Delivery Type	Treatment​Delivery​Type	CS	1
(300A,00D0)	Number of Wedges	Number​Of​Wedges	IS	1
(300A,00D1)	Wedge Sequence	Wedge​Sequence	SQ	1
(300A,00D2)	Wedge Number	Wedge​Number	IS	1
(300A,00D3)	Wedge Type	Wedge​Type	CS	1
(300A,00D4)	Wedge ID	Wedge​ID	SH	1
(300A,00D5)	Wedge Angle	Wedge​Angle	IS	1
(300A,00D6)	Wedge Factor	Wedge​Factor	DS	1
(300A,00D7)	Total Wedge Tray Water-Equivalent Thickness	Total​Wedge​Tray​Water​Equivalent​Thickness	FL	1
(300A,00D8)	Wedge Orientation	Wedge​Orientation	DS	1
(300A,00D9)	Isocenter to Wedge Tray Distance	Isocenter​To​Wedge​Tray​Distance	FL	1
(300A,00DA)	Source to Wedge Tray Distance	Source​To​Wedge​Tray​Distance	DS	1
(300A,00DB)	Wedge Thin Edge Position	Wedge​Thin​Edge​Position	FL	1
(300A,00DC)	Bolus ID	Bolus​ID	SH	1
(300A,00DD)	Bolus Description	Bolus​Description	ST	1
(300A,00DE)	Effective Wedge Angle	Effective​Wedge​Angle	DS	1
(300A,00E0)	Number of Compensators	Number​Of​Compensators	IS	1
(300A,00E1)	Material ID	Material​ID	SH	1
(300A,00E2)	Total Compensator Tray Factor	Total​Compensator​Tray​Factor	DS	1
(300A,00E3)	Compensator Sequence	Compensator​Sequence	SQ	1
(300A,00E4)	Compensator Number	Compensator​Number	IS	1
(300A,00E5)	Compensator ID	Compensator​ID	SH	1
(300A,00E6)	Source to Compensator Tray Distance	Source​To​Compensator​Tray​Distance	DS	1
(300A,00E7)	Compensator Rows	Compensator​Rows	IS	1
(300A,00E8)	Compensator Columns	Compensator​Columns	IS	1
(300A,00E9)	Compensator Pixel Spacing	Compensator​Pixel​Spacing	DS	2
(300A,00EA)	Compensator Position	Compensator​Position	DS	2
(300A,00EB)	Compensator Transmission Data	Compensator​Transmission​Data	DS	1-n
(300A,00EC)	Compensator Thickness Data	Compensator​Thickness​Data	DS	1-n
(300A,00ED)	Number of Boli	Number​Of​Boli	IS	1
(300A,00EE)	Compensator Type	Compensator​Type	CS	1
(300A,00EF)	Compensator Tray ID	Compensator​Tray​ID	SH	1
(300A,00F0)	Number of Blocks	Number​Of​Blocks	IS	1
(300A,00F2)	Total Block Tray Factor	Total​Block​Tray​Factor	DS	1
(300A,00F3)	Total Block Tray Water-Equivalent Thickness	Total​Block​Tray​Water​Equivalent​Thickness	FL	1
(300A,00F4)	Block Sequence	Block​Sequence	SQ	1
(300A,00F5)	Block Tray ID	Block​Tray​ID	SH	1
(300A,00F6)	Source to Block Tray Distance	Source​To​Block​Tray​Distance	DS	1
(300A,00F7)	Isocenter to Block Tray Distance	Isocenter​To​Block​Tray​Distance	FL	1
(300A,00F8)	Block Type	Block​Type	CS	1
(300A,00F9)	Accessory Code	Accessory​Code	LO	1
(300A,00FA)	Block Divergence	Block​Divergence	CS	1
(300A,00FB)	Block Mounting Position	Block​Mounting​Position	CS	1
(300A,00FC)	Block Number	Block​Number	IS	1
(300A,00FE)	Block Name	Block​Name	LO	1
(300A,0100)	Block Thickness	Block​Thickness	DS	1
(300A,0102)	Block Transmission	Block​Transmission	DS	1
(300A,0104)	Block Number of Points	Block​Number​Of​Points	IS	1
(300A,0106)	Block Data	Block​Data	DS	2-2n
(300A,0107)	Applicator Sequence	Applicator​Sequence	SQ	1
(300A,0108)	Applicator ID	Applicator​ID	SH	1
(300A,0109)	Applicator Type	Applicator​Type	CS	1
(300A,010A)	Applicator Description	Applicator​Description	LO	1
(300A,010C)	Cumulative Dose Reference Coefficient	Cumulative​Dose​Reference​Coefficient	DS	1
(300A,010E)	Final Cumulative Meterset Weight	Final​Cumulative​Meterset​Weight	DS	1
(300A,0110)	Number of Control Points	Number​Of​Control​Points	IS	1
(300A,0111)	Control Point Sequence	Control​Point​Sequence	SQ	1
(300A,0112)	Control Point Index	Control​Point​Index	IS	1
(300A,0114)	Nominal Beam Energy	Nominal​Beam​Energy	DS	1
(300A,0115)	Dose Rate Set	Dose​Rate​Set	DS	1
(300A,0116)	Wedge Position Sequence	Wedge​Position​Sequence	SQ	1
(300A,0118)	Wedge Position	Wedge​Position	CS	1
(300A,011A)	Beam Limiting Device Position Sequence	Beam​Limiting​Device​Position​Sequence	SQ	1
(300A,011C)	Leaf/Jaw Positions	Leaf​Jaw​Positions	DS	2-2n
(300A,011E)	Gantry Angle	Gantry​Angle	DS	1
(300A,011F)	Gantry Rotation Direction	Gantry​Rotation​Direction	CS	1
(300A,0120)	Beam Limiting Device Angle	Beam​Limiting​Device​Angle	DS	1
(300A,0121)	Beam Limiting Device Rotation Direction	Beam​Limiting​Device​Rotation​Direction	CS	1
(300A,0122)	Patient Support Angle	Patient​Support​Angle	DS	1
(300A,0123)	Patient Support Rotation Direction	Patient​Support​Rotation​Direction	CS	1
(300A,0124)	Table Top Eccentric Axis Distance	Table​Top​Eccentric​Axis​Distance	DS	1
(300A,0125)	Table Top Eccentric Angle	Table​Top​Eccentric​Angle	DS	1
(300A,0126)	Table Top Eccentric Rotation Direction	Table​Top​Eccentric​Rotation​Direction	CS	1
(300A,0128)	Table Top Vertical Position	Table​Top​Vertical​Position	DS	1
(300A,0129)	Table Top Longitudinal Position	Table​Top​Longitudinal​Position	DS	1
(300A,012A)	Table Top Lateral Position	Table​Top​Lateral​Position	DS	1
(300A,012C)	Isocenter Position	Isocenter​Position	DS	3
(300A,012E)	Surface Entry Point	Surface​Entry​Point	DS	3
(300A,0130)	Source to Surface Distance	Source​To​Surface​Distance	DS	1
(300A,0131)	Average Beam Dose Point Source to External Contour Distance	Average​Beam​Dose​Point​Source​To​External​Contour​Distance	FL	1
(300A,0132)	Source to External Contour Distance	Source​To​External​Contour​Distance	FL	1
(300A,0133)	External Contour Entry Point	External​Contour​Entry​Point	FL	3
(300A,0134)	Cumulative Meterset Weight	Cumulative​Meterset​Weight	DS	1
(300A,0140)	Table Top Pitch Angle	Table​Top​Pitch​Angle	FL	1
(300A,0142)	Table Top Pitch Rotation Direction	Table​Top​Pitch​Rotation​Direction	CS	1
(300A,0144)	Table Top Roll Angle	Table​Top​Roll​Angle	FL	1
(300A,0146)	Table Top Roll Rotation Direction	Table​Top​Roll​Rotation​Direction	CS	1
(300A,0148)	Head Fixation Angle	Head​Fixation​Angle	FL	1
(300A,014A)	Gantry Pitch Angle	Gantry​Pitch​Angle	FL	1
(300A,014C)	Gantry Pitch Rotation Direction	Gantry​Pitch​Rotation​Direction	CS	1
(300A,014E)	Gantry Pitch Angle Tolerance	Gantry​Pitch​Angle​Tolerance	FL	1
(300A,0150)	Fixation Eye	Fixation​Eye	CS	1
(300A,0151)	Chair Head Frame Position	Chair​Head​Frame​Position	DS	1
(300A,0152)	Head Fixation Angle Tolerance	Head​Fixation​Angle​Tolerance	DS	1
(300A,0153)	Chair Head Frame Position Tolerance	Chair​Head​Frame​Position​Tolerance	DS	1
(300A,0154)	Fixation Light Azimuthal Angle Tolerance	Fixation​Light​Azimuthal​Angle​Tolerance	DS	1
(300A,0155)	Fixation Light Polar Angle Tolerance	Fixation​Light​Polar​Angle​Tolerance	DS	1
(300A,0180)	Patient Setup Sequence	Patient​Setup​Sequence	SQ	1
(300A,0182)	Patient Setup Number	Patient​Setup​Number	IS	1
(300A,0183)	Patient Setup Label	Patient​Setup​Label	LO	1
(300A,0184)	Patient Additional Position	Patient​Additional​Position	LO	1
(300A,0190)	Fixation Device Sequence	Fixation​Device​Sequence	SQ	1
(300A,0192)	Fixation Device Type	Fixation​Device​Type	CS	1
(300A,0194)	Fixation Device Label	Fixation​Device​Label	SH	1
(300A,0196)	Fixation Device Description	Fixation​Device​Description	ST	1
(300A,0198)	Fixation Device Position	Fixation​Device​Position	SH	1
(300A,0199)	Fixation Device Pitch Angle	Fixation​Device​Pitch​Angle	FL	1
(300A,019A)	Fixation Device Roll Angle	Fixation​Device​Roll​Angle	FL	1
(300A,01A0)	Shielding Device Sequence	Shielding​Device​Sequence	SQ	1
(300A,01A2)	Shielding Device Type	Shielding​Device​Type	CS	1
(300A,01A4)	Shielding Device Label	Shielding​Device​Label	SH	1
(300A,01A6)	Shielding Device Description	Shielding​Device​Description	ST	1
(300A,01A8)	Shielding Device Position	Shielding​Device​Position	SH	1
(300A,01B0)	Setup Technique	Setup​Technique	CS	1
(300A,01B2)	Setup Technique Description	Setup​Technique​Description	ST	1
(300A,01B4)	Setup Device Sequence	Setup​Device​Sequence	SQ	1
(300A,01B6)	Setup Device Type	Setup​Device​Type	CS	1
(300A,01B8)	Setup Device Label	Setup​Device​Label	SH	1
(300A,01BA)	Setup Device Description	Setup​Device​Description	ST	1
(300A,01BC)	Setup Device Parameter	Setup​Device​Parameter	DS	1
(300A,01D0)	Setup Reference Description	Setup​Reference​Description	ST	1
(300A,01D2)	Table Top Vertical Setup Displacement	Table​Top​Vertical​Setup​Displacement	DS	1
(300A,01D4)	Table Top Longitudinal Setup Displacement	Table​Top​Longitudinal​Setup​Displacement	DS	1
(300A,01D6)	Table Top Lateral Setup Displacement	Table​Top​Lateral​Setup​Displacement	DS	1
(300A,0200)	Brachy Treatment Technique	Brachy​Treatment​Technique	CS	1
(300A,0202)	Brachy Treatment Type	Brachy​Treatment​Type	CS	1
(300A,0206)	Treatment Machine Sequence	Treatment​Machine​Sequence	SQ	1
(300A,0210)	Source Sequence	Source​Sequence	SQ	1
(300A,0212)	Source Number	Source​Number	IS	1
(300A,0214)	Source Type	Source​Type	CS	1
(300A,0216)	Source Manufacturer	Source​Manufacturer	LO	1
(300A,0218)	Active Source Diameter	Active​Source​Diameter	DS	1
(300A,021A)	Active Source Length	Active​Source​Length	DS	1
(300A,021B)	Source Model ID	Source​Model​ID	SH	1
(300A,021C)	Source Description	Source​Description	LO	1
(300A,0222)	Source Encapsulation Nominal Thickness	Source​Encapsulation​Nominal​Thickness	DS	1
(300A,0224)	Source Encapsulation Nominal Transmission	Source​Encapsulation​Nominal​Transmission	DS	1
(300A,0226)	Source Isotope Name	Source​Isotope​Name	LO	1
(300A,0228)	Source Isotope Half Life	Source​Isotope​Half​Life	DS	1
(300A,0229)	Source Strength Units	Source​Strength​Units	CS	1
(300A,022A)	Reference Air Kerma Rate	Reference​Air​Kerma​Rate	DS	1
(300A,022B)	Source Strength	Source​Strength	DS	1
(300A,022C)	Source Strength Reference Date	Source​Strength​Reference​Date	DA	1
(300A,022E)	Source Strength Reference Time	Source​Strength​Reference​Time	TM	1
(300A,0230)	Application Setup Sequence	Application​Setup​Sequence	SQ	1
(300A,0232)	Application Setup Type	Application​Setup​Type	CS	1
(300A,0234)	Application Setup Number	Application​Setup​Number	IS	1
(300A,0236)	Application Setup Name	Application​Setup​Name	LO	1
(300A,0238)	Application Setup Manufacturer	Application​Setup​Manufacturer	LO	1
(300A,0240)	Template Number	Template​Number	IS	1
(300A,0242)	Template Type	Template​Type	SH	1
(300A,0244)	Template Name	Template​Name	LO	1
(300A,0250)	Total Reference Air Kerma	Total​Reference​Air​Kerma	DS	1
(300A,0260)	Brachy Accessory Device Sequence	Brachy​Accessory​Device​Sequence	SQ	1
(300A,0262)	Brachy Accessory Device Number	Brachy​Accessory​Device​Number	IS	1
(300A,0263)	Brachy Accessory Device ID	Brachy​Accessory​Device​ID	SH	1
(300A,0264)	Brachy Accessory Device Type	Brachy​Accessory​Device​Type	CS	1
(300A,0266)	Brachy Accessory Device Name	Brachy​Accessory​Device​Name	LO	1
(300A,026A)	Brachy Accessory Device Nominal Thickness	Brachy​Accessory​Device​Nominal​Thickness	DS	1
(300A,026C)	Brachy Accessory Device Nominal Transmission	Brachy​Accessory​Device​Nominal​Transmission	DS	1
(300A,0271)	Channel Effective Length	Channel​Effective​Length	DS	1
(300A,0272)	Channel Inner Length	Channel​Inner​Length	DS	1
(300A,0273)	Afterloader Channel ID	Afterloader​Channel​ID	SH	1
(300A,0274)	Source Applicator Tip Length	Source​Applicator​Tip​Length	DS	1
(300A,0280)	Channel Sequence	Channel​Sequence	SQ	1
(300A,0282)	Channel Number	Channel​Number	IS	1
(300A,0284)	Channel Length	Channel​Length	DS	1
(300A,0286)	Channel Total Time	Channel​Total​Time	DS	1
(300A,0288)	Source Movement Type	Source​Movement​Type	CS	1
(300A,028A)	Number of Pulses	Number​Of​Pulses	IS	1
(300A,028C)	Pulse Repetition Interval	Pulse​Repetition​Interval	DS	1
(300A,0290)	Source Applicator Number	Source​Applicator​Number	IS	1
(300A,0291)	Source Applicator ID	Source​Applicator​ID	SH	1
(300A,0292)	Source Applicator Type	Source​Applicator​Type	CS	1
(300A,0294)	Source Applicator Name	Source​Applicator​Name	LO	1
(300A,0296)	Source Applicator Length	Source​Applicator​Length	DS	1
(300A,0298)	Source Applicator Manufacturer	Source​Applicator​Manufacturer	LO	1
(300A,029C)	Source Applicator Wall Nominal Thickness	Source​Applicator​Wall​Nominal​Thickness	DS	1
(300A,029E)	Source Applicator Wall Nominal Transmission	Source​Applicator​Wall​Nominal​Transmission	DS	1
(300A,02A0)	Source Applicator Step Size	Source​Applicator​Step​Size	DS	1
(300A,02A2)	Transfer Tube Number	Transfer​Tube​Number	IS	1
(300A,02A4)	Transfer Tube Length	Transfer​Tube​Length	DS	1
(300A,02B0)	Channel Shield Sequence	Channel​Shield​Sequence	SQ	1
(300A,02B2)	Channel Shield Number	Channel​Shield​Number	IS	1
(300A,02B3)	Channel Shield ID	Channel​Shield​ID	SH	1
(300A,02B4)	Channel Shield Name	Channel​Shield​Name	LO	1
(300A,02B8)	Channel Shield Nominal Thickness	Channel​Shield​Nominal​Thickness	DS	1
(300A,02BA)	Channel Shield Nominal Transmission	Channel​Shield​Nominal​Transmission	DS	1
(300A,02C8)	Final Cumulative Time Weight	Final​Cumulative​Time​Weight	DS	1
(300A,02D0)	Brachy Control Point Sequence	Brachy​Control​Point​Sequence	SQ	1
(300A,02D2)	Control Point Relative Position	Control​Point​Relative​Position	DS	1
(300A,02D4)	Control Point 3D Position	Control​Point3D​Position	DS	3
(300A,02D6)	Cumulative Time Weight	Cumulative​Time​Weight	DS	1
(300A,02E0)	Compensator Divergence	Compensator​Divergence	CS	1
(300A,02E1)	Compensator Mounting Position	Compensator​Mounting​Position	CS	1
(300A,02E2)	Source to Compensator Distance	Source​To​Compensator​Distance	DS	1-n
(300A,02E3)	Total Compensator Tray Water-Equivalent Thickness	Total​Compensator​Tray​Water​Equivalent​Thickness	FL	1
(300A,02E4)	Isocenter to Compensator Tray Distance	Isocenter​To​Compensator​Tray​Distance	FL	1
(300A,02E5)	Compensator Column Offset	Compensator​Column​Offset	FL	1
(300A,02E6)	Isocenter to Compensator Distances	Isocenter​To​Compensator​Distances	FL	1-n
(300A,02E7)	Compensator Relative Stopping Power Ratio	Compensator​Relative​Stopping​Power​Ratio	FL	1
(300A,02E8)	Compensator Milling Tool Diameter	Compensator​Milling​Tool​Diameter	FL	1
(300A,02EA)	Ion Range Compensator Sequence	Ion​Range​Compensator​Sequence	SQ	1
(300A,02EB)	Compensator Description	Compensator​Description	LT	1
(300A,0302)	Radiation Mass Number	Radiation​Mass​Number	IS	1
(300A,0304)	Radiation Atomic Number	Radiation​Atomic​Number	IS	1
(300A,0306)	Radiation Charge State	Radiation​Charge​State	SS	1
(300A,0308)	Scan Mode	Scan​Mode	CS	1
(300A,0309)	Modulated Scan Mode Type	Modulated​Scan​Mode​Type	CS	1
(300A,030A)	Virtual Source-Axis Distances	Virtual​Source​Axis​Distances	FL	2
(300A,030C)	Snout Sequence	Snout​Sequence	SQ	1
(300A,030D)	Snout Position	Snout​Position	FL	1
(300A,030F)	Snout ID	Snout​ID	SH	1
(300A,0312)	Number of Range Shifters	Number​Of​Range​Shifters	IS	1
(300A,0314)	Range Shifter Sequence	Range​Shifter​Sequence	SQ	1
(300A,0316)	Range Shifter Number	Range​Shifter​Number	IS	1
(300A,0318)	Range Shifter ID	Range​Shifter​ID	SH	1
(300A,0320)	Range Shifter Type	Range​Shifter​Type	CS	1
(300A,0322)	Range Shifter Description	Range​Shifter​Description	LO	1
(300A,0330)	Number of Lateral Spreading Devices	Number​Of​Lateral​Spreading​Devices	IS	1
(300A,0332)	Lateral Spreading Device Sequence	Lateral​Spreading​Device​Sequence	SQ	1
(300A,0334)	Lateral Spreading Device Number	Lateral​Spreading​Device​Number	IS	1
(300A,0336)	Lateral Spreading Device ID	Lateral​Spreading​Device​ID	SH	1
(300A,0338)	Lateral Spreading Device Type	Lateral​Spreading​Device​Type	CS	1
(300A,033A)	Lateral Spreading Device Description	Lateral​Spreading​Device​Description	LO	1
(300A,033C)	Lateral Spreading Device Water Equivalent Thickness	Lateral​Spreading​Device​Water​Equivalent​Thickness	FL	1
(300A,0340)	Number of Range Modulators	Number​Of​Range​Modulators	IS	1
(300A,0342)	Range Modulator Sequence	Range​Modulator​Sequence	SQ	1
(300A,0344)	Range Modulator Number	Range​Modulator​Number	IS	1
(300A,0346)	Range Modulator ID	Range​Modulator​ID	SH	1
(300A,0348)	Range Modulator Type	Range​Modulator​Type	CS	1
(300A,034A)	Range Modulator Description	Range​Modulator​Description	LO	1
(300A,034C)	Beam Current Modulation ID	Beam​Current​Modulation​ID	SH	1
(300A,0350)	Patient Support Type	Patient​Support​Type	CS	1
(300A,0352)	Patient Support ID	Patient​Support​ID	SH	1
(300A,0354)	Patient Support Accessory Code	Patient​Support​Accessory​Code	LO	1
(300A,0355)	Tray Accessory Code	Tray​Accessory​Code	LO	1
(300A,0356)	Fixation Light Azimuthal Angle	Fixation​Light​Azimuthal​Angle	FL	1
(300A,0358)	Fixation Light Polar Angle	Fixation​Light​Polar​Angle	FL	1
(300A,035A)	Meterset Rate	Meterset​Rate	FL	1
(300A,0360)	Range Shifter Settings Sequence	Range​Shifter​Settings​Sequence	SQ	1
(300A,0362)	Range Shifter Setting	Range​Shifter​Setting	LO	1
(300A,0364)	Isocenter to Range Shifter Distance	Isocenter​To​Range​Shifter​Distance	FL	1
(300A,0366)	Range Shifter Water Equivalent Thickness	Range​Shifter​Water​Equivalent​Thickness	FL	1
(300A,0370)	Lateral Spreading Device Settings Sequence	Lateral​Spreading​Device​Settings​Sequence	SQ	1
(300A,0372)	Lateral Spreading Device Setting	Lateral​Spreading​Device​Setting	LO	1
(300A,0374)	Isocenter to Lateral Spreading Device Distance	Isocenter​To​Lateral​Spreading​Device​Distance	FL	1
(300A,0380)	Range Modulator Settings Sequence	Range​Modulator​Settings​Sequence	SQ	1
(300A,0382)	Range Modulator Gating Start Value	Range​Modulator​Gating​Start​Value	FL	1
(300A,0384)	Range Modulator Gating Stop Value	Range​Modulator​Gating​Stop​Value	FL	1
(300A,0386)	Range Modulator Gating Start Water Equivalent Thickness	Range​Modulator​Gating​Start​Water​Equivalent​Thickness	FL	1
(300A,0388)	Range Modulator Gating Stop Water Equivalent Thickness	Range​Modulator​Gating​Stop​Water​Equivalent​Thickness	FL	1
(300A,038A)	Isocenter to Range Modulator Distance	Isocenter​To​Range​Modulator​Distance	FL	1
(300A,038F)	Scan Spot Time Offset	Scan​Spot​Time​Offset	FL	1-n
(300A,0390)	Scan Spot Tune ID	Scan​Spot​Tune​ID	SH	1
(300A,0391)	Scan Spot Prescribed Indices	Scan​Spot​Prescribed​Indices	IS	1-n
(300A,0392)	Number of Scan Spot Positions	Number​Of​Scan​Spot​Positions	IS	1
(300A,0393)	Scan Spot Reordered	Scan​Spot​Reordered	CS	1
(300A,0394)	Scan Spot Position Map	Scan​Spot​Position​Map	FL	1-n
(300A,0395)	Scan Spot Reordering Allowed	Scan​Spot​Reordering​Allowed	CS	1
(300A,0396)	Scan Spot Meterset Weights	Scan​Spot​Meterset​Weights	FL	1-n
(300A,0398)	Scanning Spot Size	Scanning​Spot​Size	FL	2
(300A,039A)	Number of Paintings	Number​Of​Paintings	IS	1
(300A,03A0)	Ion Tolerance Table Sequence	Ion​Tolerance​Table​Sequence	SQ	1
(300A,03A2)	Ion Beam Sequence	Ion​Beam​Sequence	SQ	1
(300A,03A4)	Ion Beam Limiting Device Sequence	Ion​Beam​Limiting​Device​Sequence	SQ	1
(300A,03A6)	Ion Block Sequence	Ion​Block​Sequence	SQ	1
(300A,03A8)	Ion Control Point Sequence	Ion​Control​Point​Sequence	SQ	1
(300A,03AA)	Ion Wedge Sequence	Ion​Wedge​Sequence	SQ	1
(300A,03AC)	Ion Wedge Position Sequence	Ion​Wedge​Position​Sequence	SQ	1
(300A,0401)	Referenced Setup Image Sequence	Referenced​Setup​Image​Sequence	SQ	1
(300A,0402)	Setup Image Comment	Setup​Image​Comment	ST	1
(300A,0410)	Motion Synchronization Sequence	Motion​Synchronization​Sequence	SQ	1
(300A,0412)	Control Point Orientation	Control​Point​Orientation	FL	3
(300A,0420)	General Accessory Sequence	General​Accessory​Sequence	SQ	1
(300A,0421)	General Accessory ID	General​Accessory​ID	SH	1
(300A,0422)	General Accessory Description	General​Accessory​Description	ST	1
(300A,0423)	General Accessory Type	General​Accessory​Type	CS	1
(300A,0424)	General Accessory Number	General​Accessory​Number	IS	1
(300A,0425)	Source to General Accessory Distance	Source​To​General​Accessory​Distance	FL	1
(300A,0431)	Applicator Geometry Sequence	Applicator​Geometry​Sequence	SQ	1
(300A,0432)	Applicator Aperture Shape	Applicator​Aperture​Shape	CS	1
(300A,0433)	Applicator Opening	Applicator​Opening	FL	1
(300A,0434)	Applicator Opening X	Applicator​Opening​X	FL	1
(300A,0435)	Applicator Opening Y	Applicator​Opening​Y	FL	1
(300A,0436)	Source to Applicator Mounting Position Distance	Source​To​Applicator​Mounting​Position​Distance	FL	1
(300A,0440)	Number of Block Slab Items	Number​Of​Block​Slab​Items	IS	1
(300A,0441)	Block Slab Sequence	Block​Slab​Sequence	SQ	1
(300A,0442)	Block Slab Thickness	Block​Slab​Thickness	DS	1
(300A,0443)	Block Slab Number	Block​Slab​Number	US	1
(300A,0450)	Device Motion Control Sequence	Device​Motion​Control​Sequence	SQ	1
(300A,0451)	Device Motion Execution Mode	Device​Motion​Execution​Mode	CS	1
(300A,0452)	Device Motion Observation Mode	Device​Motion​Observation​Mode	CS	1
(300A,0453)	Device Motion Parameter Code Sequence	Device​Motion​Parameter​Code​Sequence	SQ	1
(300A,0501)	Distal Depth Fraction	Distal​Depth​Fraction	FL	1
(300A,0502)	Distal Depth	Distal​Depth	FL	1
(300A,0503)	Nominal Range Modulation Fractions	Nominal​Range​Modulation​Fractions	FL	2
(300A,0504)	Nominal Range Modulated Region Depths	Nominal​Range​Modulated​Region​Depths	FL	2
(300A,0505)	Depth Dose Parameters Sequence	Depth​Dose​Parameters​Sequence	SQ	1
(300A,0506)	Delivered Depth Dose Parameters Sequence	Delivered​Depth​Dose​Parameters​Sequence	SQ	1
(300A,0507)	Delivered Distal Depth Fraction	Delivered​Distal​Depth​Fraction	FL	1
(300A,0508)	Delivered Distal Depth	Delivered​Distal​Depth	FL	1
(300A,0509)	Delivered Nominal Range Modulation Fractions	Delivered​Nominal​Range​Modulation​Fractions	FL	2
(300A,0510)	Delivered Nominal Range Modulated Region Depths	Delivered​Nominal​Range​Modulated​Region​Depths	FL	2
(300A,0511)	Delivered Reference Dose Definition	Delivered​Reference​Dose​Definition	CS	1
(300A,0512)	Reference Dose Definition	Reference​Dose​Definition	CS	1
(300C,0002)	Referenced RT Plan Sequence	Referenced​RT​Plan​Sequence	SQ	1
(300C,0004)	Referenced Beam Sequence	Referenced​Beam​Sequence	SQ	1
(300C,0006)	Referenced Beam Number	Referenced​Beam​Number	IS	1
(300C,0007)	Referenced Reference Image Number	Referenced​Reference​Image​Number	IS	1
(300C,0008)	Start Cumulative Meterset Weight	Start​Cumulative​Meterset​Weight	DS	1
(300C,0009)	End Cumulative Meterset Weight	End​Cumulative​Meterset​Weight	DS	1
(300C,000A)	Referenced Brachy Application Setup Sequence	Referenced​Brachy​Application​Setup​Sequence	SQ	1
(300C,000C)	Referenced Brachy Application Setup Number	Referenced​Brachy​Application​Setup​Number	IS	1
(300C,000E)	Referenced Source Number	Referenced​Source​Number	IS	1
(300C,0020)	Referenced Fraction Group Sequence	Referenced​Fraction​Group​Sequence	SQ	1
(300C,0022)	Referenced Fraction Group Number	Referenced​Fraction​Group​Number	IS	1
(300C,0040)	Referenced Verification Image Sequence	Referenced​Verification​Image​Sequence	SQ	1
(300C,0042)	Referenced Reference Image Sequence	Referenced​Reference​Image​Sequence	SQ	1
(300C,0050)	Referenced Dose Reference Sequence	Referenced​Dose​Reference​Sequence	SQ	1
(300C,0051)	Referenced Dose Reference Number	Referenced​Dose​Reference​Number	IS	1
(300C,0055)	Brachy Referenced Dose Reference Sequence	Brachy​Referenced​Dose​Reference​Sequence	SQ	1
(300C,0060)	Referenced Structure Set Sequence	Referenced​Structure​Set​Sequence	SQ	1
(300C,006A)	Referenced Patient Setup Number	Referenced​Patient​Setup​Number	IS	1
(300C,0080)	Referenced Dose Sequence	Referenced​Dose​Sequence	SQ	1
(300C,00A0)	Referenced Tolerance Table Number	Referenced​Tolerance​Table​Number	IS	1
(300C,00B0)	Referenced Bolus Sequence	Referenced​Bolus​Sequence	SQ	1
(300C,00C0)	Referenced Wedge Number	Referenced​Wedge​Number	IS	1
(300C,00D0)	Referenced Compensator Number	Referenced​Compensator​Number	IS	1
(300C,00E0)	Referenced Block Number	Referenced​Block​Number	IS	1
(300C,00F0)	Referenced Control Point Index	Referenced​Control​Point​Index	IS	1
(300C,00F2)	Referenced Control Point Sequence	Referenced​Control​Point​Sequence	SQ	1
(300C,00F4)	Referenced Start Control Point Index	Referenced​Start​Control​Point​Index	IS	1
(300C,00F6)	Referenced Stop Control Point Index	Referenced​Stop​Control​Point​Index	IS	1
(300C,0100)	Referenced Range Shifter Number	Referenced​Range​Shifter​Number	IS	1
(300C,0102)	Referenced Lateral Spreading Device Number	Referenced​Lateral​Spreading​Device​Number	IS	1
(300C,0104)	Referenced Range Modulator Number	Referenced​Range​Modulator​Number	IS	1
(300C,0111)	Omitted Beam Task Sequence	Omitted​Beam​Task​Sequence	SQ	1
(300C,0112)	Reason for Omission	Reason​For​Omission	CS	1
(300C,0113)	Reason for Omission Description	Reason​For​Omission​Description	LO	1
(300E,0002)	Approval Status	Approval​Status	CS	1
(300E,0004)	Review Date	Review​Date	DA	1
(300E,0005)	Review Time	Review​Time	TM	1
(300E,0008)	Reviewer Name	Reviewer​Name	PN	1
(3010,0001)	Radiobiological Dose Effect Sequence	Radiobiological​Dose​Effect​Sequence	SQ	1
(3010,0002)	Radiobiological Dose Effect Flag	Radiobiological​Dose​Effect​Flag	CS	1
(3010,0003)	Effective Dose Calculation Method Category Code Sequence	Effective​Dose​Calculation​Method​Category​Code​Sequence	SQ	1
(3010,0004)	Effective Dose Calculation Method Code Sequence	Effective​Dose​Calculation​Method​Code​Sequence	SQ	1
(3010,0005)	Effective Dose Calculation Method Description	Effective​Dose​Calculation​Method​Description	LO	1
(3010,0006)	Conceptual Volume UID	Conceptual​Volume​UID	UI	1
(3010,0007)	Originating SOP Instance Reference Sequence	Originating​SOPInstance​Reference​Sequence	SQ	1
(3010,0008)	Conceptual Volume Constituent Sequence	Conceptual​Volume​Constituent​Sequence	SQ	1
(3010,0009)	Equivalent Conceptual Volume Instance Reference Sequence	Equivalent​Conceptual​Volume​Instance​Reference​Sequence	SQ	1
(3010,000A)	Equivalent Conceptual Volumes Sequence	Equivalent​Conceptual​Volumes​Sequence	SQ	1
(3010,000B)	Referenced Conceptual Volume UID	Referenced​Conceptual​Volume​UID	UI	1
(3010,000C)	Conceptual Volume Combination Expression	Conceptual​Volume​Combination​Expression	UT	1
(3010,000D)	Conceptual Volume Constituent Index	Conceptual​Volume​Constituent​Index	US	1
(3010,000E)	Conceptual Volume Combination Flag	Conceptual​Volume​Combination​Flag	CS	1
(3010,000F)	Conceptual Volume Combination Description	Conceptual​Volume​Combination​Description	ST	1
(3010,0010)	Conceptual Volume Segmentation Defined Flag	Conceptual​Volume​Segmentation​Defined​Flag	CS	1
(3010,0011)	Conceptual Volume Segmentation Reference Sequence	Conceptual​Volume​Segmentation​Reference​Sequence	SQ	1
(3010,0012)	Conceptual Volume Constituent Segmentation Reference Sequence	Conceptual​Volume​Constituent​Segmentation​Reference​Sequence	SQ	1
(3010,0013)	Constituent Conceptual Volume UID	Constituent​Conceptual​Volume​UID	UI	1
(3010,0014)	Derivation Conceptual Volume Sequence	Derivation​Conceptual​Volume​Sequence	SQ	1
(3010,0015)	Source Conceptual Volume UID	Source​Conceptual​Volume​UID	UI	1
(3010,0016)	Conceptual Volume Derivation Algorithm Sequence	Conceptual​Volume​Derivation​Algorithm​Sequence	SQ	1
(3010,0017)	Conceptual Volume Description	Conceptual​Volume​Description	ST	1
(3010,0018)	Source Conceptual Volume Sequence	Source​Conceptual​Volume​Sequence	SQ	1
(3010,0019)	Author Identification Sequence	Author​Identification​Sequence	SQ	1
(3010,001A)	Manufacturer's Model Version	Manufacturer​Model​Version	LO	1
(3010,001B)	Device Alternate Identifier	Device​Alternate​Identifier	UC	1
(3010,001C)	Device Alternate Identifier Type	Device​Alternate​Identifier​Type	CS	1
(3010,001D)	Device Alternate Identifier Format	Device​Alternate​Identifier​Format	LT	1
(3010,001E)	Segmentation Creation Template Label	Segmentation​Creation​Template​Label	LO	1
(3010,001F)	Segmentation Template UID	Segmentation​Template​UID	UI	1
(3010,0020)	Referenced Segment Reference Index	Referenced​Segment​Reference​Index	US	1
(3010,0021)	Segment Reference Sequence	Segment​Reference​Sequence	SQ	1
(3010,0022)	Segment Reference Index	Segment​Reference​Index	US	1
(3010,0023)	Direct Segment Reference Sequence	Direct​Segment​Reference​Sequence	SQ	1
(3010,0024)	Combination Segment Reference Sequence	Combination​Segment​Reference​Sequence	SQ	1
(3010,0025)	Conceptual Volume Sequence	Conceptual​Volume​Sequence	SQ	1
(3010,0026)	Segmented RT Accessory Device Sequence	Segmented​RTAccessory​Device​Sequence	SQ	1
(3010,0027)	Segment Characteristics Sequence	Segment​Characteristics​Sequence	SQ	1
(3010,0028)	Related Segment Characteristics Sequence	Related​Segment​Characteristics​Sequence	SQ	1
(3010,0029)	Segment Characteristics Precedence	Segment​Characteristics​Precedence	US	1
(3010,002A)	RT Segment Annotation Sequence	RTSegment​Annotation​Sequence	SQ	1
(3010,002B)	Segment Annotation Category Code Sequence	Segment​Annotation​Category​Code​Sequence	SQ	1
(3010,002C)	Segment Annotation Type Code Sequence	Segment​Annotation​Type​Code​Sequence	SQ	1
(3010,002D)	Device Label	Device​Label	LO	1
(3010,002E)	Device Type Code Sequence	Device​Type​Code​Sequence	SQ	1
(3010,0030)	Patient Equipment Relationship Code Sequence	Patient​Equipment​Relationship​Code​Sequence	SQ	1
(3010,0031)	Referenced Fiducials UID	Referenced​Fiducials​UID	UI	1
(3010,0032)	Patient Treatment Orientation Sequence	Patient​Treatment​Orientation​Sequence	SQ	1
(3010,0033)	User Content Label	User​Content​Label	SH	1
(3010,0034)	User Content Long Label	User​Content​Long​Label	LO	1
(3010,0035)	Entity Label	Entity​Label	SH	1
(3010,0036)	Entity Name	Entity​Name	LO	1
(3010,0037)	Entity Description	Entity​Description	ST	1
(3010,0038)	Entity Long Label	Entity​Long​Label	LO	1
(3010,0039)	Device Index	Device​Index	US	1
(3010,003A)	RT Treatment Phase Index	RTTreatment​Phase​Index	US	1
(3010,003B)	RT Treatment Phase UID	RTTreatment​Phase​UID	UI	1
(3010,003C)	RT Prescription Index	RTPrescription​Index	US	1
(3010,003D)	RT Segment Annotation Index	RTSegment​Annotation​Index	US	1
(3010,003E)	Basis RT Treatment Phase Index	Basis​RTTreatment​Phase​Index	US	1
(3010,003F)	Related RT Treatment Phase Index	Related​RTTreatment​Phase​Index	US	1
(3010,0040)	Referenced RT Treatment Phase Index	Referenced​RTTreatment​Phase​Index	US	1
(3010,0041)	Referenced RT Prescription Index	Referenced​RTPrescription​Index	US	1
(3010,0042)	Referenced Parent RT Prescription Index	Referenced​Parent​RTPrescription​Index	US	1
(3010,0043)	Manufacturer's Device Identifier	Manufacturer​Device​Identifier	ST	1
(3010,0044)	Instance-Level Referenced Performed Procedure Step Sequence	Instance​Level​Referenced​Performed​Procedure​Step​Sequence	SQ	1
(3010,0045)	RT Treatment Phase Intent Presence Flag	RTTreatment​Phase​Intent​Presence​Flag	CS	1
(3010,0046)	Radiotherapy Treatment Type	Radiotherapy​Treatment​Type	CS	1
(3010,0047)	Teletherapy Radiation Type	Teletherapy​Radiation​Type	CS	1-n
(3010,0048)	Brachytherapy Source Type	Brachytherapy​Source​Type	CS	1-n
(3010,0049)	Referenced RT Treatment Phase Sequence	Referenced​RTTreatment​Phase​Sequence	SQ	1
(3010,004A)	Referenced Direct Segment Instance Sequence	Referenced​Direct​Segment​Instance​Sequence	SQ	1
(3010,004B)	Intended RT Treatment Phase Sequence	Intended​RTTreatment​Phase​Sequence	SQ	1
(3010,004C)	Intended Phase Start Date	Intended​Phase​Start​Date	DA	1
(3010,004D)	Intended Phase End Date	Intended​Phase​End​Date	DA	1
(3010,004E)	RT Treatment Phase Interval Sequence	RTTreatment​Phase​Interval​Sequence	SQ	1
(3010,004F)	Temporal Relationship Interval Anchor	Temporal​Relationship​Interval​Anchor	CS	1
(3010,0050)	Minimum Number of Interval Days	Minimum​Number​Of​Interval​Days	FD	1
(3010,0051)	Maximum Number of Interval Days	Maximum​Number​Of​Interval​Days	FD	1
(3010,0052)	Pertinent SOP Classes in Study	Pertinent​SOPClasses​In​Study	UI	1-n
(3010,0053)	Pertinent SOP Classes in Series	Pertinent​SOPClasses​In​Series	UI	1-n
(3010,0054)	RT Prescription Label	RTPrescription​Label	LO	1
(3010,0055)	RT Physician Intent Predecessor Sequence	RTPhysician​Intent​Predecessor​Sequence	SQ	1
(3010,0056)	RT Treatment Approach Label	RTTreatment​Approach​Label	LO	1
(3010,0057)	RT Physician Intent Sequence	RTPhysician​Intent​Sequence	SQ	1
(3010,0058)	RT Physician Intent Index	RTPhysician​Intent​Index	US	1
(3010,0059)	RT Treatment Intent Type	RTTreatment​Intent​Type	CS	1
(3010,005A)	RT Physician Intent Narrative	RTPhysician​Intent​Narrative	UT	1
(3010,005B)	RT Protocol Code Sequence	RTProtocol​Code​Sequence	SQ	1
(3010,005C)	Reason for Superseding	Reason​For​Superseding	ST	1
(3010,005D)	RT Diagnosis Code Sequence	RTDiagnosis​Code​Sequence	SQ	1
(3010,005E)	Referenced RT Physician Intent Index	Referenced​RTPhysician​Intent​Index	US	1
(3010,005F)	RT Physician Intent Input Instance Sequence	RTPhysician​Intent​Input​Instance​Sequence	SQ	1
(3010,0060)	RT Anatomic Prescription Sequence	RTAnatomic​Prescription​Sequence	SQ	1
(3010,0061)	Prior Treatment Dose Description	Prior​Treatment​Dose​Description	UT	1
(3010,0062)	Prior Treatment Reference Sequence	Prior​Treatment​Reference​Sequence	SQ	1
(3010,0063)	Dosimetric Objective Evaluation Scope	Dosimetric​Objective​Evaluation​Scope	CS	1
(3010,0064)	Therapeutic Role Category Code Sequence	Therapeutic​Role​Category​Code​Sequence	SQ	1
(3010,0065)	Therapeutic Role Type Code Sequence	Therapeutic​Role​Type​Code​Sequence	SQ	1
(3010,0066)	Conceptual Volume Optimization Precedence	Conceptual​Volume​Optimization​Precedence	US	1
(3010,0067)	Conceptual Volume Category Code Sequence	Conceptual​Volume​Category​Code​Sequence	SQ	1
(3010,0068)	Conceptual Volume Blocking Constraint	Conceptual​Volume​Blocking​Constraint	CS	1
(3010,0069)	Conceptual Volume Type Code Sequence	Conceptual​Volume​Type​Code​Sequence	SQ	1
(3010,006A)	Conceptual Volume Type Modifier Code Sequence	Conceptual​Volume​Type​Modifier​Code​Sequence	SQ	1
(3010,006B)	RT Prescription Sequence	RTPrescription​Sequence	SQ	1
(3010,006C)	Dosimetric Objective Sequence	Dosimetric​Objective​Sequence	SQ	1
(3010,006D)	Dosimetric Objective Type Code Sequence	Dosimetric​Objective​Type​Code​Sequence	SQ	1
(3010,006E)	Dosimetric Objective UID	Dosimetric​Objective​UID	UI	1
(3010,006F)	Referenced Dosimetric Objective UID	Referenced​Dosimetric​Objective​UID	UI	1
(3010,0070)	Dosimetric Objective Parameter Sequence	Dosimetric​Objective​Parameter​Sequence	SQ	1
(3010,0071)	Referenced Dosimetric Objectives Sequence	Referenced​Dosimetric​Objectives​Sequence	SQ	1
(3010,0073)	Absolute Dosimetric Objective Flag	Absolute​Dosimetric​Objective​Flag	CS	1
(3010,0074)	Dosimetric Objective Weight	Dosimetric​Objective​Weight	FD	1
(3010,0075)	Dosimetric Objective Purpose	Dosimetric​Objective​Purpose	CS	1
(3010,0076)	Planning Input Information Sequence	Planning​Input​Information​Sequence	SQ	1
(3010,0077)	Treatment Site	Treatment​Site	LO	1
(3010,0078)	Treatment Site Code Sequence	Treatment​Site​Code​Sequence	SQ	1
(3010,0079)	Fraction Pattern Sequence	Fraction​Pattern​Sequence	SQ	1
(3010,007A)	Treatment Technique Notes	Treatment​Technique​Notes	UT	1
(3010,007B)	Prescription Notes	Prescription​Notes	UT	1
(3010,007C)	Number of Interval Fractions	Number​Of​Interval​Fractions	IS	1
(3010,007D)	Number of Fractions	Number​Of​Fractions	US	1
(3010,007E)	Intended Delivery Duration	Intended​Delivery​Duration	US	1
(3010,007F)	Fractionation Notes	Fractionation​Notes	UT	1
(3010,0080)	RT Treatment Technique Code Sequence	RTTreatment​Technique​Code​Sequence	SQ	1
(3010,0081)	Prescription Notes Sequence	Prescription​Notes​Sequence	SQ	1
(3010,0082)	Fraction-Based Relationship Sequence	Fraction​Based​Relationship​Sequence	SQ	1
(3010,0083)	Fraction-Based Relationship Interval Anchor	Fraction​Based​Relationship​Interval​Anchor	CS	1
(3010,0084)	Minimum Hours between Fractions	Minimum​Hours​Between​Fractions	FD	1
(3010,0085)	Intended Fraction Start Time	Intended​Fraction​Start​Time	TM	1-n
(3010,0086)	Intended Start Day of Week	Intended​Start​Day​Of​Week	LT	1
(3010,0087)	Weekday Fraction Pattern Sequence	Weekday​Fraction​Pattern​Sequence	SQ	1
(3010,0088)	Delivery Time Structure Code Sequence	Delivery​Time​Structure​Code​Sequence	SQ	1
(4000,0010)	Arbitrary	Arbitrary	LT	1	RET
(4000,4000)	Text Comments	Text​Comments	LT	1	RET
(4008,0040)	Results ID	Results​ID	SH	1	RET (2004)
(4008,0042)	Results ID Issuer	Results​ID​Issuer	LO	1	RET (2004)
(4008,0050)	Referenced Interpretation Sequence	Referenced​Interpretation​Sequence	SQ	1	RET (2004)
(4008,00FF)	Report Production Status (Trial)	Report​Production​Status​Trial	CS	1	RET (2009)
(4008,0100)	Interpretation Recorded Date	Interpretation​Recorded​Date	DA	1	RET (2004)
(4008,0101)	Interpretation Recorded Time	Interpretation​Recorded​Time	TM	1	RET (2004)
(4008,0102)	Interpretation Recorder	Interpretation​Recorder	PN	1	RET (2004)
(4008,0103)	Reference to Recorded Sound	Reference​To​Recorded​Sound	LO	1	RET (2004)
(4008,0108)	Interpretation Transcription Date	Interpretation​Transcription​Date	DA	1	RET (2004)
(4008,0109)	Interpretation Transcription Time	Interpretation​Transcription​Time	TM	1	RET (2004)
(4008,010A)	Interpretation Transcriber	Interpretation​Transcriber	PN	1	RET (2004)
(4008,010B)	Interpretation Text	Interpretation​Text	ST	1	RET (2004)
(4008,010C)	Interpretation Author	Interpretation​Author	PN	1	RET (2004)
(4008,0111)	Interpretation Approver Sequence	Interpretation​Approver​Sequence	SQ	1	RET (2004)
(4008,0112)	Interpretation Approval Date	Interpretation​Approval​Date	DA	1	RET (2004)
(4008,0113)	Interpretation Approval Time	Interpretation​Approval​Time	TM	1	RET (2004)
(4008,0114)	Physician Approving Interpretation	Physician​Approving​Interpretation	PN	1	RET (2004)
(4008,0115)	Interpretation Diagnosis Description	Interpretation​Diagnosis​Description	LT	1	RET (2004)
(4008,0117)	Interpretation Diagnosis Code Sequence	Interpretation​Diagnosis​Code​Sequence	SQ	1	RET (2004)
(4008,0118)	Results Distribution List Sequence	Results​Distribution​List​Sequence	SQ	1	RET (2004)
(4008,0119)	Distribution Name	Distribution​Name	PN	1	RET (2004)
(4008,011A)	Distribution Address	Distribution​Address	LO	1	RET (2004)
(4008,0200)	Interpretation ID	Interpretation​ID	SH	1	RET (2004)
(4008,0202)	Interpretation ID Issuer	Interpretation​ID​Issuer	LO	1	RET (2004)
(4008,0210)	Interpretation Type ID	Interpretation​Type​ID	CS	1	RET (2004)
(4008,0212)	Interpretation Status ID	Interpretation​Status​ID	CS	1	RET (2004)
(4008,0300)	Impressions	Impressions	ST	1	RET (2004)
(4008,4000)	Results Comments	Results​Comments	ST	1	RET (2004)
(4010,0001)	Low Energy Detectors	Low​Energy​Detectors	CS	1	DICOS
(4010,0002)	High Energy Detectors	High​Energy​Detectors	CS	1	DICOS
(4010,0004)	Detector Geometry Sequence	Detector​Geometry​Sequence	SQ	1	DICOS
(4010,1001)	Threat ROI Voxel Sequence	Threat​ROI​Voxel​Sequence	SQ	1	DICOS
(4010,1004)	Threat ROI Base	Threat​ROI​Base	FL	3	DICOS
(4010,1005)	Threat ROI Extents	Threat​ROI​Extents	FL	3	DICOS
(4010,1006)	Threat ROI Bitmap	Threat​ROI​Bitmap	OB	1	DICOS
(4010,1007)	Route Segment ID	Route​Segment​ID	SH	1	DICOS
(4010,1008)	Gantry Type	Gantry​Type	CS	1	DICOS
(4010,1009)	OOI Owner Type	OOI​Owner​Type	CS	1	DICOS
(4010,100A)	Route Segment Sequence	Route​Segment​Sequence	SQ	1	DICOS
(4010,1010)	Potential Threat Object ID	Potential​Threat​Object​ID	US	1	DICOS
(4010,1011)	Threat Sequence	Threat​Sequence	SQ	1	DICOS
(4010,1012)	Threat Category	Threat​Category	CS	1	DICOS
(4010,1013)	Threat Category Description	Threat​Category​Description	LT	1	DICOS
(4010,1014)	ATD Ability Assessment	ATD​Ability​Assessment	CS	1	DICOS
(4010,1015)	ATD Assessment Flag	ATD​Assessment​Flag	CS	1	DICOS
(4010,1016)	ATD Assessment Probability	ATD​Assessment​Probability	FL	1	DICOS
(4010,1017)	Mass	Mass	FL	1	DICOS
(4010,1018)	Density	Density	FL	1	DICOS
(4010,1019)	Z Effective	Z​Effective	FL	1	DICOS
(4010,101A)	Boarding Pass ID	Boarding​Pass​ID	SH	1	DICOS
(4010,101B)	Center of Mass	Center​Of​Mass	FL	3	DICOS
(4010,101C)	Center of PTO	Center​Of​PTO	FL	3	DICOS
(4010,101D)	Bounding Polygon	Bounding​Polygon	FL	6-n	DICOS
(4010,101E)	Route Segment Start Location ID	Route​Segment​Start​Location​ID	SH	1	DICOS
(4010,101F)	Route Segment End Location ID	Route​Segment​End​Location​ID	SH	1	DICOS
(4010,1020)	Route Segment Location ID Type	Route​Segment​Location​ID​Type	CS	1	DICOS
(4010,1021)	Abort Reason	Abort​Reason	CS	1-n	DICOS
(4010,1023)	Volume of PTO	Volume​Of​PTO	FL	1	DICOS
(4010,1024)	Abort Flag	Abort​Flag	CS	1	DICOS
(4010,1025)	Route Segment Start Time	Route​Segment​Start​Time	DT	1	DICOS
(4010,1026)	Route Segment End Time	Route​Segment​End​Time	DT	1	DICOS
(4010,1027)	TDR Type	TDR​Type	CS	1	DICOS
(4010,1028)	International Route Segment	International​Route​Segment	CS	1	DICOS
(4010,1029)	Threat Detection Algorithm and Version	Threat​Detection​Algorithmand​Version	LO	1-n	DICOS
(4010,102A)	Assigned Location	Assigned​Location	SH	1	DICOS
(4010,102B)	Alarm Decision Time	Alarm​Decision​Time	DT	1	DICOS
(4010,1031)	Alarm Decision	Alarm​Decision	CS	1	DICOS
(4010,1033)	Number of Total Objects	Number​Of​Total​Objects	US	1	DICOS
(4010,1034)	Number of Alarm Objects	Number​Of​Alarm​Objects	US	1	DICOS
(4010,1037)	PTO Representation Sequence	PTO​Representation​Sequence	SQ	1	DICOS
(4010,1038)	ATD Assessment Sequence	ATD​Assessment​Sequence	SQ	1	DICOS
(4010,1039)	TIP Type	TIP​Type	CS	1	DICOS
(4010,103A)	DICOS Version	DICOS​Version	CS	1	DICOS
(4010,1041)	OOI Owner Creation Time	OOI​Owner​Creation​Time	DT	1	DICOS
(4010,1042)	OOI Type	OOI​Type	CS	1	DICOS
(4010,1043)	OOI Size	OOI​Size	FL	3	DICOS
(4010,1044)	Acquisition Status	Acquisition​Status	CS	1	DICOS
(4010,1045)	Basis Materials Code Sequence	Basis​Materials​Code​Sequence	SQ	1	DICOS
(4010,1046)	Phantom Type	Phantom​Type	CS	1	DICOS
(4010,1047)	OOI Owner Sequence	OOI​Owner​Sequence	SQ	1	DICOS
(4010,1048)	Scan Type	Scan​Type	CS	1	DICOS
(4010,1051)	Itinerary ID	Itinerary​ID	LO	1	DICOS
(4010,1052)	Itinerary ID Type	Itinerary​ID​Type	SH	1	DICOS
(4010,1053)	Itinerary ID Assigning Authority	Itinerary​ID​Assigning​Authority	LO	1	DICOS
(4010,1054)	Route ID	Route​ID	SH	1	DICOS
(4010,1055)	Route ID Assigning Authority	Route​ID​Assigning​Authority	SH	1	DICOS
(4010,1056)	Inbound Arrival Type	Inbound​Arrival​Type	CS	1	DICOS
(4010,1058)	Carrier ID	Carrier​ID	SH	1	DICOS
(4010,1059)	Carrier ID Assigning Authority	Carrier​ID​Assigning​Authority	CS	1	DICOS
(4010,1060)	Source Orientation	Source​Orientation	FL	3	DICOS
(4010,1061)	Source Position	Source​Position	FL	3	DICOS
(4010,1062)	Belt Height	Belt​Height	FL	1	DICOS
(4010,1064)	Algorithm Routing Code Sequence	Algorithm​Routing​Code​Sequence	SQ	1	DICOS
(4010,1067)	Transport Classification	Transport​Classification	CS	1	DICOS
(4010,1068)	OOI Type Descriptor	OOI​Type​Descriptor	LT	1	DICOS
(4010,1069)	Total Processing Time	Total​Processing​Time	FL	1	DICOS
(4010,106C)	Detector Calibration Data	Detector​Calibration​Data	OB	1	DICOS
(4010,106D)	Additional Screening Performed	Additional​Screening​Performed	CS	1	DICOS
(4010,106E)	Additional Inspection Selection Criteria	Additional​Inspection​Selection​Criteria	CS	1	DICOS
(4010,106F)	Additional Inspection Method Sequence	Additional​Inspection​Method​Sequence	SQ	1	DICOS
(4010,1070)	AIT Device Type	AIT​Device​Type	CS	1	DICOS
(4010,1071)	QR Measurements Sequence	QR​Measurements​Sequence	SQ	1	DICOS
(4010,1072)	Target Material Sequence	Target​Material​Sequence	SQ	1	DICOS
(4010,1073)	SNR Threshold	SNR​Threshold	FD	1	DICOS
(4010,1075)	Image Scale Representation	Image​Scale​Representation	DS	1	DICOS
(4010,1076)	Referenced PTO Sequence	Referenced​PTO​Sequence	SQ	1	DICOS
(4010,1077)	Referenced TDR Instance Sequence	Referenced​TDR​Instance​Sequence	SQ	1	DICOS
(4010,1078)	PTO Location Description	PTO​Location​Description	ST	1	DICOS
(4010,1079)	Anomaly Locator Indicator Sequence	Anomaly​Locator​Indicator​Sequence	SQ	1	DICOS
(4010,107A)	Anomaly Locator Indicator	Anomaly​Locator​Indicator	FL	3	DICOS
(4010,107B)	PTO Region Sequence	PTO​Region​Sequence	SQ	1	DICOS
(4010,107C)	Inspection Selection Criteria	Inspection​Selection​Criteria	CS	1	DICOS
(4010,107D)	Secondary Inspection Method Sequence	Secondary​Inspection​Method​Sequence	SQ	1	DICOS
(4010,107E)	PRCS to RCS Orientation	PRCS​To​RCS​Orientation	DS	6	DICOS
(4FFE,0001)	MAC Parameters Sequence	MAC​Parameters​Sequence	SQ	1
(50xx,0005)	Curve Dimensions	Curve​Dimensions	US	1	RET (2004)
(50xx,0010)	Number of Points	Number​Of​Points	US	1	RET (2004)
(50xx,0020)	Type of Data	Type​Of​Data	CS	1	RET (2004)
(50xx,0022)	Curve Description	Curve​Description	LO	1	RET (2004)
(50xx,0030)	Axis Units	Axis​Units	SH	1-n	RET (2004)
(50xx,0040)	Axis Labels	Axis​Labels	SH	1-n	RET (2004)
(50xx,0103)	Data Value Representation	Data​Value​Representation	US	1	RET (2004)
(50xx,0104)	Minimum Coordinate Value	Minimum​Coordinate​Value	US	1-n	RET (2004)
(50xx,0105)	Maximum Coordinate Value	Maximum​Coordinate​Value	US	1-n	RET (2004)
(50xx,0106)	Curve Range	Curve​Range	SH	1-n	RET (2004)
(50xx,0110)	Curve Data Descriptor	Curve​Data​Descriptor	US	1-n	RET (2004)
(50xx,0112)	Coordinate Start Value	Coordinate​Start​Value	US	1-n	RET (2004)
(50xx,0114)	Coordinate Step Value	Coordinate​Step​Value	US	1-n	RET (2004)
(50xx,1001)	Curve Activation Layer	Curve​Activation​Layer	CS	1	RET (2004)
(50xx,2000)	Audio Type	Audio​Type	US	1	RET (2004)
(50xx,2002)	Audio Sample Format	Audio​Sample​Format	US	1	RET (2004)
(50xx,2004)	Number of Channels	Number​Of​Channels	US	1	RET (2004)
(50xx,2006)	Number of Samples	Number​Of​Samples	UL	1	RET (2004)
(50xx,2008)	Sample Rate	Sample​Rate	UL	1	RET (2004)
(50xx,200A)	Total Time	Total​Time	UL	1	RET (2004)
(50xx,200C)	Audio Sample Data	Audio​Sample​Data	OB or OW	1	RET (2004)
(50xx,200E)	Audio Comments	Audio​Comments	LT	1	RET (2004)
(50xx,2500)	Curve Label	Curve​Label	LO	1	RET (2004)
(50xx,2600)	Curve Referenced Overlay Sequence	Curve​Referenced​Overlay​Sequence	SQ	1	RET (2004)
(50xx,2610)	Curve Referenced Overlay Group	Curve​Referenced​Overlay​Group	US	1	RET (2004)
(50xx,3000)	Curve Data	Curve​Data	OB or OW	1	RET (2004)
(5200,9229)	Shared Functional Groups Sequence	Shared​Functional​Groups​Sequence	SQ	1
(5200,9230)	Per-frame Functional Groups Sequence	Per​Frame​Functional​Groups​Sequence	SQ	1
(5400,0100)	Waveform Sequence	Waveform​Sequence	SQ	1
(5400,0110)	Channel Minimum Value	Channel​Minimum​Value	OB or OW	1
(5400,0112)	Channel Maximum Value	Channel​Maximum​Value	OB or OW	1
(5400,1004)	Waveform Bits Allocated	Waveform​Bits​Allocated	US	1
(5400,1006)	Waveform Sample Interpretation	Waveform​Sample​Interpretation	CS	1
(5400,100A)	Waveform Padding Value	Waveform​Padding​Value	OB or OW	1
(5400,1010)	Waveform Data	Waveform​Data	OB or OW	1
(5600,0010)	First Order Phase Correction Angle	First​Order​Phase​Correction​Angle	OF	1
(5600,0020)	Spectroscopy Data	Spectroscopy​Data	OF	1
(60xx,0010)	Overlay Rows	Overlay​Rows	US	1
(60xx,0011)	Overlay Columns	Overlay​Columns	US	1
(60xx,0012)	Overlay Planes	Overlay​Planes	US	1	RET (2007)
(60xx,0015)	Number of Frames in Overlay	Number​Of​Frames​In​Overlay	IS	1
(60xx,0022)	Overlay Description	Overlay​Description	LO	1
(60xx,0040)	Overlay Type	Overlay​Type	CS	1
(60xx,0045)	Overlay Subtype	Overlay​Subtype	LO	1
(60xx,0050)	Overlay Origin	Overlay​Origin	SS	2
(60xx,0051)	Image Frame Origin	Image​Frame​Origin	US	1
(60xx,0052)	Overlay Plane Origin	Overlay​Plane​Origin	US	1	RET (2007)
(60xx,0060)	Overlay Compression Code	Overlay​Compression​Code	CS	1	RET
(60xx,0061)	Overlay Compression Originator	Overlay​Compression​Originator	SH	1	RET (2007)
(60xx,0062)	Overlay Compression Label	Overlay​Compression​Label	SH	1	RET (2007)
(60xx,0063)	Overlay Compression Description	Overlay​Compression​Description	CS	1	RET (2007)
(60xx,0066)	Overlay Compression Step Pointers	Overlay​Compression​Step​Pointers	AT	1-n	RET (2007)
(60xx,0068)	Overlay Repeat Interval	Overlay​Repeat​Interval	US	1	RET (2007)
(60xx,0069)	Overlay Bits Grouped	Overlay​Bits​Grouped	US	1	RET (2007)
(60xx,0100)	Overlay Bits Allocated	Overlay​Bits​Allocated	US	1
(60xx,0102)	Overlay Bit Position	Overlay​Bit​Position	US	1
(60xx,0110)	Overlay Format	Overlay​Format	CS	1	RET
(60xx,0200)	Overlay Location	Overlay​Location	US	1	RET
(60xx,0800)	Overlay Code Label	Overlay​Code​Label	CS	1-n	RET (2007)
(60xx,0802)	Overlay Number of Tables	Overlay​Number​Of​Tables	US	1	RET (2007)
(60xx,0803)	Overlay Code Table Location	Overlay​Code​Table​Location	AT	1-n	RET (2007)
(60xx,0804)	Overlay Bits For Code Word	Overlay​Bits​For​Code​Word	US	1	RET (2007)
(60xx,1001)	Overlay Activation Layer	Overlay​Activation​Layer	CS	1
(60xx,1100)	Overlay Descriptor - Gray	Overlay​Descriptor​Gray	US	1	RET (1996)
(60xx,1101)	Overlay Descriptor - Red	Overlay​Descriptor​Red	US	1	RET (1996)
(60xx,1102)	Overlay Descriptor - Green	Overlay​Descriptor​Green	US	1	RET (1996)
(60xx,1103)	Overlay Descriptor - Blue	Overlay​Descriptor​Blue	US	1	RET (1996)
(60xx,1200)	Overlays - Gray	Overlays​Gray	US	1-n	RET (1996)
(60xx,1201)	Overlays - Red	Overlays​Red	US	1-n	RET (1996)
(60xx,1202)	Overlays - Green	Overlays​Green	US	1-n	RET (1996)
(60xx,1203)	Overlays - Blue	Overlays​Blue	US	1-n	RET (1996)
(60xx,1301)	ROI Area	ROI​Area	IS	1
(60xx,1302)	ROI Mean	ROI​Mean	DS	1
(60xx,1303)	ROI Standard Deviation	ROI​Standard​Deviation	DS	1
(60xx,1500)	Overlay Label	Overlay​Label	LO	1
(60xx,3000)	Overlay Data	Overlay​Data	OB or OW	1
(60xx,4000)	Overlay Comments	Overlay​Comments	LT	1	RET
(7FE0,0001)	Extended Offset Table	Extended​Offset​Table	OV	1
(7FE0,0002)	Extended Offset Table Lengths	Extended​Offset​Table​Lengths	OV	1
(7FE0,0008)	Float Pixel Data	Float​​Pixel​​Data	OF	1
(7FE0,0009)	Double Float Pixel Data	Double​Float​Pixel​​Data	OD	1
(7FE0,0010)	Pixel Data	Pixel​Data	OB or OW	1
(7FE0,0020)	Coefficients SDVN	Coefficients​SDVN	OW	1	RET (2007)
(7FE0,0030)	Coefficients SDHN	Coefficients​SDHN	OW	1	RET (2007)
(7FE0,0040)	Coefficients SDDN	Coefficients​SDDN	OW	1	RET (2007)
(7Fxx,0010)	Variable Pixel Data	Variable​Pixel​Data	OB or OW	1	RET (2007)
(7Fxx,0011)	Variable Next Data Group	Variable​Next​Data​Group	US	1	RET (2007)
(7Fxx,0020)	Variable Coefficients SDVN	Variable​Coefficients​SDVN	OW	1	RET (2007)
(7Fxx,0030)	Variable Coefficients SDHN	Variable​Coefficients​SDHN	OW	1	RET (2007)
(7Fxx,0040)	Variable Coefficients SDDN	Variable​Coefficients​SDDN	OW	1	RET (2007)
(FFFA,FFFA)	Digital Signatures Sequence	Digital​Signatures​Sequence	SQ	1
(FFFC,FFFC)	Data Set Trailing Padding	Data​Set​Trailing​Padding	OB	1
(FFFE,E000)	Item	Item	See Note 2	1
(FFFE,E00D)	Item Delimitation Item	Item​Delimitation​Item	See Note 2	1
(FFFE,E0DD)	Sequence Delimitation Item	Sequence​Delimitation​Item	See Note 2	1

Note
1. Tag (0040,A170) was defined as Observation Class with a VR of "CS" in the Frozen Draft version of Supplement 23 "Structured Reporting" (from November 20, 1997). Implementers of the Standard should be warned that old objects of the associated SOP Classes exist and that they use this VR instead of "SQ". In particular, when reading objects with Implicit VR Little Endian transfer syntax, this inconsistency might result in parsing errors if not handled appropriately.
2. The VR for Data Elements, Item (FFFE,E000), Item Delimitation Item (FFFE,E00D), and Sequence Delimitation Item (FFFE,E0DD) do not exist. See PS3.5 for explanation.
3. For some Data Elements, no Name or Keyword or VR or VM is specified; these are "placeholders" that are not assigned but will not be reused.
7 Registry of DICOM File Meta Elements
This section specifies the File Meta Elements needed to support the formatting of the File Meta Information of the DICOM File Format (see PS3.10).
Table 7-1. Registry of DICOM File Meta Elements
Tag	Name	Keyword	VR	VM
(0002,0000)	File Meta Information Group Length	File​Meta​Information​Group​Length	UL	1
(0002,0001)	File Meta Information Version	File​Meta​Information​Version	OB	1
(0002,0002)	Media Storage SOP Class UID	Media​Storage​SOP​Class​UID	UI	1
(0002,0003)	Media Storage SOP Instance UID	Media​Storage​SOP​Instance​UID	UI	1
(0002,0010)	Transfer Syntax UID	Transfer​Syntax​UID	UI	1
(0002,0012)	Implementation Class UID	Implementation​Class​UID	UI	1
(0002,0013)	Implementation Version Name	Implementation​Version​Name	SH	1
(0002,0016)	Source Application Entity Title	Source​Application​Entity​Title	AE	1
(0002,0017)	Sending Application Entity Title	Sending​Application​Entity​Title	AE	1
(0002,0018)	Receiving Application Entity Title	Receiving​Application​Entity​Title	AE	1
(0002,0100)	Private Information Creator UID	Private​Information​Creator​UID	UI	1
(0002,0102)	Private Information	Private​Information	OB	1

8 Registry of DICOM Directory Structuring Elements
For retired elements, the last edition of the Standard that documented the element is indicated in parentheses.
Table 8-1. Registry of DICOM Directory Structuring Elements
Tag	Name	Keyword	VR	VM
(0004,1130)	File-set ID	File​Set​ID	CS	1
(0004,1141)	File-set Descriptor File ID	File​Set​Descriptor​File​ID	CS	1-8
(0004,1142)	Specific Character Set of File-set Descriptor File	Specific​Character​Set​Of​File​Set​Descriptor​File	CS	1
(0004,1200)	Offset of the First Directory Record of the Root Directory Entity	Offset​Of​The​First​Directory​Record​Of​The​Root​Directory​Entity	UL	1
(0004,1202)	Offset of the Last Directory Record of the Root Directory Entity	Offset​Of​The​Last​Directory​Record​Of​The​Root​Directory​Entity	UL	1
(0004,1212)	File-set Consistency Flag	File​Set​Consistency​Flag	US	1
(0004,1220)	Directory Record Sequence	Directory​Record​Sequence	SQ	1
(0004,1400)	Offset of the Next Directory Record	Offset​Of​The​Next​Directory​Record	UL	1
(0004,1410)	Record In-use Flag	Record​In​Use​Flag	US	1
(0004,1420)	Offset of Referenced Lower-Level Directory Entity	Offset​Of​Referenced​Lower​Level​Directory​Entity	UL	1
(0004,1430)	Directory Record Type	Directory​Record​Type	CS	1
(0004,1432)	Private Record UID	Private​Record​UID	UI	1
(0004,1500)	Referenced File ID	Referenced​File​ID	CS	1-8
(0004,1504)	MRDR Directory Record Offset	MRDR​Directory​Record​Offset	UL	1	RET (2004)
(0004,1510)	Referenced SOP Class UID in File	Referenced​SOP​Class​UID​In​File	UI	1
(0004,1511)	Referenced SOP Instance UID in File	Referenced​SOP​Instance​UID​In​File	UI	1
(0004,1512)	Referenced Transfer Syntax UID in File	Referenced​Transfer​Syntax​UID​In​File	UI	1
(0004,151A)	Referenced Related General SOP Class UID in File	Referenced​Related​General​SOP​Class​UID​In​File	UI	1-n
(0004,1600)	Number of References	Number​Of​References	UL	1	RET (2004)

A Registry of DICOM Unique Identifiers (UIDs) (Normative)
Table A-1 lists the UID values that are registered and used throughout the Parts of the DICOM Standard. This central registry ensures that when additional UIDs are assigned, non-duplicate values are assigned. For retired UIDs, the edition of the Standard in parentheses is the edition in which the item last appeared before it was retired.
Table A-1. UID Values
UID Value	UID Name	UID Type	Part
1.2.840.10008.1.​1	Verification SOP Class	SOP Class	PS3.4
1.2.840.10008.1.​2	Implicit VR Little Endian: Default Transfer Syntax for DICOM	Transfer Syntax	PS3.5
1.2.840.10008.1.​2.​1	Explicit VR Little Endian	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​1.​99	Deflated Explicit VR Little Endian	Transfer Syntax	PS3.5
1.2.840.10008.1.​2.​2	Explicit VR Big Endian (Retired)	Transfer Syntax	PS3.5 (2011)
1.2.840.10008.1.2.​4.​50	JPEG Baseline (Process 1): Default Transfer Syntax for Lossy JPEG 8 Bit Image Compression	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​51	JPEG Extended (Process 2 & 4): Default Transfer Syntax for Lossy JPEG 12 Bit Image Compression (Process 4 only)	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​52	JPEG Extended (Process 3 & 5) (Retired)	Transfer Syntax	PS3.5 (2001)
1.2.840.10008.1.2.​4.​53	JPEG Spectral Selection, Non-Hierarchical (Process 6 & 8) (Retired)	Transfer Syntax	PS3.5 (2001)
1.2.840.10008.1.2.​4.​54	JPEG Spectral Selection, Non-Hierarchical (Process 7 & 9) (Retired)	Transfer Syntax	PS3.5 (2001)
1.2.840.10008.1.2.​4.​55	JPEG Full Progression, Non-Hierarchical (Process 10 & 12) (Retired)	Transfer Syntax	PS3.5 (2001)
1.2.840.10008.1.2.​4.​56	JPEG Full Progression, Non-Hierarchical (Process 11 & 13) (Retired)	Transfer Syntax	PS3.5 (2001)
1.2.840.10008.1.2.​4.​57	JPEG Lossless, Non-Hierarchical (Process 14)	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​58	JPEG Lossless, Non-Hierarchical (Process 15) (Retired)	Transfer Syntax	PS3.5 (2001)
1.2.840.10008.1.2.​4.​59	JPEG Extended, Hierarchical (Process 16 & 18) (Retired)	Transfer Syntax	PS3.5 (2001)
1.2.840.10008.1.2.​4.​60	JPEG Extended, Hierarchical (Process 17 & 19) (Retired)	Transfer Syntax	PS3.5 (2001)
1.2.840.10008.1.2.​4.​61	JPEG Spectral Selection, Hierarchical (Process 20 & 22) (Retired)	Transfer Syntax	PS3.5 (2001)
1.2.840.10008.1.2.​4.​62	JPEG Spectral Selection, Hierarchical (Process 21 & 23) (Retired)	Transfer Syntax	PS3.5 (2001)
1.2.840.10008.1.2.​4.​63	JPEG Full Progression, Hierarchical (Process 24 & 26) (Retired)	Transfer Syntax	PS3.5 (2001)
1.2.840.10008.1.2.​4.​64	JPEG Full Progression, Hierarchical (Process 25 & 27) (Retired)	Transfer Syntax	PS3.5 (2001)
1.2.840.10008.1.2.​4.​65	JPEG Lossless, Hierarchical (Process 28) (Retired)	Transfer Syntax	PS3.5 (2001)
1.2.840.10008.1.2.​4.​66	JPEG Lossless, Hierarchical (Process 29) (Retired)	Transfer Syntax	PS3.5 (2001)
1.2.840.10008.1.2.​4.​70	JPEG Lossless, Non-Hierarchical, First-Order Prediction (Process 14 [Selection Value 1]): Default Transfer Syntax for Lossless JPEG Image Compression	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​80	JPEG-LS Lossless Image Compression	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​81	JPEG-LS Lossy (Near-Lossless) Image Compression	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​90	JPEG 2000 Image Compression (Lossless Only)	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​91	JPEG 2000 Image Compression	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​92	JPEG 2000 Part 2 Multi-component Image Compression (Lossless Only)	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​93	JPEG 2000 Part 2 Multi-component Image Compression	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​94	JPIP Referenced	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​95	JPIP Referenced Deflate	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​100	MPEG2 Main Profile / Main Level	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​101	MPEG2 Main Profile / High Level	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​102	MPEG-4 AVC/H.264 High Profile / Level 4.1	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​103	MPEG-4 AVC/H.264 BD-compatible High Profile / Level 4.1	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​104	MPEG-4 AVC/H.264 High Profile / Level 4.2 For 2D Video	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​105	MPEG-4 AVC/H.264 High Profile / Level 4.2 For 3D Video	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.​106	MPEG-4 AVC/H.264 Stereo High Profile / Level 4.2	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.107	HEVC/H.265 Main Profile / Level 5.1	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​4.108	HEVC/H.265 Main 10 Profile / Level 5.1	Transfer Syntax	PS3.5
1.2.840.10008.1.​2.​5	RLE Lossless	Transfer Syntax	PS3.5
1.2.840.10008.1.2.​6.​1	RFC 2557 MIME encapsulation (Retired)	Transfer Syntax	PS3.10 (2018b)
1.2.840.10008.1.2.​6.​2	XML Encoding (Retired)	Transfer Syntax	PS3.10 (2018b)
1.2.840.10008.1.​3.​10	Media Storage Directory Storage	SOP Class	PS3.4
1.2.840.10008.1.4.​1.​1	Talairach Brain Atlas Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​2	SPM2 T1 Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​3	SPM2 T2 Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​4	SPM2 PD Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​5	SPM2 EPI Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​6	SPM2 FIL T1 Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​7	SPM2 PET Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​8	SPM2 TRANSM Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​9	SPM2 SPECT Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​10	SPM2 GRAY Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​11	SPM2 WHITE Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​12	SPM2 CSF Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​13	SPM2 BRAINMASK Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​14	SPM2 AVG305T1 Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​15	SPM2 AVG152T1 Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​16	SPM2 AVG152T2 Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​17	SPM2 AVG152PD Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​1.​18	SPM2 SINGLESUBJT1 Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​2.​1	ICBM 452 T1 Frame of Reference	Well-known frame of reference
1.2.840.10008.1.4.​2.​2	ICBM Single Subject MRI Frame of Reference	Well-known frame of reference
1.2.840.10008.1.​5.​1	Hot Iron Color Palette SOP Instance	Well-known SOP Instance	PS 3.6
1.2.840.10008.1.​5.​2	PET Color Palette SOP Instance	Well-known SOP Instance	PS 3.6
1.2.840.10008.1.​5.​3	Hot Metal Blue Color Palette SOP Instance	Well-known SOP Instance	PS 3.6
1.2.840.10008.1.​5.​4	PET 20 Step Color Palette SOP Instance	Well-known SOP Instance	PS 3.6
1.2.840.10008.1.​5.​5	Spring Color Palette SOP Instance	Well-known SOP Instance	PS 3.6
1.2.840.10008.1.​5.​6	Summer Color Palette SOP Instance	Well-known SOP Instance	PS 3.6
1.2.840.10008.1.​5.​7	Fall Color Palette SOP Instance	Well-known SOP Instance	PS 3.6
1.2.840.10008.1.​5.​8	Winter Color Palette SOP Instance	Well-known SOP Instance	PS 3.6
1.2.840.10008.1.​9	Basic Study Content Notification SOP Class (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.1.​20	Papyrus 3 Implicit VR Little Endian (Retired)	Transfer Syntax	(2015c)
1.2.840.10008.1.​20.​1	Storage Commitment Push Model SOP Class	SOP Class	PS3.4
1.2.840.10008.1.20.​1.​1	Storage Commitment Push Model SOP Instance	Well-known SOP Instance	PS3.4
1.2.840.10008.1.​20.​2	Storage Commitment Pull Model SOP Class (Retired)	SOP Class	PS3.4 (2001)
1.2.840.10008.1.20.​2.​1	Storage Commitment Pull Model SOP Instance (Retired)	Well-known SOP Instance	PS3.4 (2001)
1.2.840.10008.1.​40	Procedural Event Logging SOP Class	SOP Class	PS3.4
1.2.840.10008.1.​40.​1	Procedural Event Logging SOP Instance	Well-known SOP Instance	PS3.4
1.2.840.10008.1.​42	Substance Administration Logging SOP Class	SOP Class	PS3.4
1.2.840.10008.1.​42.​1	Substance Administration Logging SOP Instance	Well-known SOP Instance	PS3.4
1.2.840.10008.2.​6.​1	DICOM UID Registry	DICOM UIDs as a Coding Scheme	PS 3.6
1.2.840.10008.2.​16.​4	DICOM Controlled Terminology	Coding Scheme	PS3.16
1.2.840.10008.2.​16.​5	Adult Mouse Anatomy Ontology	Coding Scheme	PS3.16
1.2.840.10008.2.​16.​6	Uberon Ontology	Coding Scheme	PS3.16
1.2.840.10008.2.​16.​7	Integrated Taxonomic Information System (ITIS) Taxonomic Serial Number (TSN)	Coding Scheme	PS3.16
1.2.840.10008.2.​16.​8	Mouse Genome Initiative (MGI)	Coding Scheme	PS3.16
1.2.840.10008.2.​16.​9	PubChem Compound CID	Coding Scheme	PS3.16
1.2.840.10008.2.​16.​10	ICD-11	Coding Scheme	PS3.16
1.2.840.10008.2.​16.​11	New York University Melanoma Clinical Cooperative Group	Coding Scheme	PS3.16
1.2.840.10008.2.​16.​12	Mayo Clinic Non-radiological Images Specific Body Structure Anatomical Surface Region Guide	Coding Scheme	PS3.16
1.2.840.10008.2.​16.​13	Image Biomarker Standardisation Initiative	Coding Scheme	PS3.16
1.2.840.10008.2.​16.​14	Radiomics Ontology	Coding Scheme	PS3.16
1.2.840.10008.3.1.​1.​1	DICOM Application Context Name	Application Context Name	PS3.7
1.2.840.10008.3.1.2.​1.​1	Detached Patient Management SOP Class (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.3.1.2.​1.​4	Detached Patient Management Meta SOP Class (Retired)	Meta SOP Class	PS3.4 (2004)
1.2.840.10008.3.1.2.​2.​1	Detached Visit Management SOP Class (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.3.1.2.​3.​1	Detached Study Management SOP Class (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.3.1.2.​3.​2	Study Component Management SOP Class (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.3.1.2.​3.​3	Modality Performed Procedure Step SOP Class	SOP Class	PS3.4
1.2.840.10008.3.1.2.​3.​4	Modality Performed Procedure Step Retrieve SOP Class	SOP Class	PS3.4
1.2.840.10008.3.1.2.​3.​5	Modality Performed Procedure Step Notification SOP Class	SOP Class	PS3.4
1.2.840.10008.3.1.2.​5.​1	Detached Results Management SOP Class (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.3.1.2.​5.​4	Detached Results Management Meta SOP Class (Retired)	Meta SOP Class	PS3.4 (2004)
1.2.840.10008.3.1.2.​5.​5	Detached Study Management Meta SOP Class (Retired)	Meta SOP Class	PS3.4 (2004)
1.2.840.10008.3.1.2.​6.​1	Detached Interpretation Management SOP Class (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.4.​2	Storage Service Class	Service Class	PS3.4
1.2.840.10008.5.1.​1.​1	Basic Film Session SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.​1.​2	Basic Film Box SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.​1.​4	Basic Grayscale Image Box SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.1.​4.​1	Basic Color Image Box SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.1.​4.​2	Referenced Image Box SOP Class (Retired)	SOP Class	PS3.4 (1998)
1.2.840.10008.5.1.​1.​9	Basic Grayscale Print Management Meta SOP Class	Meta SOP Class	PS3.4
1.2.840.10008.5.1.1.​9.​1	Referenced Grayscale Print Management Meta SOP Class (Retired)	Meta SOP Class	PS3.4 (1998)
1.2.840.10008.5.1.​1.​14	Print Job SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.​1.​15	Basic Annotation Box SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.​1.​16	Printer SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.1.​16.​376	Printer Configuration Retrieval SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.​1.​17	Printer SOP Instance	Well-known Printer SOP Instance	PS3.4
1.2.840.10008.5.1.1.​17.​376	Printer Configuration Retrieval SOP Instance	Well-known Printer SOP Instance	PS3.4
1.2.840.10008.5.1.​1.​18	Basic Color Print Management Meta SOP Class	Meta SOP Class	PS3.4
1.2.840.10008.5.1.1.​18.​1	Referenced Color Print Management Meta SOP Class (Retired)	Meta SOP Class	PS3.4 (1998)
1.2.840.10008.5.1.​1.​22	VOI LUT Box SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.​1.​23	Presentation LUT SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.​1.​24	Image Overlay Box SOP Class (Retired)	SOP Class	PS3.4 (1998)
1.2.840.10008.5.1.1.​24.​1	Basic Print Image Overlay Box SOP Class (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.5.1.​1.​25	Print Queue SOP Instance (Retired)	Well-known Print Queue SOP Instance	PS3.4 (2004)
1.2.840.10008.5.1.​1.​26	Print Queue Management SOP Class (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.5.1.​1.​27	Stored Print Storage SOP Class (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.5.1.​1.​29	Hardcopy Grayscale Image Storage SOP Class (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.5.1.​1.​30	Hardcopy Color Image Storage SOP Class (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.5.1.​1.​31	Pull Print Request SOP Class (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.5.1.​1.​32	Pull Stored Print Management Meta SOP Class (Retired)	Meta SOP Class	PS3.4 (2004)
1.2.840.10008.5.1.​1.​33	Media Creation Management SOP Class UID	SOP Class	PS3.4
1.2.840.10008.5.1.1.40	Display System SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.1.40.1	Display System SOP Instance	Well-known SOP Instance	PS3.4
1.2.840.10008.5.1.4.1.​1.​1	Computed Radiography Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​1.​1	Digital X-Ray Image Storage - For Presentation	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.1.​1.​1	Digital X-Ray Image Storage - For Processing	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​1.​2	Digital Mammography X-Ray Image Storage - For Presentation	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.1.​2.​1	Digital Mammography X-Ray Image Storage - For Processing	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​1.​3	Digital Intra-Oral X-Ray Image Storage - For Presentation	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.1.​3.​1	Digital Intra-Oral X-Ray Image Storage - For Processing	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.​1.​2	CT Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​2.​1	Enhanced CT Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​2.​2	Legacy Converted Enhanced CT Image Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.​1.​3	Ultrasound Multi-frame Image Storage (Retired)	SOP Class	PS3.4 (1993)
1.2.840.10008.​5.​1.​4.​1.​1.​3.​1	Ultrasound Multi-frame Image Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.​1.​4	MR Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​4.​1	Enhanced MR Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​4.​2	MR Spectroscopy Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​4.​3	Enhanced MR Color Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​4.​4	Legacy Converted Enhanced MR Image Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.​1.​5	Nuclear Medicine Image Storage (Retired)	SOP Class	PS3.4 (1993)
1.2.840.10008.5.1.4.1.​1.​6	Ultrasound Image Storage (Retired)	SOP Class	PS3.4 (1993)
1.2.840.10008.​5.​1.​4.​1.​1.​6.​1	Ultrasound Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​6.​2	Enhanced US Volume Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.​1.​7	Secondary Capture Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​7.​1	Multi-frame Single Bit Secondary Capture Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​7.​2	Multi-frame Grayscale Byte Secondary Capture Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​7.​3	Multi-frame Grayscale Word Secondary Capture Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​7.​4	Multi-frame True Color Secondary Capture Image Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.​1.​8	Standalone Overlay Storage (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.5.1.4.1.​1.​9	Standalone Curve Storage (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.​5.​1.​4.​1.​1.​9.​1	Waveform Storage - Trial (Retired)	SOP Class	PS3.4 (2007)
1.2.840.10008.​5.​1.​4.​1.​1.9.​1.​1	12-lead ECG Waveform Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.9.​1.​2	General ECG Waveform Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.9.​1.​3	Ambulatory ECG Waveform Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.9.​2.​1	Hemodynamic Waveform Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.9.​3.​1	Cardiac Electrophysiology Waveform Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.9.​4.​1	Basic Voice Audio Waveform Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.9.​4.​2	General Audio Waveform Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.9.​5.​1	Arterial Pulse Waveform Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.9.​6.​1	Respiratory Waveform Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.​1.​10	Standalone Modality LUT Storage (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.5.1.4.1.​1.​11	Standalone VOI LUT Storage (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.​5.​1.​4.​1.​1.​11.​1	Grayscale Softcopy Presentation State Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​11.​2	Color Softcopy Presentation State Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​11.​3	Pseudo-Color Softcopy Presentation State Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​11.​4	Blending Softcopy Presentation State Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​11.​5	XA/XRF Grayscale Softcopy Presentation State Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​11.​6	Grayscale Planar MPR Volumetric Presentation State Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​11.​7	Compositing Planar MPR Volumetric Presentation State Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​11.​8	Advanced Blending Presentation State Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​11.​9	Volume Rendering Volumetric Presentation State Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​11.​10	Segmented Volume Rendering Volumetric Presentation State Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​11.​11	Multiple Volume Rendering Volumetric Presentation State Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​12.​1	X-Ray Angiographic Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.12.​1.​1	Enhanced XA Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​12.​2	X-Ray Radiofluoroscopic Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.12.​2.​1	Enhanced XRF Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​12.​3	X-Ray Angiographic Bi-Plane Image Storage (Retired)	SOP Class	PS3.4 (1998)
1.2.840.10008.​5.​1.​4.​1.​1.​12.​77	(Retired)	SOP Class	(2015c)
1.2.840.10008.​5.​1.​4.​1.​1.13.​1.​1	X-Ray 3D Angiographic Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.13.​1.​2	X-Ray 3D Craniofacial Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.13.​1.​3	Breast Tomosynthesis Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.1.4.1.1.13.1.4	Breast Projection X-Ray Image Storage - For Presentation	SOP Class	PS3.4
1.2.840.10008.​5.1.4.1.1.13.1.5	Breast Projection X-Ray Image Storage - For Processing	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​14.​1	Intravascular Optical Coherence Tomography Image Storage - For Presentation	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​14.​2	Intravascular Optical Coherence Tomography Image Storage - For Processing	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.​1.​20	Nuclear Medicine Image Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.​1.​30	Parametric Map Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.​1.​40	(Retired)	SOP Class	(2015c)
1.2.840.10008.5.1.4.1.​1.​66	Raw Data Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​66.​1	Spatial Registration Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​66.​2	Spatial Fiducials Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​66.​3	Deformable Spatial Registration Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​66.​4	Segmentation Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​66.​5	Surface Segmentation Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​66.​6	Tractography Results Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.​1.​67	Real World Value Mapping Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​68.​1	Surface Scan Mesh Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​68.​2	Surface Scan Point Cloud Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​77.​1	VL Image Storage - Trial (Retired)	SOP Class	PS3.4 (1998)
1.2.840.10008.​5.​1.​4.​1.​1.​77.​2	VL Multi-frame Image Storage - Trial (Retired)	SOP Class	PS3.4 (1998)
1.2.840.10008.​5.​1.​4.​1.​1.77.​1.​1	VL Endoscopic Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.77.1.​1.​1	Video Endoscopic Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.77.​1.​2	VL Microscopic Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.77.1.​2.​1	Video Microscopic Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.77.​1.​3	VL Slide-Coordinates Microscopic Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.77.​1.​4	VL Photographic Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.77.1.​4.​1	Video Photographic Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.77.1.​5.​1	Ophthalmic Photography 8 Bit Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.77.1.​5.​2	Ophthalmic Photography 16 Bit Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.77.1.​5.​3	Stereometric Relationship Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.77.1.​5.​4	Ophthalmic Tomography Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.77.1.​5.​5	Wide Field Ophthalmic Photography Stereographic Projection Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.77.1.​5.​6	Wide Field Ophthalmic Photography 3D Coordinates Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.77.1.​5.​7	Ophthalmic Optical Coherence Tomography En Face Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.77.1.​5.​8	Ophthalmic Optical Coherence Tomography B-scan Volume Analysis Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.77.​1.​6	VL Whole Slide Microscopy Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​78.​1	Lensometry Measurements Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​78.​2	Autorefraction Measurements Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​78.​3	Keratometry Measurements Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​78.​4	Subjective Refraction Measurements Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​78.​5	Visual Acuity Measurements Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​78.​6	Spectacle Prescription Report Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​78.​7	Ophthalmic Axial Measurements Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​78.​8	Intraocular Lens Calculations Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​79.​1	Macular Grid Thickness and Volume Report Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​80.​1	Ophthalmic Visual Field Static Perimetry Measurements Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​81.​1	Ophthalmic Thickness Map Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​82.​1	Corneal Topography Map Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​88.​1	Text SR Storage - Trial (Retired)	SOP Class	PS3.4 (2007)
1.2.840.10008.​5.​1.​4.​1.​1.​88.​2	Audio SR Storage - Trial (Retired)	SOP Class	PS3.4 (2007)
1.2.840.10008.​5.​1.​4.​1.​1.​88.​3	Detail SR Storage - Trial (Retired)	SOP Class	PS3.4 (2007)
1.2.840.10008.​5.​1.​4.​1.​1.​88.​4	Comprehensive SR Storage - Trial (Retired)	SOP Class	PS3.4 (2007)
1.2.840.10008.​5.​1.​4.​1.​1.​88.​11	Basic Text SR Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​88.​22	Enhanced SR Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​88.​33	Comprehensive SR Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​88.​34	Comprehensive 3D SR Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​88.​35	Extensible SR Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​88.​40	Procedure Log Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​88.​50	Mammography CAD SR Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​88.​59	Key Object Selection Document Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​88.​65	Chest CAD SR Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​88.​67	X-Ray Radiation Dose SR Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​88.​68	Radiopharmaceutical Radiation Dose SR Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​88.​69	Colon CAD SR Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​88.​70	Implantation Plan SR Storage	SOP Class	PS3.4
1.2.840.10008.5.​1.​4.​1.​1.​88.​71	Acquisition Context SR Storage	SOP Class	PS3.4
1.2.840.10008.5.​1.​4.​1.​1.​88.​72	Simplified Adult Echo SR Storage	SOP Class	PS3.4
1.2.840.10008.5.​1.​4.​1.​1.​88.​73	Patient Radiation Dose SR Storage	SOP Class	PS3.4
1.2.840.10008.5.​1.​4.​1.​1.​88.​74	Planned Imaging Agent Administration SR Storage	SOP Class	PS3.4
1.2.840.10008.5.​1.​4.​1.​1.​88.​75	Performed Imaging Agent Administration SR Storage	SOP Class	PS3.4
1.2.840.10008.5.​1.​4.​1.​1.​90.​1	Content Assessment Results Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​104.​1	Encapsulated PDF Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​104.​2	Encapsulated CDA Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​104.​3	Encapsulated STL Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.​1.​128	Positron Emission Tomography Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​128.​1	Legacy Converted Enhanced PET Image Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.​1.​129	Standalone PET Curve Storage (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.5.1.4.1.​1.​130	Enhanced PET Image Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.​1.​131	Basic Structured Display Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.1.200.1	CT Defined Procedure Protocol Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.1.200.2	CT Performed Procedure Protocol Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.1.200.3	Protocol Approval Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.1.200.4	Protocol Approval Information Model - FIND	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.1.200.5	Protocol Approval Information Model - MOVE	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.1.200.6	Protocol Approval Information Model - GET	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​481.​1	RT Image Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​481.​2	RT Dose Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​481.​3	RT Structure Set Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​481.​4	RT Beams Treatment Record Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​481.​5	RT Plan Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​481.​6	RT Brachy Treatment Record Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​481.​7	RT Treatment Summary Record Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​481.​8	RT Ion Plan Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​481.​9	RT Ion Beams Treatment Record Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​481.​10	RT Physician Intent Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​481.​11	RT Segment Annotation Storage	SOP Class	PS3.4
1.2.840.10008.​5.​1.​4.​1.​1.​501.​1	DICOS CT Image Storage	SOP Class	DICOS
1.2.840.10008.​5.​1.​4.​1.​1.501.​2.​1	DICOS Digital X-Ray Image Storage - For Presentation	SOP Class	DICOS
1.2.840.10008.​5.​1.​4.​1.​1.501.​2.​2	DICOS Digital X-Ray Image Storage - For Processing	SOP Class	DICOS
1.2.840.10008.​5.​1.​4.​1.​1.​501.​3	DICOS Threat Detection Report Storage	SOP Class	DICOS
1.2.840.10008.​5.​1.​4.​1.​1.​501.​4	DICOS 2D AIT Storage	SOP Class	DICOS
1.2.840.10008.​5.​1.​4.​1.​1.​501.​5	DICOS 3D AIT Storage	SOP Class	DICOS
1.2.840.10008.​5.​1.​4.​1.​1.​501.​6	DICOS Quadrupole Resonance (QR) Storage	SOP Class	DICOS
1.2.840.10008.​5.​1.​4.​1.​1.​601.​1	Eddy Current Image Storage	SOP Class	DICONDE ASTM E2934
1.2.840.10008.​5.​1.​4.​1.​1.​601.​2	Eddy Current Multi-frame Image Storage	SOP Class	DICONDE ASTM E2934
1.2.840.10008.5.1.4.1.2.​1.​1	Patient Root Query/Retrieve Information Model - FIND	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.2.​1.​2	Patient Root Query/Retrieve Information Model - MOVE	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.2.​1.​3	Patient Root Query/Retrieve Information Model - GET	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.2.​2.​1	Study Root Query/Retrieve Information Model - FIND	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.2.​2.​2	Study Root Query/Retrieve Information Model - MOVE	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.2.​2.​3	Study Root Query/Retrieve Information Model - GET	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.2.​3.​1	Patient/Study Only Query/Retrieve Information Model - FIND (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.5.1.4.1.2.​3.​2	Patient/Study Only Query/Retrieve Information Model - MOVE (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.5.1.4.1.2.​3.​3	Patient/Study Only Query/Retrieve Information Model - GET (Retired)	SOP Class	PS3.4 (2004)
1.2.840.10008.5.1.4.1.2.​4.​2	Composite Instance Root Retrieve - MOVE	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.2.​4.​3	Composite Instance Root Retrieve - GET	SOP Class	PS3.4
1.2.840.10008.5.1.4.1.2.​5.​3	Composite Instance Retrieve Without Bulk Data - GET	SOP Class	PS3.4
1.2.840.10008.5.1.4.20.1	Defined Procedure Protocol Information Model - FIND	SOP Class	PS3.4
1.2.840.10008.5.1.4.20.2	Defined Procedure Protocol Information Model - MOVE	SOP Class	PS3.4
1.2.840.10008.5.1.4.20.3	Defined Procedure Protocol Information Model - GET	SOP Class	PS3.4
1.2.840.10008.5.1.​4.​31	Modality Worklist Information Model - FIND	SOP Class	PS3.4
1.2.840.10008.5.1.​4.​32	General Purpose Worklist Management Meta SOP Class (Retired)	Meta SOP Class	PS3.4 (2011)
1.2.840.10008.5.1.4.​32.​1	General Purpose Worklist Information Model - FIND (Retired)	SOP Class	PS3.4 (2011)
1.2.840.10008.5.1.4.​32.​2	General Purpose Scheduled Procedure Step SOP Class (Retired)	SOP Class	PS3.4 (2011)
1.2.840.10008.5.1.4.​32.​3	General Purpose Performed Procedure Step SOP Class (Retired)	SOP Class	PS3.4 (2011)
1.2.840.10008.5.1.​4.​33	Instance Availability Notification SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.4.​34.​1	RT Beams Delivery Instruction Storage - Trial (Retired)	SOP Class	PS3.4 (2009)
1.2.840.10008.5.1.4.​34.​2	RT Conventional Machine Verification - Trial (Retired)	SOP Class	PS3.4 (2009)
1.2.840.10008.5.1.4.​34.​3	RT Ion Machine Verification - Trial (Retired)	SOP Class	PS3.4 (2009)
1.2.840.10008.5.1.4.​34.​4	Unified Worklist and Procedure Step Service Class - Trial (Retired)	Service Class	PS3.4 (2009)
1.2.840.10008.5.1.4.34.​4.​1	Unified Procedure Step - Push SOP Class - Trial (Retired)	SOP Class	PS3.4 (2009)
1.2.840.10008.5.1.4.34.​4.​2	Unified Procedure Step - Watch SOP Class - Trial (Retired)	SOP Class	PS3.4 (2009)
1.2.840.10008.5.1.4.34.​4.​3	Unified Procedure Step - Pull SOP Class - Trial (Retired)	SOP Class	PS3.4 (2009)
1.2.840.10008.5.1.4.34.​4.​4	Unified Procedure Step - Event SOP Class - Trial (Retired)	SOP Class	PS3.4 (2009)
1.2.840.10008.5.1.4.​34.​5	UPS Global Subscription SOP Instance	Well-known SOP Instance	PS3.4
1.2.840.10008.5.1.4.​34.​5.​1	UPS Filtered Global Subscription SOP Instance	Well-known SOP Instance	PS3.4
1.2.840.10008.5.1.4.​34.​6	Unified Worklist and Procedure Step Service Class	Service Class	PS3.4
1.2.840.10008.5.1.4.34.​6.​1	Unified Procedure Step - Push SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.4.34.​6.​2	Unified Procedure Step - Watch SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.4.34.​6.​3	Unified Procedure Step - Pull SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.4.34.​6.​4	Unified Procedure Step - Event SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.4.​34.​7	RT Beams Delivery Instruction Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.​34.​8	RT Conventional Machine Verification	SOP Class	PS3.4
1.2.840.10008.5.1.4.​34.​9	RT Ion Machine Verification	SOP Class	PS3.4
1.2.840.10008.5.1.4.​34.​10	RT Brachy Application Setup Delivery Instruction Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.​37.​1	General Relevant Patient Information Query	SOP Class	PS3.4
1.2.840.10008.5.1.4.​37.​2	Breast Imaging Relevant Patient Information Query	SOP Class	PS3.4
1.2.840.10008.5.1.4.​37.​3	Cardiac Relevant Patient Information Query	SOP Class	PS3.4
1.2.840.10008.5.1.4.​38.​1	Hanging Protocol Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.​38.​2	Hanging Protocol Information Model - FIND	SOP Class	PS3.4
1.2.840.10008.5.1.4.​38.​3	Hanging Protocol Information Model - MOVE	SOP Class	PS3.4
1.2.840.10008.5.1.4.​38.​4	Hanging Protocol Information Model - GET	SOP Class	PS3.4
1.2.840.10008.5.1.4.​39.​1	Color Palette Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.​39.​2	Color Palette Query/Retrieve Information Model - FIND	SOP Class	PS3.4
1.2.840.10008.5.1.4.​39.​3	Color Palette Query/Retrieve Information Model - MOVE	SOP Class	PS3.4
1.2.840.10008.5.1.4.​39.​4	Color Palette Query/Retrieve Information Model - GET	SOP Class	PS3.4
1.2.840.10008.5.1.​4.​41	Product Characteristics Query SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.​4.​42	Substance Approval Query SOP Class	SOP Class	PS3.4
1.2.840.10008.5.1.4.​43.​1	Generic Implant Template Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.​43.​2	Generic Implant Template Information Model - FIND	SOP Class	PS3.4
1.2.840.10008.5.1.4.​43.​3	Generic Implant Template Information Model - MOVE	SOP Class	PS3.4
1.2.840.10008.5.1.4.​43.​4	Generic Implant Template Information Model - GET	SOP Class	PS3.4
1.2.840.10008.5.1.4.​44.​1	Implant Assembly Template Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.​44.​2	Implant Assembly Template Information Model - FIND	SOP Class	PS3.4
1.2.840.10008.5.1.4.​44.​3	Implant Assembly Template Information Model - MOVE	SOP Class	PS3.4
1.2.840.10008.5.1.4.​44.​4	Implant Assembly Template Information Model - GET	SOP Class	PS3.4
1.2.840.10008.5.1.4.​45.​1	Implant Template Group Storage	SOP Class	PS3.4
1.2.840.10008.5.1.4.​45.​2	Implant Template Group Information Model - FIND	SOP Class	PS3.4
1.2.840.10008.5.1.4.​45.​3	Implant Template Group Information Model - MOVE	SOP Class	PS3.4
1.2.840.10008.5.1.4.​45.​4	Implant Template Group Information Model - GET	SOP Class	PS3.4
1.2.840.10008.7.​1.​1	Native DICOM Model	Application Hosting Model	PS3.19
1.2.840.10008.7.​1.​2	Abstract Multi-Dimensional Image Model	Application Hosting Model	PS3.19
1.2.840.10008.8.​1.​1	DICOM Content Mapping Resource	Mapping Resource	PS3.16
1.2.840.10008.15.0.​3.​1	dicom​Device​Name	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​2	dicom​Description	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​3	dicom​Manufacturer	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​4	dicom​Manufacturer​Model​Name	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​5	dicom​Software​Version	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​6	dicom​Vendor​Data	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​7	dicomAE​Title	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​8	dicom​Network​Connection​Reference	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​9	dicom​Application​Cluster	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​10	dicom​Association​Initiator	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​11	dicom​Association​Acceptor	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​12	dicom​Hostname	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​13	dicom​Port	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​14	dicomSOP​Class	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​15	dicom​Transfer​Role	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​16	dicom​Transfer​Syntax	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​17	dicom​Primary​Device​Type	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​18	dicom​Related​Device​Reference	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​19	dicom​Preferred​CalledAE​Title	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​20	dicomTLS​Cyphersuite	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​21	dicom​Authorized​Node​Certificate​Reference	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​22	dicom​This​Node​Certificate​Reference	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​23	dicom​Installed	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​24	dicom​Station​Name	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​25	dicom​Device​Serial​Number	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​26	dicom​Institution​Name	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​27	dicom​Institution​Address	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​28	dicom​Institution​Department​Name	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​29	dicom​Issuer​OfPatientID	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​30	dicom​Preferred​CallingAE​Title	LDAP OID	PS3.15
1.2.840.10008.15.0.​3.​31	dicom​Supported​Character​Set	LDAP OID	PS3.15
1.2.840.10008.15.0.​4.​1	dicom​Configuration​Root	LDAP OID	PS3.15
1.2.840.10008.15.0.​4.​2	dicom​Devices​Root	LDAP OID	PS3.15
1.2.840.10008.15.0.​4.​3	dicom​UniqueAE​Titles​Registry​Root	LDAP OID	PS3.15
1.2.840.10008.15.0.​4.​4	dicom​Device	LDAP OID	PS3.15
1.2.840.10008.15.0.​4.​5	dicom​NetworkAE	LDAP OID	PS3.15
1.2.840.10008.15.0.​4.​6	dicom​Network​Connection	LDAP OID	PS3.15
1.2.840.10008.15.0.​4.​7	dicom​UniqueAE​Title	LDAP OID	PS3.15
1.2.840.10008.15.0.​4.​8	dicom​Transfer​Capability	LDAP OID	PS3.15
1.2.840.10008.15.​1.​1	Universal Coordinated Time	Synchronization Frame of Reference	PS3.3

Table A-2. Well-known Frames of Reference
UID Value	UID Name	Normative Reference
1.2.840.10008.1.4.​1.​1	Talairach Brain Atlas Frame of Reference	Talairach J. and Tournoux P. Co-Planar stereotactic atlas of the human brain. Stutgart: Georg Thieme Verlag, 1988.
1.2.840.10008.1.4.​1.​2	SPM2 T1 Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/templates/T1.mnc
1.2.840.10008.1.4.​1.​3	SPM2 T2 Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/templates/T2.mnc
1.2.840.10008.1.4.​1.​4	SPM2 PD Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/templates/PD.mnc
1.2.840.10008.1.4.​1.​5	SPM2 EPI Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/templates/EPI.mnc
1.2.840.10008.1.4.​1.​6	SPM2 FIL T1 Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/templates/filT1.mnc
1.2.840.10008.1.4.​1.​7	SPM2 PET Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/templates/PET.mnc
1.2.840.10008.1.4.​1.​8	SPM2 TRANSM Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/templates/Transm.mnc
1.2.840.10008.1.4.​1.​9	SPM2 SPECT Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/templates/SPECT.mnc
1.2.840.10008.1.4.​1.​10	SPM2 GRAY Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/apriori/gray.mnc
1.2.840.10008.1.4.​1.​11	SPM2 WHITE Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/apriori/white.mnc
1.2.840.10008.1.4.​1.​12	SPM2 CSF Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/apriori/csf.mnc
1.2.840.10008.1.4.​1.​13	SPM2 BRAINMASK Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/apriori/brainmask.mnc
1.2.840.10008.1.4.​1.​14	SPM2 AVG305T1 Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/canonical/avg305T1.mnc
1.2.840.10008.1.4.​1.​15	SPM2 AVG152T1 Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/canonical/avg152T1.mnc
1.2.840.10008.1.4.​1.​16	SPM2 AVG152T2 Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/canonical/avg152T2.mnc
1.2.840.10008.1.4.​1.​17	SPM2 AVG152PD Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/canonical/avg152PD.mnc
1.2.840.10008.1.4.​1.​18	SPM2 SINGLESUBJT1 Frame of Reference	http://​github.com/​spm/​spm2/​blob/​master/canonical/single_subj_T1.mnc
1.2.840.10008.1.4.​2.​1	ICBM 452 T1 Frame of Reference	ICBM452 T1 Atlas
1.2.840.10008.1.4.​2.​2	ICBM Single Subject MRI Frame of Reference	ICBM Single Subject MRI Anatomical Template

SPM2 (Statistical Parametric Mapping) templates are available at http://​github.com/​spm/​spm2/, and they are described at https://​github.com/​spm/​spm2/​blob/​master/​spm_templates.man.
ICBM templates are available at http://​www.loni.ucla.edu/​ICBM/​ICBM_ICBMAtlases.html.
A-3 lists the Context Groups and their UID values. For retired Context Groups, the edition of the Standard in parentheses is the edition in which the item last appeared before it was retired
Table A-3. Context Group UID Values
Context UID	Context Identifier	Context Group Name	Comment
1.2.840.10008.6.​1.​1	CID 2	Anatomic Modifier
1.2.840.10008.6.​1.​2	CID 4	Anatomic Region
1.2.840.10008.6.​1.​3	CID 5	Transducer Approach
1.2.840.10008.6.​1.​4	CID 6	Transducer Orientation
1.2.840.10008.6.​1.​5	CID 7	Ultrasound Beam Path
1.2.840.10008.6.​1.​6	CID 8	Angiographic Interventional Devices
1.2.840.10008.6.​1.​7	CID 9	Image Guided Therapeutic Procedures
1.2.840.10008.6.​1.​8	CID 10	Interventional Drug
1.2.840.10008.6.​1.​9	CID 11	Route of Administration
1.2.840.10008.6.​1.​10	CID 12	Radiographic Contrast Agent
1.2.840.10008.6.​1.​11	CID 13	Radiographic Contrast Agent Ingredient
1.2.840.10008.6.​1.​12	CID 18	Isotopes in Radiopharmaceuticals
1.2.840.10008.6.​1.​13	CID 19	Patient Orientation
1.2.840.10008.6.​1.​14	CID 20	Patient Orientation Modifier
1.2.840.10008.6.​1.​15	CID 21	Patient Equipment Relationship
1.2.840.10008.6.​1.​16	CID 23	Cranio-Caudad Angulation
1.2.840.10008.6.​1.​17	CID 25	Radiopharmaceuticals
1.2.840.10008.6.​1.​18	CID 26	Nuclear Medicine Projections
1.2.840.10008.6.​1.​19	CID 29	Acquisition Modality
1.2.840.10008.6.​1.​20	CID 30	DICOM Devices
1.2.840.10008.6.​1.​21	CID 31	Abstract Priors
1.2.840.10008.6.​1.​22	CID 42	Numeric Value Qualifier
1.2.840.10008.6.​1.​23	CID 82	Units of Measurement
1.2.840.10008.6.​1.​24	CID 83	Units for Real World Value Mapping
1.2.840.10008.6.​1.​25	CID 220	Level of Significance
1.2.840.10008.6.​1.​26	CID 221	Measurement Range Concepts
1.2.840.10008.6.​1.​27	CID 222	Normality Codes
1.2.840.10008.6.​1.​28	CID 223	Normal Range Values
1.2.840.10008.6.​1.​29	CID 224	Selection Method
1.2.840.10008.6.​1.​30	CID 225	Measurement Uncertainty Concepts
1.2.840.10008.6.​1.​31	CID 226	Population Statistical Descriptors
1.2.840.10008.6.​1.​32	CID 227	Sample Statistical Descriptors
1.2.840.10008.6.​1.​33	CID 228	Equation or Table
1.2.840.10008.6.​1.​34	CID 230	Yes-No
1.2.840.10008.6.​1.​35	CID 240	Present-Absent
1.2.840.10008.6.​1.​36	CID 242	Normal-Abnormal
1.2.840.10008.6.​1.​37	CID 244	Laterality
1.2.840.10008.6.​1.​38	CID 250	Positive-Negative
1.2.840.10008.6.​1.​39	CID 251	Severity of Complication
1.2.840.10008.6.​1.​40	CID 270	Observer Type
1.2.840.10008.6.​1.​41	CID 271	Observation Subject Class
1.2.840.10008.6.​1.​42	CID 3000	Audio Channel Source
1.2.840.10008.6.​1.​43	CID 3001	ECG Leads
1.2.840.10008.6.​1.​44	CID 3003	Hemodynamic Waveform Sources
1.2.840.10008.6.​1.​45	CID 3010	Cardiovascular Anatomic Locations
1.2.840.10008.6.​1.​46	CID 3011	Electrophysiology Anatomic Locations
1.2.840.10008.6.​1.​47	CID 3014	Coronary Artery Segments
1.2.840.10008.6.​1.​48	CID 3015	Coronary Arteries
1.2.840.10008.6.​1.​49	CID 3019	Cardiovascular Anatomic Location Modifiers
1.2.840.10008.6.​1.​50	CID 3082	Cardiology Units of Measurement (Retired)	RET (2011)
1.2.840.10008.6.​1.​51	CID 3090	Time Synchronization Channel Types
1.2.840.10008.6.​1.​52	CID 3101	Cardiac Procedural State Values
1.2.840.10008.6.​1.​53	CID 3240	Electrophysiology Measurement Functions and Techniques
1.2.840.10008.6.​1.​54	CID 3241	Hemodynamic Measurement Techniques
1.2.840.10008.6.​1.​55	CID 3250	Catheterization Procedure Phase
1.2.840.10008.6.​1.​56	CID 3254	Electrophysiology Procedure Phase
1.2.840.10008.6.​1.​57	CID 3261	Stress Protocols
1.2.840.10008.6.​1.​58	CID 3262	ECG Patient State Values
1.2.840.10008.6.​1.​59	CID 3263	Electrode Placement Values
1.2.840.10008.6.​1.​60	CID 3264	XYZ Electrode Placement Values (Retired)	RET (2013)
1.2.840.10008.6.​1.​61	CID 3271	Hemodynamic Physiological Challenges
1.2.840.10008.6.​1.​62	CID 3335	ECG Annotations
1.2.840.10008.6.​1.​63	CID 3337	Hemodynamic Annotations
1.2.840.10008.6.​1.​64	CID 3339	Electrophysiology Annotations
1.2.840.10008.6.​1.​65	CID 3400	Procedure Log Titles
1.2.840.10008.6.​1.​66	CID 3401	Types of Log Notes
1.2.840.10008.6.​1.​67	CID 3402	Patient Status and Events
1.2.840.10008.6.​1.​68	CID 3403	Percutaneous Entry
1.2.840.10008.6.​1.​69	CID 3404	Staff Actions
1.2.840.10008.6.​1.​70	CID 3405	Procedure Action Values
1.2.840.10008.6.​1.​71	CID 3406	Non-coronary Transcatheter Interventions
1.2.840.10008.6.​1.​72	CID 3407	Purpose of Reference to Object
1.2.840.10008.6.​1.​73	CID 3408	Actions With Consumables
1.2.840.10008.6.​1.​74	CID 3409	Administration of Drugs/Contrast
1.2.840.10008.6.​1.​75	CID 3410	Numeric Parameters of Drugs/Contrast
1.2.840.10008.6.​1.​76	CID 3411	Intracoronary Devices
1.2.840.10008.6.​1.​77	CID 3412	Intervention Actions and Status
1.2.840.10008.6.​1.​78	CID 3413	Adverse Outcomes
1.2.840.10008.6.​1.​79	CID 3414	Procedure Urgency
1.2.840.10008.6.​1.​80	CID 3415	Cardiac Rhythms
1.2.840.10008.6.​1.​81	CID 3416	Respiration Rhythms
1.2.840.10008.6.​1.​82	CID 3418	Lesion Risk
1.2.840.10008.6.​1.​83	CID 3419	Findings Titles
1.2.840.10008.6.​1.​84	CID 3421	Procedure Action
1.2.840.10008.6.​1.​85	CID 3422	Device Use Actions
1.2.840.10008.6.​1.​86	CID 3423	Numeric Device Characteristics
1.2.840.10008.6.​1.​87	CID 3425	Intervention Parameters
1.2.840.10008.6.​1.​88	CID 3426	Consumables Parameters
1.2.840.10008.6.​1.​89	CID 3427	Equipment Events
1.2.840.10008.6.​1.​90	CID 3428	Imaging Procedures
1.2.840.10008.6.​1.​91	CID 3429	Catheterization Devices
1.2.840.10008.6.​1.​92	CID 3430	DateTime Qualifiers
1.2.840.10008.6.​1.​93	CID 3440	Peripheral Pulse Locations
1.2.840.10008.6.​1.​94	CID 3441	Patient Assessments
1.2.840.10008.6.​1.​95	CID 3442	Peripheral Pulse Methods
1.2.840.10008.6.​1.​96	CID 3446	Skin Condition
1.2.840.10008.6.​1.​97	CID 3448	Airway Assessment
1.2.840.10008.6.​1.​98	CID 3451	Calibration Objects
1.2.840.10008.6.​1.​99	CID 3452	Calibration Methods
1.2.840.10008.6.​1.​100	CID 3453	Cardiac Volume Methods
1.2.840.10008.6.​1.​101	CID 3455	Index Methods
1.2.840.10008.6.​1.​102	CID 3456	Sub-segment Methods
1.2.840.10008.6.​1.​103	CID 3458	Contour Realignment
1.2.840.10008.6.​1.​104	CID 3460	Circumferential Extent
1.2.840.10008.6.​1.​105	CID 3461	Regional Extent
1.2.840.10008.6.​1.​106	CID 3462	Chamber Identification
1.2.840.10008.6.​1.​107	CID 3465	QA Reference Methods
1.2.840.10008.6.​1.​108	CID 3466	Plane Identification
1.2.840.10008.6.​1.​109	CID 3467	Ejection Fraction
1.2.840.10008.6.​1.​110	CID 3468	ED Volume
1.2.840.10008.6.​1.​111	CID 3469	ES Volume
1.2.840.10008.6.​1.​112	CID 3470	Vessel Lumen Cross-sectional Area Calculation Methods
1.2.840.10008.6.​1.​113	CID 3471	Estimated Volumes
1.2.840.10008.6.​1.​114	CID 3472	Cardiac Contraction Phase
1.2.840.10008.6.​1.​115	CID 3480	IVUS Procedure Phases
1.2.840.10008.6.​1.​116	CID 3481	IVUS Distance Measurements
1.2.840.10008.6.​1.​117	CID 3482	IVUS Area Measurements
1.2.840.10008.6.​1.​118	CID 3483	IVUS Longitudinal Measurements
1.2.840.10008.6.​1.​119	CID 3484	IVUS Indices and Ratios
1.2.840.10008.6.​1.​120	CID 3485	IVUS Volume Measurements
1.2.840.10008.6.​1.​121	CID 3486	Vascular Measurement Sites
1.2.840.10008.6.​1.​122	CID 3487	Intravascular Volumetric Regions
1.2.840.10008.6.​1.​123	CID 3488	Min/Max/Mean
1.2.840.10008.6.​1.​124	CID 3489	Calcium Distribution
1.2.840.10008.6.​1.​125	CID 3491	IVUS Lesion Morphologies
1.2.840.10008.6.​1.​126	CID 3492	Vascular Dissection Classifications
1.2.840.10008.6.​1.​127	CID 3493	IVUS Relative Stenosis Severities
1.2.840.10008.6.​1.​128	CID 3494	IVUS Non Morphological Findings
1.2.840.10008.6.​1.​129	CID 3495	IVUS Plaque Composition
1.2.840.10008.6.​1.​130	CID 3496	IVUS Fiducial Points
1.2.840.10008.6.​1.​131	CID 3497	IVUS Arterial Morphology
1.2.840.10008.6.​1.​132	CID 3500	Pressure Units
1.2.840.10008.6.​1.​133	CID 3502	Hemodynamic Resistance Units
1.2.840.10008.6.​1.​134	CID 3503	Indexed Hemodynamic Resistance Units
1.2.840.10008.6.​1.​135	CID 3510	Catheter Size Units
1.2.840.10008.6.​1.​136	CID 3515	Specimen Collection
1.2.840.10008.6.​1.​137	CID 3520	Blood Source Type
1.2.840.10008.6.​1.​138	CID 3524	Blood Gas Pressures
1.2.840.10008.6.​1.​139	CID 3525	Blood Gas Content
1.2.840.10008.6.​1.​140	CID 3526	Blood Gas Saturation
1.2.840.10008.6.​1.​141	CID 3527	Blood Base Excess
1.2.840.10008.6.​1.​142	CID 3528	Blood pH
1.2.840.10008.6.​1.​143	CID 3529	Arterial / Venous Content
1.2.840.10008.6.​1.​144	CID 3530	Oxygen Administration Actions
1.2.840.10008.6.​1.​145	CID 3531	Oxygen Administration
1.2.840.10008.6.​1.​146	CID 3550	Circulatory Support Actions
1.2.840.10008.6.​1.​147	CID 3551	Ventilation Actions
1.2.840.10008.6.​1.​148	CID 3552	Pacing Actions
1.2.840.10008.6.​1.​149	CID 3553	Circulatory Support
1.2.840.10008.6.​1.​150	CID 3554	Ventilation
1.2.840.10008.6.​1.​151	CID 3555	Pacing
1.2.840.10008.6.​1.​152	CID 3560	Blood Pressure Methods
1.2.840.10008.6.​1.​153	CID 3600	Relative Times
1.2.840.10008.6.​1.​154	CID 3602	Hemodynamic Patient State
1.2.840.10008.6.​1.​155	CID 3604	Arterial Lesion Locations
1.2.840.10008.6.​1.​156	CID 3606	Arterial Source Locations
1.2.840.10008.6.​1.​157	CID 3607	Venous Source Locations
1.2.840.10008.6.​1.​158	CID 3608	Atrial Source Locations
1.2.840.10008.6.​1.​159	CID 3609	Ventricular Source Locations
1.2.840.10008.6.​1.​160	CID 3610	Gradient Source Locations
1.2.840.10008.6.​1.​161	CID 3611	Pressure Measurements
1.2.840.10008.6.​1.​162	CID 3612	Blood Velocity Measurements
1.2.840.10008.6.​1.​163	CID 3613	Hemodynamic Time Measurements
1.2.840.10008.6.​1.​164	CID 3614	Valve Areas, Non-mitral
1.2.840.10008.6.​1.​165	CID 3615	Valve Areas
1.2.840.10008.6.​1.​166	CID 3616	Hemodynamic Period Measurements
1.2.840.10008.6.​1.​167	CID 3617	Valve Flows
1.2.840.10008.6.​1.​168	CID 3618	Hemodynamic Flows
1.2.840.10008.6.​1.​169	CID 3619	Hemodynamic Resistance Measurements
1.2.840.10008.6.​1.​170	CID 3620	Hemodynamic Ratios
1.2.840.10008.6.​1.​171	CID 3621	Fractional Flow Reserve
1.2.840.10008.6.​1.​172	CID 3627	Measurement Type
1.2.840.10008.6.​1.​173	CID 3628	Cardiac Output Methods
1.2.840.10008.6.​1.​174	CID 3629	Procedure Intent
1.2.840.10008.6.​1.​175	CID 3630	Cardiovascular Anatomic Locations
1.2.840.10008.6.​1.​176	CID 3640	Hypertension
1.2.840.10008.6.​1.​177	CID 3641	Hemodynamic Assessments
1.2.840.10008.6.​1.​178	CID 3642	Degree Findings
1.2.840.10008.6.​1.​179	CID 3651	Hemodynamic Measurement Phase
1.2.840.10008.6.​1.​180	CID 3663	Body Surface Area Equations
1.2.840.10008.6.​1.​181	CID 3664	Oxygen Consumption Equations and Tables
1.2.840.10008.6.​1.​182	CID 3666	P50 Equations
1.2.840.10008.6.​1.​183	CID 3667	Framingham Scores
1.2.840.10008.6.​1.​184	CID 3668	Framingham Tables
1.2.840.10008.6.​1.​185	CID 3670	ECG Procedure Types
1.2.840.10008.6.​1.​186	CID 3671	Reason for ECG Exam
1.2.840.10008.6.​1.​187	CID 3672	Pacemakers
1.2.840.10008.6.​1.​188	CID 3673	Diagnosis (Retired)	RET (2013)
1.2.840.10008.6.​1.​189	CID 3675	Other Filters (Retired)	RET (2013)
1.2.840.10008.6.​1.​190	CID 3676	Lead Measurement Technique
1.2.840.10008.6.​1.​191	CID 3677	Summary Codes ECG
1.2.840.10008.6.​1.​192	CID 3678	QT Correction Algorithms
1.2.840.10008.6.​1.​193	CID 3679	ECG Morphology Descriptions (Retired)	RET (2013)
1.2.840.10008.6.​1.​194	CID 3680	ECG Lead Noise Descriptions
1.2.840.10008.6.​1.​195	CID 3681	ECG Lead Noise Modifiers (Retired)
1.2.840.10008.6.​1.​196	CID 3682	Probability (Retired)	RET (2013)
1.2.840.10008.6.​1.​197	CID 3683	Modifiers (Retired)	RET (2013)
1.2.840.10008.6.​1.​198	CID 3684	Trend (Retired)	RET (2013)
1.2.840.10008.6.​1.​199	CID 3685	Conjunctive Terms (Retired)	RET (2013)
1.2.840.10008.6.​1.​200	CID 3686	ECG Interpretive Statements (Retired)	RET (2013)
1.2.840.10008.6.​1.​201	CID 3687	Electrophysiology Waveform Durations
1.2.840.10008.6.​1.​202	CID 3688	Electrophysiology Waveform Voltages
1.2.840.10008.6.​1.​203	CID 3700	Cath Diagnosis
1.2.840.10008.6.​1.​204	CID 3701	Cardiac Valves and Tracts
1.2.840.10008.6.​1.​205	CID 3703	Wall Motion
1.2.840.10008.6.​1.​206	CID 3704	Myocardium Wall Morphology Findings
1.2.840.10008.6.​1.​207	CID 3705	Chamber Size
1.2.840.10008.6.​1.​208	CID 3706	Overall Contractility
1.2.840.10008.6.​1.​209	CID 3707	VSD Description
1.2.840.10008.6.​1.​210	CID 3709	Aortic Root Description
1.2.840.10008.6.​1.​211	CID 3710	Coronary Dominance
1.2.840.10008.6.​1.​212	CID 3711	Valvular Abnormalities
1.2.840.10008.6.​1.​213	CID 3712	Vessel Descriptors
1.2.840.10008.6.​1.​214	CID 3713	TIMI Flow Characteristics
1.2.840.10008.6.​1.​215	CID 3714	Thrombus
1.2.840.10008.6.​1.​216	CID 3715	Lesion Margin
1.2.840.10008.6.​1.​217	CID 3716	Severity
1.2.840.10008.6.​1.​218	CID 3717	Myocardial Wall Segments
1.2.840.10008.6.​1.​219	CID 3718	Myocardial Wall Segments in Projection
1.2.840.10008.6.​1.​220	CID 3719	Canadian Clinical Classification
1.2.840.10008.6.​1.​221	CID 3720	Cardiac History Dates (Retired)	RET (2007)
1.2.840.10008.6.​1.​222	CID 3721	Cardiovascular Surgeries
1.2.840.10008.6.​1.​223	CID 3722	Diabetic Therapy
1.2.840.10008.6.​1.​224	CID 3723	MI Types
1.2.840.10008.6.​1.​225	CID 3724	Smoking History
1.2.840.10008.6.​1.​226	CID 3726	Indications for Coronary Intervention
1.2.840.10008.6.​1.​227	CID 3727	Indications for Catheterization
1.2.840.10008.6.​1.​228	CID 3728	Cath Findings
1.2.840.10008.6.​1.​229	CID 3729	Admission Status
1.2.840.10008.6.​1.​230	CID 3730	Insurance Payor
1.2.840.10008.6.​1.​231	CID 3733	Primary Cause of Death
1.2.840.10008.6.​1.​232	CID 3735	Acute Coronary Syndrome Time Period
1.2.840.10008.6.​1.​233	CID 3736	NYHA Classification
1.2.840.10008.6.​1.​234	CID 3737	Non-invasive Test - Ischemia
1.2.840.10008.6.​1.​235	CID 3738	Pre-Cath Angina Type
1.2.840.10008.6.​1.​236	CID 3739	Cath Procedure Type
1.2.840.10008.6.​1.​237	CID 3740	Thrombolytic Administration
1.2.840.10008.6.​1.​238	CID 3741	Medication Administration, Lab Visit
1.2.840.10008.6.​1.​239	CID 3742	Medication Administration, PCI
1.2.840.10008.6.​1.​240	CID 3743	Clopidogrel/Ticlopidine Administration
1.2.840.10008.6.​1.​241	CID 3744	EF Testing Method
1.2.840.10008.6.​1.​242	CID 3745	Calculation Method
1.2.840.10008.6.​1.​243	CID 3746	Percutaneous Entry Site
1.2.840.10008.6.​1.​244	CID 3747	Percutaneous Closure
1.2.840.10008.6.​1.​245	CID 3748	Angiographic EF Testing Method
1.2.840.10008.6.​1.​246	CID 3749	PCI Procedure Result
1.2.840.10008.6.​1.​247	CID 3750	Previously Dilated Lesion
1.2.840.10008.6.​1.​248	CID 3752	Guidewire Crossing
1.2.840.10008.6.​1.​249	CID 3754	Vascular Complications
1.2.840.10008.6.​1.​250	CID 3755	Cath Complications
1.2.840.10008.6.​1.​251	CID 3756	Cardiac Patient Risk Factors
1.2.840.10008.6.​1.​252	CID 3757	Cardiac Diagnostic Procedures
1.2.840.10008.6.​1.​253	CID 3758	Cardiovascular Family History
1.2.840.10008.6.​1.​254	CID 3760	Hypertension Therapy
1.2.840.10008.6.​1.​255	CID 3761	Antilipemic Agents
1.2.840.10008.6.​1.​256	CID 3762	Antiarrhythmic Agents
1.2.840.10008.6.​1.​257	CID 3764	Myocardial Infarction Therapies
1.2.840.10008.6.​1.​258	CID 3769	Concern Types
1.2.840.10008.6.​1.​259	CID 3770	Problem Status
1.2.840.10008.6.​1.​260	CID 3772	Health Status
1.2.840.10008.6.​1.​261	CID 3773	Use Status
1.2.840.10008.6.​1.​262	CID 3774	Social History
1.2.840.10008.6.​1.​263	CID 3777	Implanted Devices
1.2.840.10008.6.​1.​264	CID 3802	Plaque Structures
1.2.840.10008.6.​1.​265	CID 3804	Stenosis Measurement Methods
1.2.840.10008.6.​1.​266	CID 3805	Stenosis Types
1.2.840.10008.6.​1.​267	CID 3806	Stenosis Shape
1.2.840.10008.6.​1.​268	CID 3807	Volume Measurement Methods
1.2.840.10008.6.​1.​269	CID 3808	Aneurysm Types
1.2.840.10008.6.​1.​270	CID 3809	Associated Conditions
1.2.840.10008.6.​1.​271	CID 3810	Vascular Morphology
1.2.840.10008.6.​1.​272	CID 3813	Stent Findings
1.2.840.10008.6.​1.​273	CID 3814	Stent Composition
1.2.840.10008.6.​1.​274	CID 3815	Source of Vascular Finding
1.2.840.10008.6.​1.​275	CID 3817	Vascular Sclerosis Types
1.2.840.10008.6.​1.​276	CID 3820	Non-invasive Vascular Procedures
1.2.840.10008.6.​1.​277	CID 3821	Papillary Muscle Included/Excluded
1.2.840.10008.6.​1.​278	CID 3823	Respiratory Status
1.2.840.10008.6.​1.​279	CID 3826	Heart Rhythm
1.2.840.10008.6.​1.​280	CID 3827	Vessel Segments
1.2.840.10008.6.​1.​281	CID 3829	Pulmonary Arteries
1.2.840.10008.6.​1.​282	CID 3831	Stenosis Length
1.2.840.10008.6.​1.​283	CID 3832	Stenosis Grade
1.2.840.10008.6.​1.​284	CID 3833	Cardiac Ejection Fraction
1.2.840.10008.6.​1.​285	CID 3835	Cardiac Volume Measurements
1.2.840.10008.6.​1.​286	CID 3836	Time-based Perfusion Measurements
1.2.840.10008.6.​1.​287	CID 3837	Fiducial Feature
1.2.840.10008.6.​1.​288	CID 3838	Diameter Derivation
1.2.840.10008.6.​1.​289	CID 3839	Coronary Veins
1.2.840.10008.6.​1.​290	CID 3840	Pulmonary Veins
1.2.840.10008.6.​1.​291	CID 3843	Myocardial Subsegment
1.2.840.10008.6.​1.​292	CID 4005	Partial View Section for Mammography
1.2.840.10008.6.​1.​293	CID 4009	DX Anatomy Imaged
1.2.840.10008.6.​1.​294	CID 4010	DX View
1.2.840.10008.6.​1.​295	CID 4011	DX View Modifier
1.2.840.10008.6.​1.​296	CID 4012	Projection Eponymous Name
1.2.840.10008.6.​1.​297	CID 4013	Anatomic Region for Mammography
1.2.840.10008.6.​1.​298	CID 4014	View for Mammography
1.2.840.10008.6.​1.​299	CID 4015	View Modifier for Mammography
1.2.840.10008.6.​1.​300	CID 4016	Anatomic Region for Intra-oral Radiography
1.2.840.10008.6.​1.​301	CID 4017	Anatomic Region Modifier for Intra-oral Radiography
1.2.840.10008.6.​1.​302	CID 4018	Primary Anatomic Structure for Intra-oral Radiography (Permanent Dentition - Designation of Teeth)
1.2.840.10008.6.​1.​303	CID 4019	Primary Anatomic Structure for Intra-oral Radiography (Deciduous Dentition - Designation of Teeth)
1.2.840.10008.6.​1.​304	CID 4020	PET Radionuclide
1.2.840.10008.6.​1.​305	CID 4021	PET Radiopharmaceutical
1.2.840.10008.6.​1.​306	CID 4028	Craniofacial Anatomic Regions
1.2.840.10008.6.​1.​307	CID 4030	CT, MR and PET Anatomy Imaged
1.2.840.10008.6.​1.​308	CID 4031	Common Anatomic Regions
1.2.840.10008.6.​1.​309	CID 4032	MR Spectroscopy Metabolites
1.2.840.10008.6.​1.​310	CID 4033	MR Proton Spectroscopy Metabolites
1.2.840.10008.6.​1.​311	CID 4040	Endoscopy Anatomic Regions
1.2.840.10008.6.​1.​312	CID 4042	XA/XRF Anatomy Imaged
1.2.840.10008.6.​1.​313	CID 4050	Drug or Contrast Agent Characteristics
1.2.840.10008.6.​1.​314	CID 4051	General Devices
1.2.840.10008.6.​1.​315	CID 4052	Phantom Devices
1.2.840.10008.6.​1.​316	CID 4200	Ophthalmic Imaging Agent
1.2.840.10008.6.​1.​317	CID 4201	Patient Eye Movement Command
1.2.840.10008.6.​1.​318	CID 4202	Ophthalmic Photography Acquisition Device
1.2.840.10008.6.​1.​319	CID 4203	Ophthalmic Photography Illumination
1.2.840.10008.6.​1.​320	CID 4204	Ophthalmic Filter
1.2.840.10008.6.​1.​321	CID 4205	Ophthalmic Lens
1.2.840.10008.6.​1.​322	CID 4206	Ophthalmic Channel Description
1.2.840.10008.6.​1.​323	CID 4207	Ophthalmic Image Position
1.2.840.10008.6.​1.​324	CID 4208	Mydriatic Agent
1.2.840.10008.6.​1.​325	CID 4209	Ophthalmic Anatomic Structure Imaged
1.2.840.10008.6.​1.​326	CID 4210	Ophthalmic Tomography Acquisition Device
1.2.840.10008.6.​1.​327	CID 4211	Ophthalmic OCT Anatomic Structure Imaged
1.2.840.10008.6.​1.​328	CID 5000	Languages
1.2.840.10008.6.​1.​329	CID 5001	Countries
1.2.840.10008.6.​1.​330	CID 6000	Overall Breast Composition
1.2.840.10008.6.​1.​331	CID 6001	Overall Breast Composition from BI-RADS®
1.2.840.10008.6.​1.​332	CID 6002	Change Since Last Mammogram or Prior Surgery
1.2.840.10008.6.​1.​333	CID 6003	Change Since Last Mammogram or Prior Surgery from BI-RADS®
1.2.840.10008.6.​1.​334	CID 6004	Mammography Characteristics of Shape
1.2.840.10008.6.​1.​335	CID 6005	Characteristics of Shape from BI-RADS®
1.2.840.10008.6.​1.​336	CID 6006	Mammography Characteristics of Margin
1.2.840.10008.6.​1.​337	CID 6007	Characteristics of Margin from BI-RADS®
1.2.840.10008.6.​1.​338	CID 6008	Density Modifier
1.2.840.10008.6.​1.​339	CID 6009	Density Modifier from BI-RADS®
1.2.840.10008.6.​1.​340	CID 6010	Mammography Calcification Types
1.2.840.10008.6.​1.​341	CID 6011	Calcification Types from BI-RADS®
1.2.840.10008.6.​1.​342	CID 6012	Calcification Distribution Modifier
1.2.840.10008.6.​1.​343	CID 6013	Calcification Distribution Modifier from BI-RADS®
1.2.840.10008.6.​1.​344	CID 6014	Mammography Single Image Finding
1.2.840.10008.6.​1.​345	CID 6015	Single Image Finding from BI-RADS®
1.2.840.10008.6.​1.​346	CID 6016	Mammography Composite Feature
1.2.840.10008.6.​1.​347	CID 6017	Composite Feature from BI-RADS®
1.2.840.10008.6.​1.​348	CID 6018	Clockface Location or Region
1.2.840.10008.6.​1.​349	CID 6019	Clockface Location or Region from BI-RADS®
1.2.840.10008.6.​1.​350	CID 6020	Quadrant Location
1.2.840.10008.6.​1.​351	CID 6021	Quadrant Location from BI-RADS®
1.2.840.10008.6.​1.​352	CID 6022	Side
1.2.840.10008.6.​1.​353	CID 6023	Side from BI-RADS®
1.2.840.10008.6.​1.​354	CID 6024	Depth
1.2.840.10008.6.​1.​355	CID 6025	Depth from BI-RADS®
1.2.840.10008.6.​1.​356	CID 6026	Mammography Assessment
1.2.840.10008.6.​1.​357	CID 6027	Assessment from BI-RADS®
1.2.840.10008.6.​1.​358	CID 6028	Mammography Recommended Follow-up
1.2.840.10008.6.​1.​359	CID 6029	Recommended Follow-up from BI-RADS®
1.2.840.10008.6.​1.​360	CID 6030	Mammography Pathology Codes
1.2.840.10008.6.​1.​361	CID 6031	Benign Pathology Codes from BI-RADS®
1.2.840.10008.6.​1.​362	CID 6032	High Risk Lesions Pathology Codes from BI-RADS®
1.2.840.10008.6.​1.​363	CID 6033	Malignant Pathology Codes from BI-RADS®
1.2.840.10008.6.​1.​364	CID 6034	Intended Use of CAD Output
1.2.840.10008.6.​1.​365	CID 6035	Composite Feature Relations
1.2.840.10008.6.​1.​366	CID 6036	Scope of Feature
1.2.840.10008.6.​1.​367	CID 6037	Mammography Quantitative Temporal Difference Type
1.2.840.10008.6.​1.​368	CID 6038	Mammography Qualitative Temporal Difference Type
1.2.840.10008.6.​1.​369	CID 6039	Nipple Characteristic
1.2.840.10008.6.​1.​370	CID 6040	Non-lesion Object Type
1.2.840.10008.6.​1.​371	CID 6041	Mammography Image Quality Finding
1.2.840.10008.6.​1.​372	CID 6042	Status of Results
1.2.840.10008.6.​1.​373	CID 6043	Types of Mammography CAD Analysis
1.2.840.10008.6.​1.​374	CID 6044	Types of Image Quality Assessment
1.2.840.10008.6.​1.​375	CID 6045	Mammography Types of Quality Control Standard
1.2.840.10008.6.​1.​376	CID 6046	Units of Follow-up Interval
1.2.840.10008.6.​1.​377	CID 6047	CAD Processing and Findings Summary
1.2.840.10008.6.​1.​378	CID 6048	CAD Operating Point Axis Label
1.2.840.10008.6.​1.​379	CID 6050	Breast Procedure Reported
1.2.840.10008.6.​1.​380	CID 6051	Breast Procedure Reason
1.2.840.10008.6.​1.​381	CID 6052	Breast Imaging Report Section Title
1.2.840.10008.6.​1.​382	CID 6053	Breast Imaging Report Elements
1.2.840.10008.6.​1.​383	CID 6054	Breast Imaging Findings
1.2.840.10008.6.​1.​384	CID 6055	Breast Clinical Finding or Indicated Problem
1.2.840.10008.6.​1.​385	CID 6056	Associated Findings for Breast
1.2.840.10008.6.​1.​386	CID 6057	Ductography Findings for Breast
1.2.840.10008.6.​1.​387	CID 6058	Procedure Modifiers for Breast
1.2.840.10008.6.​1.​388	CID 6059	Breast Implant Types
1.2.840.10008.6.​1.​389	CID 6060	Breast Biopsy Techniques
1.2.840.10008.6.​1.​390	CID 6061	Breast Imaging Procedure Modifiers
1.2.840.10008.6.​1.​391	CID 6062	Interventional Procedure Complications
1.2.840.10008.6.​1.​392	CID 6063	Interventional Procedure Results
1.2.840.10008.6.​1.​393	CID 6064	Ultrasound Findings for Breast
1.2.840.10008.6.​1.​394	CID 6065	Instrument Approach
1.2.840.10008.6.​1.​395	CID 6066	Target Confirmation
1.2.840.10008.6.​1.​396	CID 6067	Fluid Color
1.2.840.10008.6.​1.​397	CID 6068	Tumor Stages From AJCC
1.2.840.10008.6.​1.​398	CID 6069	Nottingham Combined Histologic Grade
1.2.840.10008.6.​1.​399	CID 6070	Bloom-Richardson Histologic Grade
1.2.840.10008.6.​1.​400	CID 6071	Histologic Grading Method
1.2.840.10008.6.​1.​401	CID 6072	Breast Implant Findings
1.2.840.10008.6.​1.​402	CID 6080	Gynecological Hormones
1.2.840.10008.6.​1.​403	CID 6081	Breast Cancer Risk Factors
1.2.840.10008.6.​1.​404	CID 6082	Gynecological Procedures
1.2.840.10008.6.​1.​405	CID 6083	Procedures for Breast
1.2.840.10008.6.​1.​406	CID 6084	Mammoplasty Procedures
1.2.840.10008.6.​1.​407	CID 6085	Therapies for Breast
1.2.840.10008.6.​1.​408	CID 6086	Menopausal Phase
1.2.840.10008.6.​1.​409	CID 6087	General Risk Factors
1.2.840.10008.6.​1.​410	CID 6088	OB-GYN Maternal Risk Factors
1.2.840.10008.6.​1.​411	CID 6089	Substances
1.2.840.10008.6.​1.​412	CID 6090	Relative Usage, Exposure Amount
1.2.840.10008.6.​1.​413	CID 6091	Relative Frequency of Event Values
1.2.840.10008.6.​1.​414	CID 6092	Quantitative Concepts for Usage, Exposure
1.2.840.10008.6.​1.​415	CID 6093	Qualitative Concepts for Usage, Exposure Amount
1.2.840.10008.6.​1.​416	CID 6094	Qualitative Concepts for Usage, Exposure Frequency
1.2.840.10008.6.​1.​417	CID 6095	Numeric Properties of Procedures
1.2.840.10008.6.​1.​418	CID 6096	Pregnancy Status
1.2.840.10008.6.​1.​419	CID 6097	Side of Family
1.2.840.10008.6.​1.​420	CID 6100	Chest Component Categories
1.2.840.10008.6.​1.​421	CID 6101	Chest Finding or Feature
1.2.840.10008.6.​1.​422	CID 6102	Chest Finding or Feature Modifier
1.2.840.10008.6.​1.​423	CID 6103	Abnormal Lines Finding or Feature
1.2.840.10008.6.​1.​424	CID 6104	Abnormal Opacity Finding or Feature
1.2.840.10008.6.​1.​425	CID 6105	Abnormal Lucency Finding or Feature
1.2.840.10008.6.​1.​426	CID 6106	Abnormal Texture Finding or Feature
1.2.840.10008.6.​1.​427	CID 6107	Width Descriptor
1.2.840.10008.6.​1.​428	CID 6108	Chest Anatomic Structure Abnormal Distribution
1.2.840.10008.6.​1.​429	CID 6109	Radiographic Anatomy Finding or Feature
1.2.840.10008.6.​1.​430	CID 6110	Lung Anatomy Finding or Feature
1.2.840.10008.6.​1.​431	CID 6111	Bronchovascular Anatomy Finding or Feature
1.2.840.10008.6.​1.​432	CID 6112	Pleura Anatomy Finding or Feature
1.2.840.10008.6.​1.​433	CID 6113	Mediastinum Anatomy Finding or Feature
1.2.840.10008.6.​1.​434	CID 6114	Osseous Anatomy Finding or Feature
1.2.840.10008.6.​1.​435	CID 6115	Osseous Anatomy Modifiers
1.2.840.10008.6.​1.​436	CID 6116	Muscular Anatomy
1.2.840.10008.6.​1.​437	CID 6117	Vascular Anatomy
1.2.840.10008.6.​1.​438	CID 6118	Size Descriptor
1.2.840.10008.6.​1.​439	CID 6119	Chest Border Shape
1.2.840.10008.6.​1.​440	CID 6120	Chest Border Definition
1.2.840.10008.6.​1.​441	CID 6121	Chest Orientation Descriptor
1.2.840.10008.6.​1.​442	CID 6122	Chest Content Descriptor
1.2.840.10008.6.​1.​443	CID 6123	Chest Opacity Descriptor
1.2.840.10008.6.​1.​444	CID 6124	Location in Chest
1.2.840.10008.6.​1.​445	CID 6125	General Chest Location
1.2.840.10008.6.​1.​446	CID 6126	Location in Lung
1.2.840.10008.6.​1.​447	CID 6127	Segment Location in Lung
1.2.840.10008.6.​1.​448	CID 6128	Chest Distribution Descriptor
1.2.840.10008.6.​1.​449	CID 6129	Chest Site Involvement
1.2.840.10008.6.​1.​450	CID 6130	Severity Descriptor
1.2.840.10008.6.​1.​451	CID 6131	Chest Texture Descriptor
1.2.840.10008.6.​1.​452	CID 6132	Chest Calcification Descriptor
1.2.840.10008.6.​1.​453	CID 6133	Chest Quantitative Temporal Difference Type
1.2.840.10008.6.​1.​454	CID 6134	Chest Qualitative Temporal Difference Type
1.2.840.10008.6.​1.​455	CID 6135	Image Quality Finding
1.2.840.10008.6.​1.​456	CID 6136	Chest Types of Quality Control Standard
1.2.840.10008.6.​1.​457	CID 6137	Types of CAD Analysis
1.2.840.10008.6.​1.​458	CID 6138	Chest Non-lesion Object Type
1.2.840.10008.6.​1.​459	CID 6139	Non-lesion Modifiers
1.2.840.10008.6.​1.​460	CID 6140	Calculation Methods
1.2.840.10008.6.​1.​461	CID 6141	Attenuation Coefficient Measurements
1.2.840.10008.6.​1.​462	CID 6142	Calculated Value
1.2.840.10008.6.​1.​463	CID 6143	Lesion Response
1.2.840.10008.6.​1.​464	CID 6144	RECIST Defined Lesion Response
1.2.840.10008.6.​1.​465	CID 6145	Baseline Category
1.2.840.10008.6.​1.​466	CID 6151	Background Echotexture
1.2.840.10008.6.​1.​467	CID 6152	Orientation
1.2.840.10008.6.​1.​468	CID 6153	Lesion Boundary
1.2.840.10008.6.​1.​469	CID 6154	Echo Pattern
1.2.840.10008.6.​1.​470	CID 6155	Posterior Acoustic Features
1.2.840.10008.6.​1.​471	CID 6157	Vascularity
1.2.840.10008.6.​1.​472	CID 6158	Correlation to Other Findings
1.2.840.10008.6.​1.​473	CID 6159	Malignancy Type
1.2.840.10008.6.​1.​474	CID 6160	Breast Primary Tumor Assessment From AJCC
1.2.840.10008.6.​1.​475	CID 6161	Clinical Regional Lymph Node Assessment for Breast
1.2.840.10008.6.​1.​476	CID 6162	Assessment of Metastasis for Breast
1.2.840.10008.6.​1.​477	CID 6163	Menstrual Cycle Phase
1.2.840.10008.6.​1.​478	CID 6164	Time Intervals
1.2.840.10008.6.​1.​479	CID 6165	Breast Linear Measurements
1.2.840.10008.6.​1.​480	CID 6166	CAD Geometry Secondary Graphical Representation
1.2.840.10008.6.​1.​481	CID 7000	Diagnostic Imaging Report Document Titles
1.2.840.10008.6.​1.​482	CID 7001	Diagnostic Imaging Report Headings
1.2.840.10008.6.​1.​483	CID 7002	Diagnostic Imaging Report Elements
1.2.840.10008.6.​1.​484	CID 7003	Diagnostic Imaging Report Purposes of Reference
1.2.840.10008.6.​1.​485	CID 7004	Waveform Purposes of Reference
1.2.840.10008.6.​1.​486	CID 7005	Contributing Equipment Purposes of Reference
1.2.840.10008.6.​1.​487	CID 7006	SR Document Purposes of Reference
1.2.840.10008.6.​1.​488	CID 7007	Signature Purpose
1.2.840.10008.6.​1.​489	CID 7008	Media Import
1.2.840.10008.6.​1.​490	CID 7010	Key Object Selection Document Title
1.2.840.10008.6.​1.​491	CID 7011	Rejected for Quality Reasons
1.2.840.10008.6.​1.​492	CID 7012	Best in Set
1.2.840.10008.6.​1.​493	CID 7020	Document Titles
1.2.840.10008.6.​1.​494	CID 7100	RCS Registration Method Type
1.2.840.10008.6.​1.​495	CID 7101	Brain Atlas Fiducials
1.2.840.10008.6.​1.​496	CID 7150	Segmentation Property Categories
1.2.840.10008.6.​1.​497	CID 7151	Segmentation Property Types
1.2.840.10008.6.​1.​498	CID 7152	Cardiac Structure Segmentation Types
1.2.840.10008.6.​1.​499	CID 7153	CNS Segmentation Types
1.2.840.10008.6.​1.​500	CID 7154	Abdominal Segmentation Types
1.2.840.10008.6.​1.​501	CID 7155	Thoracic Segmentation Types
1.2.840.10008.6.​1.​502	CID 7156	Vascular Segmentation Types
1.2.840.10008.6.​1.​503	CID 7157	Device Segmentation Types
1.2.840.10008.6.​1.​504	CID 7158	Artifact Segmentation Types
1.2.840.10008.6.​1.​505	CID 7159	Lesion Segmentation Types
1.2.840.10008.6.​1.​506	CID 7160	Pelvic Organ Segmentation Types
1.2.840.10008.6.​1.​507	CID 7161	Physiology Segmentation Types
1.2.840.10008.6.​1.​508	CID 7201	Referenced Image Purposes of Reference
1.2.840.10008.6.​1.​509	CID 7202	Source Image Purposes of Reference
1.2.840.10008.6.​1.​510	CID 7203	Image Derivation
1.2.840.10008.6.​1.​511	CID 7205	Purpose of Reference to Alternate Representation
1.2.840.10008.6.​1.​512	CID 7210	Related Series Purposes of Reference
1.2.840.10008.6.​1.​513	CID 7250	Multi-Frame Subset Type
1.2.840.10008.6.​1.​514	CID 7450	Person Roles
1.2.840.10008.6.​1.​515	CID 7451	Family Member
1.2.840.10008.6.​1.​516	CID 7452	Organizational Roles
1.2.840.10008.6.​1.​517	CID 7453	Performing Roles
1.2.840.10008.6.​1.​518	CID 7454	Animal Taxonomic Rank Values
1.2.840.10008.6.​1.​519	CID 7455	Sex
1.2.840.10008.6.​1.​520	CID 7456	Units of Measure for Age
1.2.840.10008.6.​1.​521	CID 7460	Units of Linear Measurement
1.2.840.10008.6.​1.​522	CID 7461	Units of Area Measurement
1.2.840.10008.6.​1.​523	CID 7462	Units of Volume Measurement
1.2.840.10008.6.​1.​524	CID 7470	Linear Measurements
1.2.840.10008.6.​1.​525	CID 7471	Area Measurements
1.2.840.10008.6.​1.​526	CID 7472	Volume Measurements
1.2.840.10008.6.​1.​527	CID 7473	General Area Calculation Methods
1.2.840.10008.6.​1.​528	CID 7474	General Volume Calculation Methods
1.2.840.10008.6.​1.​529	CID 7480	Breed
1.2.840.10008.6.​1.​530	CID 7481	Breed Registry
1.2.840.10008.6.​1.​531	CID 9231	Workitem Definition
1.2.840.10008.6.​1.​532	CID 9232	Non-DICOM Output Types (Retired)	RET (2011)
1.2.840.10008.6.​1.​533	CID 9300	Procedure Discontinuation Reasons
1.2.840.10008.6.​1.​534	CID 10000	Scope of Accumulation
1.2.840.10008.6.​1.​535	CID 10001	UID Types
1.2.840.10008.6.​1.​536	CID 10002	Irradiation Event Types
1.2.840.10008.6.​1.​537	CID 10003	Equipment Plane Identification
1.2.840.10008.6.​1.​538	CID 10004	Fluoro Modes
1.2.840.10008.6.​1.​539	CID 10006	X-Ray Filter Materials
1.2.840.10008.6.​1.​540	CID 10007	X-Ray Filter Types
1.2.840.10008.6.​1.​541	CID 10008	Dose Related Distance Measurements
1.2.840.10008.6.​1.​542	CID 10009	Measured/Calculated
1.2.840.10008.6.​1.​543	CID 10010	Dose Measurement Devices
1.2.840.10008.6.​1.​544	CID 10011	Effective Dose Evaluation Method
1.2.840.10008.6.​1.​545	CID 10013	CT Acquisition Type
1.2.840.10008.6.​1.​546	CID 10014	Contrast Imaging Technique
1.2.840.10008.6.​1.​547	CID 10015	CT Dose Reference Authorities
1.2.840.10008.6.​1.​548	CID 10016	Anode Target Material
1.2.840.10008.6.​1.​549	CID 10017	X-Ray Grid
1.2.840.10008.6.​1.​550	CID 12001	Ultrasound Protocol Types
1.2.840.10008.6.​1.​551	CID 12002	Ultrasound Protocol Stage Types
1.2.840.10008.6.​1.​552	CID 12003	OB-GYN Dates
1.2.840.10008.6.​1.​553	CID 12004	Fetal Biometry Ratios
1.2.840.10008.6.​1.​554	CID 12005	Fetal Biometry Measurements
1.2.840.10008.6.​1.​555	CID 12006	Fetal Long Bones Biometry Measurements
1.2.840.10008.6.​1.​556	CID 12007	Fetal Cranium
1.2.840.10008.6.​1.​557	CID 12008	OB-GYN Amniotic Sac
1.2.840.10008.6.​1.​558	CID 12009	Early Gestation Biometry Measurements
1.2.840.10008.6.​1.​559	CID 12011	Ultrasound Pelvis and Uterus
1.2.840.10008.6.​1.​560	CID 12012	OB Equations and Tables
1.2.840.10008.6.​1.​561	CID 12013	Gestational Age Equations and Tables
1.2.840.10008.6.​1.​562	CID 12014	OB Fetal Body Weight Equations and Tables
1.2.840.10008.6.​1.​563	CID 12015	Fetal Growth Equations and Tables
1.2.840.10008.6.​1.​564	CID 12016	Estimated Fetal Weight Percentile Equations and Tables
1.2.840.10008.6.​1.​565	CID 12017	Growth Distribution Rank
1.2.840.10008.6.​1.​566	CID 12018	OB-GYN Summary
1.2.840.10008.6.​1.​567	CID 12019	OB-GYN Fetus Summary
1.2.840.10008.6.​1.​568	CID 12101	Vascular Summary
1.2.840.10008.6.​1.​569	CID 12102	Temporal Periods Relating to Procedure or Therapy
1.2.840.10008.6.​1.​570	CID 12103	Vascular Ultrasound Anatomic Location
1.2.840.10008.6.​1.​571	CID 12104	Extracranial Arteries
1.2.840.10008.6.​1.​572	CID 12105	Intracranial Cerebral Vessels
1.2.840.10008.6.​1.​573	CID 12106	Intracranial Cerebral Vessels (Unilateral)
1.2.840.10008.6.​1.​574	CID 12107	Upper Extremity Arteries
1.2.840.10008.6.​1.​575	CID 12108	Upper Extremity Veins
1.2.840.10008.6.​1.​576	CID 12109	Lower Extremity Arteries
1.2.840.10008.6.​1.​577	CID 12110	Lower Extremity Veins
1.2.840.10008.6.​1.​578	CID 12111	Abdominal Arteries (Lateral)
1.2.840.10008.6.​1.​579	CID 12112	Abdominal Arteries (Unilateral)
1.2.840.10008.6.​1.​580	CID 12113	Abdominal Veins (Lateral)
1.2.840.10008.6.​1.​581	CID 12114	Abdominal Veins (Unilateral)
1.2.840.10008.6.​1.​582	CID 12115	Renal Vessels
1.2.840.10008.6.​1.​583	CID 12116	Vessel Segment Modifiers
1.2.840.10008.6.​1.​584	CID 12117	Vessel Branch Modifiers
1.2.840.10008.6.​1.​585	CID 12119	Vascular Ultrasound Property
1.2.840.10008.6.​1.​586	CID 12120	Blood Velocity Measurements by Ultrasound
1.2.840.10008.6.​1.​587	CID 12121	Vascular Indices and Ratios
1.2.840.10008.6.​1.​588	CID 12122	Other Vascular Properties
1.2.840.10008.6.​1.​589	CID 12123	Carotid Ratios
1.2.840.10008.6.​1.​590	CID 12124	Renal Ratios
1.2.840.10008.6.​1.​591	CID 12140	Pelvic Vasculature Anatomical Location
1.2.840.10008.6.​1.​592	CID 12141	Fetal Vasculature Anatomical Location
1.2.840.10008.6.​1.​593	CID 12200	Echocardiography Left Ventricle
1.2.840.10008.6.​1.​594	CID 12201	Left Ventricle Linear
1.2.840.10008.6.​1.​595	CID 12202	Left Ventricle Volume
1.2.840.10008.6.​1.​596	CID 12203	Left Ventricle Other
1.2.840.10008.6.​1.​597	CID 12204	Echocardiography Right Ventricle
1.2.840.10008.6.​1.​598	CID 12205	Echocardiography Left Atrium
1.2.840.10008.6.​1.​599	CID 12206	Echocardiography Right Atrium
1.2.840.10008.6.​1.​600	CID 12207	Echocardiography Mitral Valve
1.2.840.10008.6.​1.​601	CID 12208	Echocardiography Tricuspid Valve
1.2.840.10008.6.​1.​602	CID 12209	Echocardiography Pulmonic Valve
1.2.840.10008.6.​1.​603	CID 12210	Echocardiography Pulmonary Artery
1.2.840.10008.6.​1.​604	CID 12211	Echocardiography Aortic Valve
1.2.840.10008.6.​1.​605	CID 12212	Echocardiography Aorta
1.2.840.10008.6.​1.​606	CID 12214	Echocardiography Pulmonary Veins
1.2.840.10008.6.​1.​607	CID 12215	Echocardiography Vena Cavae
1.2.840.10008.6.​1.​608	CID 12216	Echocardiography Hepatic Veins
1.2.840.10008.6.​1.​609	CID 12217	Echocardiography Cardiac Shunt
1.2.840.10008.6.​1.​610	CID 12218	Echocardiography Congenital
1.2.840.10008.6.​1.​611	CID 12219	Pulmonary Vein Modifiers
1.2.840.10008.6.​1.​612	CID 12220	Echocardiography Common Measurements
1.2.840.10008.6.​1.​613	CID 12221	Flow Direction
1.2.840.10008.6.​1.​614	CID 12222	Orifice Flow Properties
1.2.840.10008.6.​1.​615	CID 12223	Echocardiography Stroke Volume Origin
1.2.840.10008.6.​1.​616	CID 12224	Ultrasound Image Modes
1.2.840.10008.6.​1.​617	CID 12226	Echocardiography Image View
1.2.840.10008.6.​1.​618	CID 12227	Echocardiography Measurement Method
1.2.840.10008.6.​1.​619	CID 12228	Echocardiography Volume Methods
1.2.840.10008.6.​1.​620	CID 12229	Echocardiography Area Methods
1.2.840.10008.6.​1.​621	CID 12230	Gradient Methods
1.2.840.10008.6.​1.​622	CID 12231	Volume Flow Methods
1.2.840.10008.6.​1.​623	CID 12232	Myocardium Mass Methods
1.2.840.10008.6.​1.​624	CID 12233	Cardiac Phase
1.2.840.10008.6.​1.​625	CID 12234	Respiration State
1.2.840.10008.6.​1.​626	CID 12235	Mitral Valve Anatomic Sites
1.2.840.10008.6.​1.​627	CID 12236	Echo Anatomic Sites
1.2.840.10008.6.​1.​628	CID 12237	Echocardiography Anatomic Site Modifiers
1.2.840.10008.6.​1.​629	CID 12238	Wall Motion Scoring Schemes
1.2.840.10008.6.​1.​630	CID 12239	Cardiac Output Properties
1.2.840.10008.6.​1.​631	CID 12240	Left Ventricle Area
1.2.840.10008.6.​1.​632	CID 12241	Tricuspid Valve Finding Sites
1.2.840.10008.6.​1.​633	CID 12242	Aortic Valve Finding Sites
1.2.840.10008.6.​1.​634	CID 12243	Left Ventricle Finding Sites
1.2.840.10008.6.​1.​635	CID 12244	Congenital Finding Sites
1.2.840.10008.6.​1.​636	CID 7162	Surface Processing Algorithm Families
1.2.840.10008.6.​1.​637	CID 3207	Stress Test Procedure Phases
1.2.840.10008.6.​1.​638	CID 3778	Stages
1.2.840.10008.6.​1.​735	CID 252	S-M-L Size Descriptor
1.2.840.10008.6.​1.​736	CID 3016	Major Coronary Arteries
1.2.840.10008.6.​1.​737	CID 3083	Units of Radioactivity
1.2.840.10008.6.​1.​738	CID 3102	Rest-Stress
1.2.840.10008.6.​1.​739	CID 3106	PET Cardiology Protocols
1.2.840.10008.6.​1.​740	CID 3107	PET Cardiology Radiopharmaceuticals
1.2.840.10008.6.​1.​741	CID 3108	NM/PET Procedures
1.2.840.10008.6.​1.​742	CID 3110	Nuclear Cardiology Protocols
1.2.840.10008.6.​1.​743	CID 3111	Nuclear Cardiology Radiopharmaceuticals
1.2.840.10008.6.​1.​744	CID 3112	Attenuation Correction
1.2.840.10008.6.​1.​745	CID 3113	Types of Perfusion Defects
1.2.840.10008.6.​1.​746	CID 3114	Study Quality
1.2.840.10008.6.​1.​747	CID 3115	Stress Imaging Quality Issues
1.2.840.10008.6.​1.​748	CID 3116	NM Extracardiac Findings
1.2.840.10008.6.​1.​749	CID 3117	Attenuation Correction Methods
1.2.840.10008.6.​1.​750	CID 3118	Level of Risk
1.2.840.10008.6.​1.​751	CID 3119	LV Function
1.2.840.10008.6.​1.​752	CID 3120	Perfusion Findings
1.2.840.10008.6.​1.​753	CID 3121	Perfusion Morphology
1.2.840.10008.6.​1.​754	CID 3122	Ventricular Enlargement
1.2.840.10008.6.​1.​755	CID 3200	Stress Test Procedure
1.2.840.10008.6.​1.​756	CID 3201	Indications for Stress Test
1.2.840.10008.6.​1.​757	CID 3202	Chest Pain
1.2.840.10008.6.​1.​758	CID 3203	Exerciser Device
1.2.840.10008.6.​1.​759	CID 3204	Stress Agents
1.2.840.10008.6.​1.​760	CID 3205	Indications for Pharmacological Stress Test
1.2.840.10008.6.​1.​761	CID 3206	Non-invasive Cardiac Imaging Procedures
1.2.840.10008.6.​1.​763	CID 3208	Summary Codes Exercise ECG
1.2.840.10008.6.​1.​764	CID 3209	Summary Codes Stress Imaging
1.2.840.10008.6.​1.​765	CID 3210	Speed of Response
1.2.840.10008.6.​1.​766	CID 3211	BP Response
1.2.840.10008.6.​1.​767	CID 3212	Treadmill Speed
1.2.840.10008.6.​1.​768	CID 3213	Stress Hemodynamic Findings
1.2.840.10008.6.​1.​769	CID 3215	Perfusion Finding Method
1.2.840.10008.6.​1.​770	CID 3217	Comparison Finding
1.2.840.10008.6.​1.​771	CID 3220	Stress Symptoms
1.2.840.10008.6.​1.​772	CID 3221	Stress Test Termination Reasons
1.2.840.10008.6.​1.​773	CID 3227	QTc Measurements
1.2.840.10008.6.​1.​774	CID 3228	ECG Timing Measurements
1.2.840.10008.6.​1.​775	CID 3229	ECG Axis Measurements
1.2.840.10008.6.​1.​776	CID 3230	ECG Findings
1.2.840.10008.6.​1.​777	CID 3231	ST Segment Findings
1.2.840.10008.6.​1.​778	CID 3232	ST Segment Location
1.2.840.10008.6.​1.​779	CID 3233	ST Segment Morphology
1.2.840.10008.6.​1.​780	CID 3234	Ectopic Beat Morphology
1.2.840.10008.6.​1.​781	CID 3235	Perfusion Comparison Findings
1.2.840.10008.6.​1.​782	CID 3236	Tolerance Comparison Findings
1.2.840.10008.6.​1.​783	CID 3237	Wall Motion Comparison Findings
1.2.840.10008.6.​1.​784	CID 3238	Stress Scoring Scales
1.2.840.10008.6.​1.​785	CID 3239	Perceived Exertion Scales
1.2.840.10008.6.​1.​786	CID 3463	Ventricle Identification
1.2.840.10008.6.​1.​787	CID 6200	Colon Overall Assessment
1.2.840.10008.6.​1.​788	CID 6201	Colon Finding or Feature
1.2.840.10008.6.​1.​789	CID 6202	Colon Finding or Feature Modifier
1.2.840.10008.6.​1.​790	CID 6203	Colon Non-lesion Object Type
1.2.840.10008.6.​1.​791	CID 6204	Anatomic Non-colon Findings
1.2.840.10008.6.​1.​792	CID 6205	Clockface Location for Colon
1.2.840.10008.6.​1.​793	CID 6206	Recumbent Patient Orientation for Colon
1.2.840.10008.6.​1.​794	CID 6207	Colon Quantitative Temporal Difference Type
1.2.840.10008.6.​1.​795	CID 6208	Colon Types of Quality Control Standard
1.2.840.10008.6.​1.​796	CID 6209	Colon Morphology Descriptor
1.2.840.10008.6.​1.​797	CID 6210	Location in Intestinal Tract
1.2.840.10008.6.​1.​798	CID 6211	Colon CAD Material Description
1.2.840.10008.6.​1.​799	CID 6212	Calculated Value for Colon Findings
1.2.840.10008.6.​1.​800	CID 4214	Ophthalmic Horizontal Directions
1.2.840.10008.6.​1.​801	CID 4215	Ophthalmic Vertical Directions
1.2.840.10008.6.​1.​802	CID 4216	Ophthalmic Visual Acuity Type
1.2.840.10008.6.​1.​803	CID 3004	Arterial Pulse Waveform
1.2.840.10008.6.​1.​804	CID 3005	Respiration Waveform
1.2.840.10008.6.​1.​805	CID 12030	Ultrasound Contrast/Bolus Agents
1.2.840.10008.6.​1.​806	CID 12031	Protocol Interval Events
1.2.840.10008.6.​1.​807	CID 12032	Transducer Scan Pattern
1.2.840.10008.6.​1.​808	CID 12033	Ultrasound Transducer Geometry
1.2.840.10008.6.​1.​809	CID 12034	Ultrasound Transducer Beam Steering
1.2.840.10008.6.​1.​810	CID 12035	Ultrasound Transducer Application
1.2.840.10008.6.​1.​811	CID 50	Instance Availability Status
1.2.840.10008.6.​1.​812	CID 9301	Modality PPS Discontinuation Reasons
1.2.840.10008.6.​1.​813	CID 9302	Media Import PPS Discontinuation Reasons
1.2.840.10008.6.​1.​814	CID 7482	DX Anatomy Imaged for Animals
1.2.840.10008.6.​1.​815	CID 7483	Common Anatomic Regions for Animals
1.2.840.10008.6.​1.​816	CID 7484	DX View for Animals
1.2.840.10008.6.​1.​817	CID 7030	Institutional Departments, Units and Services
1.2.840.10008.6.​1.​818	CID 7009	Purpose of Reference to Predecessor Report
1.2.840.10008.6.​1.​819	CID 4220	Visual Fixation Quality During Acquisition
1.2.840.10008.6.​1.​820	CID 4221	Visual Fixation Quality Problem
1.2.840.10008.6.​1.​821	CID 4222	Ophthalmic Macular Grid Problem
1.2.840.10008.6.​1.​822	CID 5002	Organizations
1.2.840.10008.6.​1.​823	CID 7486	Mixed Breeds
1.2.840.10008.6.​1.​824	CID 7040	Broselow-Luten Pediatric Size Categories
1.2.840.10008.6.​1.​825	CID 7042	CMDCTECC Calcium Scoring Patient Size Categories
1.2.840.10008.6.​1.​826	CID 12245	Cardiac Ultrasound Report Titles
1.2.840.10008.6.​1.​827	CID 12246	Cardiac Ultrasound Indication for Study
1.2.840.10008.6.​1.​828	CID 12247	Pediatric, Fetal and Congenital Cardiac Surgical Interventions
1.2.840.10008.6.​1.​829	CID 12248	Cardiac Ultrasound Summary Codes
1.2.840.10008.6.​1.​830	CID 12249	Cardiac Ultrasound Fetal Summary Codes
1.2.840.10008.6.​1.​831	CID 12250	Cardiac Ultrasound Common Linear Measurements
1.2.840.10008.6.​1.​832	CID 12251	Cardiac Ultrasound Linear Valve Measurements
1.2.840.10008.6.​1.​833	CID 12252	Cardiac Ultrasound Cardiac Function
1.2.840.10008.6.​1.​834	CID 12253	Cardiac Ultrasound Area Measurements
1.2.840.10008.6.​1.​835	CID 12254	Cardiac Ultrasound Hemodynamic Measurements
1.2.840.10008.6.​1.​836	CID 12255	Cardiac Ultrasound Myocardium Measurements
1.2.840.10008.6.​1.​837
1.2.840.10008.6.​1.​838	CID 12257	Cardiac Ultrasound Left Ventricle
1.2.840.10008.6.​1.​839	CID 12258	Cardiac Ultrasound Right Ventricle
1.2.840.10008.6.​1.​840	CID 12259	Cardiac Ultrasound Ventricles Measurements
1.2.840.10008.6.​1.​841	CID 12260	Cardiac Ultrasound Pulmonary Artery
1.2.840.10008.6.​1.​842	CID 12261	Cardiac Ultrasound Pulmonary Vein
1.2.840.10008.6.​1.​843	CID 12262	Cardiac Ultrasound Pulmonary Valve
1.2.840.10008.6.​1.​844	CID 12263	Cardiac Ultrasound Venous Return Pulmonary Measurements
1.2.840.10008.6.​1.​845	CID 12264	Cardiac Ultrasound Venous Return Systemic Measurements
1.2.840.10008.6.​1.​846	CID 12265	Cardiac Ultrasound Atria and Atrial Septum Measurements
1.2.840.10008.6.​1.​847	CID 12266	Cardiac Ultrasound Mitral Valve
1.2.840.10008.6.​1.​848	CID 12267	Cardiac Ultrasound Tricuspid Valve
1.2.840.10008.6.​1.​849	CID 12268	Cardiac Ultrasound Atrioventricular Valves Measurements
1.2.840.10008.6.​1.​850	CID 12269	Cardiac Ultrasound Interventricular Septum Measurements
1.2.840.10008.6.​1.​851	CID 12270	Cardiac Ultrasound Aortic Valve
1.2.840.10008.6.​1.​852	CID 12271	Cardiac Ultrasound Outflow Tracts Measurements
1.2.840.10008.6.​1.​853	CID 12272	Cardiac Ultrasound Semilunar Valves, Annulate and Sinuses Measurements
1.2.840.10008.6.​1.​854	CID 12273	Cardiac Ultrasound Aortic Sinotubular Junction
1.2.840.10008.6.​1.​855	CID 12274	Cardiac Ultrasound Aorta Measurements
1.2.840.10008.6.​1.​856	CID 12275	Cardiac Ultrasound Coronary Arteries Measurements
1.2.840.10008.6.​1.​857	CID 12276	Cardiac Ultrasound Aorto Pulmonary Connections Measurements
1.2.840.10008.6.​1.​858	CID 12277	Cardiac Ultrasound Pericardium and Pleura Measurements
1.2.840.10008.6.​1.​859	CID 12279	Cardiac Ultrasound Fetal General Measurements
1.2.840.10008.6.​1.​860	CID 12280	Cardiac Ultrasound Target Sites
1.2.840.10008.6.​1.​861	CID 12281	Cardiac Ultrasound Target Site Modifiers
1.2.840.10008.6.​1.​862	CID 12282	Cardiac Ultrasound Venous Return Systemic Finding Sites
1.2.840.10008.6.​1.​863	CID 12283	Cardiac Ultrasound Venous Return Pulmonary Finding Sites
1.2.840.10008.6.​1.​864	CID 12284	Cardiac Ultrasound Atria and Atrial Septum Finding Sites
1.2.840.10008.6.​1.​865	CID 12285	Cardiac Ultrasound Atrioventricular Valves Finding Sites
1.2.840.10008.6.​1.​866	CID 12286	Cardiac Ultrasound Interventricular Septum Finding Sites
1.2.840.10008.6.​1.​867	CID 12287	Cardiac Ultrasound Ventricles Finding Sites
1.2.840.10008.6.​1.​868	CID 12288	Cardiac Ultrasound Outflow Tracts Finding Sites
1.2.840.10008.6.​1.​869	CID 12289	Cardiac Ultrasound Semilunar Valves, Annulus and Sinuses Finding Sites
1.2.840.10008.6.​1.​870	CID 12290	Cardiac Ultrasound Pulmonary Arteries Finding Sites
1.2.840.10008.6.​1.​871	CID 12291	Cardiac Ultrasound Aorta Finding Sites
1.2.840.10008.6.​1.​872	CID 12292	Cardiac Ultrasound Coronary Arteries Finding Sites
1.2.840.10008.6.​1.​873	CID 12293	Cardiac Ultrasound Aortopulmonary Connections Finding Sites
1.2.840.10008.6.​1.​874	CID 12294	Cardiac Ultrasound Pericardium and Pleura Finding Sites
1.2.840.10008.6.​1.​876	CID 4230	Ophthalmic Ultrasound Axial Measurements Type
1.2.840.10008.6.​1.​877	CID 4231	Lens Status
1.2.840.10008.6.​1.​878	CID 4232	Vitreous Status
1.2.840.10008.6.​1.​879	CID 4233	Ophthalmic Axial Length Measurements Segment Names
1.2.840.10008.6.​1.​880	CID 4234	Refractive Surgery Types
1.2.840.10008.6.​1.​881	CID 4235	Keratometry Descriptors
1.2.840.10008.6.​1.​882	CID 4236	IOL Calculation Formula
1.2.840.10008.6.​1.​883	CID 4237	Lens Constant Type
1.2.840.10008.6.​1.​884	CID 4238	Refractive Error Types
1.2.840.10008.6.​1.​885	CID 4239	Anterior Chamber Depth Definition
1.2.840.10008.6.​1.​886	CID 4240	Ophthalmic Measurement or Calculation Data Source
1.2.840.10008.6.​1.​887	CID 4241	Ophthalmic Axial Length Selection Method
1.2.840.10008.6.​1.​888
1.2.840.10008.6.​1.​889	CID 4243	Ophthalmic Quality Metric Type
1.2.840.10008.6.​1.​890	CID 4244	Ophthalmic Agent Concentration Units
1.2.840.10008.6.​1.​891	CID 91	Functional Condition Present During Acquisition
1.2.840.10008.6.​1.​892	CID 92	Joint Position During Acquisition
1.2.840.10008.6.​1.​893	CID 93	Joint Positioning Method
1.2.840.10008.6.​1.​894	CID 94	Physical Force Applied During Acquisition
1.2.840.10008.6.​1.​895	CID 3690	ECG Control Variables Numeric
1.2.840.10008.6.​1.​896	CID 3691	ECG Control Variables Text
1.2.840.10008.6.​1.​897	CID 8120	WSI Referenced Image Purposes of Reference
1.2.840.10008.6.​1.​898	CID 8121	Microscopy Lens Type
1.2.840.10008.6.​1.​899	CID 8122	Microscopy Illuminator and Sensor Color
1.2.840.10008.6.​1.​900	CID 8123	Microscopy Illumination Method
1.2.840.10008.6.​1.​901	CID 8124	Microscopy Filter
1.2.840.10008.6.​1.​902	CID 8125	Microscopy Illuminator Type
1.2.840.10008.6.​1.​903	CID 400	Audit Event ID
1.2.840.10008.6.​1.​904	CID 401	Audit Event Type Code
1.2.840.10008.6.​1.​905	CID 402	Audit Active Participant Role ID Code
1.2.840.10008.6.​1.​906	CID 403	Security Alert Type Code
1.2.840.10008.6.​1.​907	CID 404	Audit Participant Object ID Type Code
1.2.840.10008.6.​1.​908	CID 405	Media Type Code
1.2.840.10008.6.​1.​909	CID 4250	Visual Field Static Perimetry Test Patterns
1.2.840.10008.6.​1.​910	CID 4251	Visual Field Static Perimetry Test Strategies
1.2.840.10008.6.​1.​911	CID 4252	Visual Field Static Perimetry Screening Test Modes
1.2.840.10008.6.​1.​912	CID 4253	Visual Field Static Perimetry Fixation Strategy
1.2.840.10008.6.​1.​913	CID 4254	Visual Field Static Perimetry Test Analysis Results
1.2.840.10008.6.​1.​914	CID 4255	Visual Field Illumination Color
1.2.840.10008.6.​1.​915	CID 4256	Visual Field Procedure Modifier
1.2.840.10008.6.​1.​916	CID 4257	Visual Field Global Index Name
1.2.840.10008.6.​1.​917	CID 7180	Abstract Multi-dimensional Image Model Component Semantics
1.2.840.10008.6.​1.​918	CID 7181	Abstract Multi-dimensional Image Model Component Units
1.2.840.10008.6.​1.​919	CID 7182	Abstract Multi-dimensional Image Model Dimension Semantics
1.2.840.10008.6.​1.​920	CID 7183	Abstract Multi-dimensional Image Model Dimension Units
1.2.840.10008.6.​1.​921	CID 7184	Abstract Multi-dimensional Image Model Axis Direction
1.2.840.10008.6.​1.​922	CID 7185	Abstract Multi-dimensional Image Model Axis Orientation
1.2.840.10008.6.​1.​923	CID 7186	Abstract Multi-dimensional Image Model Qualitative Dimension Sample Semantics
1.2.840.10008.6.​1.​924	CID 7320	Planning Methods
1.2.840.10008.6.​1.​925	CID 7050	De-identification Method
1.2.840.10008.6.​1.​926	CID 12118	Measurement Orientation
1.2.840.10008.6.​1.​927	CID 3689	ECG Global Waveform Durations
1.2.840.10008.6.​1.​930	CID 3692	ICDs
1.2.840.10008.6.​1.​931	CID 9241	Radiotherapy General Workitem Definition
1.2.840.10008.6.​1.​932	CID 9242	Radiotherapy Acquisition Workitem Definition
1.2.840.10008.6.​1.​933	CID 9243	Radiotherapy Registration Workitem Definition
1.2.840.10008.6.​1.​934	CID 3850	Contrast Bolus Substance
1.2.840.10008.6.​1.​935	CID 10022	Label Types
1.2.840.10008.6.​1.​936	CID 4260	Ophthalmic Mapping Units for Real World Value Mapping
1.2.840.10008.6.​1.​937	CID 4261	Ophthalmic Mapping Acquisition Method
1.2.840.10008.6.​1.​938	CID 4262	Retinal Thickness Definition
1.2.840.10008.6.​1.​939	CID 4263	Ophthalmic Thickness Map Value Type
1.2.840.10008.6.​1.​940	CID 4264	Ophthalmic Map Purposes of Reference
1.2.840.10008.6.​1.​941	CID 4265	Ophthalmic Thickness Deviation Categories
1.2.840.10008.6.​1.​942	CID 4266	Ophthalmic Anatomic Structure Reference Point
1.2.840.10008.6.​1.​943	CID 3104	Cardiac Synchronization Technique
1.2.840.10008.6.​1.​944	CID 8130	Staining Protocols
1.2.840.10008.6.​1.​945
1.2.840.10008.6.​1.​946
1.2.840.10008.6.​1.​947	CID 10023	Size Specific Dose Estimation Method for CT
1.2.840.10008.6.​1.​948	CID 8131	Pathology Imaging Protocols
1.2.840.10008.6.​1.​949	CID 8132	Magnification Selection
1.2.840.10008.6.​1.​950	CID 8133	Tissue Selection
1.2.840.10008.6.​1.​951	CID 7464	General Region of Interest Measurement Modifiers
1.2.840.10008.6.​1.​952	CID 7465	Measurements Derived From Multiple ROI Measurements
1.2.840.10008.6.​1.​953	CID 8201	Surface Scan Acquisition Types
1.2.840.10008.6.​1.​954	CID 8202	Surface Scan Mode Types
1.2.840.10008.6.​1.​956	CID 8203	Surface Scan Registration Method Types
1.2.840.10008.6.​1.​957	CID 27	Basic Cardiac Views
1.2.840.10008.6.​1.​958	CID 10033	CT Reconstruction Algorithm
1.2.840.10008.6.​1.​959	CID 10030	Detector Types
1.2.840.10008.6.​1.​960	CID 10031	CR/DR Mechanical Configuration
1.2.840.10008.6.​1.​961	CID 10032	Projection X-Ray Acquisition Device Types
1.2.840.10008.6.​1.​962	CID 7165	Abstract Segmentation Types
1.2.840.10008.6.​1.​963	CID 7166	Common Tissue Segmentation Types
1.2.840.10008.6.​1.​964	CID 7167	Peripheral Nervous System Segmentation Types
1.2.840.10008.6.​1.​965	CID 4267	Corneal Topography Mapping Units for Real World Value Mapping
1.2.840.10008.6.​1.​966	CID 4268	Corneal Topography Map Value Type
1.2.840.10008.6.​1.​967	CID 7140	Brain Structures for Volumetric Measurements
1.2.840.10008.6.​1.​968	CID 7220	RT Dose Derivation
1.2.840.10008.6.​1.​969	CID 7221	RT Dose Purpose of Reference
1.2.840.10008.6.1.970	CID 7215	Spectroscopy Purpose of Reference
1.2.840.10008.6.​1.​971	CID 9250	Scheduled Processing Parameter Concept Codes for RT Treatment
1.2.840.10008.6.1.972	CID 10040	Radiopharmaceutical Organ Dose Reference Authority
1.2.840.10008.6.1.973	CID 10041	Source of Radioisotope Activity Information
1.2.840.10008.6.1.975	CID 10043	Intravenous Extravasation Symptoms
1.2.840.10008.6.1.976	CID 10044	Radiosensitive Organs
1.2.840.10008.6.1.977	CID 10045	Radiopharmaceutical Patient State
1.2.840.10008.6.1.978	CID 10046	GFR Measurements
1.2.840.10008.6.1.979	CID 10047	GFR Measurement Methods
1.2.840.10008.6.1.980	CID 8300	Visual Evaluation Methods
1.2.840.10008.6.1.981	CID 8301	Test Pattern Codes
1.2.840.10008.6.1.982	CID 8302	Measurement Pattern Codes
1.2.840.10008.6.1.983	CID 8303	Display Device Type
1.2.840.10008.6.1.984	CID 85	SUV Units
1.2.840.10008.6.1.985	CID 4100	T1 Measurement Methods
1.2.840.10008.6.1.986	CID 4101	Tracer Kinetic Models
1.2.840.10008.6.1.987	CID 4102	Perfusion Measurement Methods
1.2.840.10008.6.1.988	CID 4103	Arterial Input Function Measurement Methods
1.2.840.10008.6.1.989	CID 4104	Bolus Arrival Time Derivation Methods
1.2.840.10008.6.1.990	CID 4105	Perfusion Analysis Methods
1.2.840.10008.6.1.991	CID 4106	Quantitative Methods used for Perfusion And Tracer Kinetic Models
1.2.840.10008.6.1.992	CID 4107	Tracer Kinetic Model Parameters
1.2.840.10008.6.1.993	CID 4108	Perfusion Model Parameters
1.2.840.10008.6.1.994	CID 4109	Model-Independent Dynamic Contrast Analysis Parameters
1.2.840.10008.6.1.995	CID 4110	Tracer Kinetic Modeling Covariates
1.2.840.10008.6.1.996	CID 4111	Contrast Characteristics
1.2.840.10008.6.1.997	CID 7021	Measurement Report Document Titles
1.2.840.10008.6.1.998	CID 100	Quantitative Diagnostic Imaging Procedures
1.2.840.10008.6.1.999	CID 7466	PET Region of Interest Measurements
1.2.840.10008.6.1.1000	CID 7467	Gray Level Co-occurrence Matrix Measurements
1.2.840.10008.6.1.1001	CID 7468	Texture Measurements
1.2.840.10008.6.1.1002	CID 6146	Time Point Types
1.2.840.10008.6.1.1003	CID 7469	Generic Intensity and Size Measurements
1.2.840.10008.6.1.1004	CID 6147	Response Criteria
1.2.840.10008.6.1.1005	CID 12020	Fetal Biometry Anatomic Sites
1.2.840.10008.6.1.1006	CID 12021	Fetal Long Bone Anatomic Sites
1.2.840.10008.6.1.1007	CID 12022	Fetal Cranium Anatomic Sites
1.2.840.10008.6.1.1008	CID 12023	Pelvis and Uterus Anatomic Sites
1.2.840.10008.6.1.1009	CID 7222	Parametric Map Derivation Image Purpose of Reference
1.2.840.10008.6.1.1010	CID 9000	Physical Quantity Descriptors
1.2.840.10008.6.1.1011	CID 7600	Lymph Node Anatomic Sites
1.2.840.10008.6.1.1012	CID 7601	Head and Neck Cancer Anatomic Sites
1.2.840.10008.6.1.1013	CID 7701	Fiber Tracts In Brainstem
1.2.840.10008.6.1.1014	CID 7702	Projection and Thalamic Fibers
1.2.840.10008.6.1.1015	CID 7703	Association Fibers
1.2.840.10008.6.1.1016	CID 7704	Limbic System Tracts
1.2.840.10008.6.1.1017	CID 7705	Commissural Fibers
1.2.840.10008.6.1.1018	CID 7706	Cranial Nerves
1.2.840.10008.6.1.1019	CID 7707	Spinal Cord Fibers
1.2.840.10008.6.1.1020	CID 7710	Tractography Anatomic Sites
1.2.840.10008.6.1.1021	CID 4025	Primary Anatomic Structure for Intra-oral Radiography (Supernumerary Dentition - Designation of Teeth)
1.2.840.10008.6.1.1022	CID 4026	Primary Anatomic Structure for Intra-oral and Craniofacial Radiography - Teeth
1.2.840.10008.6.1.1023	CID 9401	IEC61217 Device Position Parameters
1.2.840.10008.6.1.1024	CID 9402	IEC61217 Gantry Position Parameters
1.2.840.10008.6.1.1025	CID 9403	IEC61217 Patient Support Position Parameters
1.2.840.10008.6.1.1026	CID 7035	Actionable Finding Classification
1.2.840.10008.6.1.1027	CID 7036	Image Quality Assessment
1.2.840.10008.6.1.1028	CID 10050	Summary Radiation Exposure Quantities
1.2.840.10008.6.1.1029	CID 4245	Wide Field Ophthalmic Photography Transformation Method
1.2.840.10008.6.1.1030	CID 84	PET Units
1.2.840.10008.6.1.1031	CID 7300	Implant Materials
1.2.840.10008.6.1.1032	CID 7301	Intervention Types
1.2.840.10008.6.1.1033	CID 7302	Implant Templates View Orientations
1.2.840.10008.6.1.1034	CID 7303	Implant Templates Modified View Orientations
1.2.840.10008.6.1.1035	CID 7304	Implant Target Anatomy
1.2.840.10008.6.1.1036	CID 7305	Implant Planning Landmarks
1.2.840.10008.6.1.1037	CID 7306	Human Hip Implant Planning Landmarks
1.2.840.10008.6.1.1038	CID 7307	Implant Component Types
1.2.840.10008.6.1.1039	CID 7308	Human Hip Implant Component Types
1.2.840.10008.6.1.1040	CID 7309	Human Trauma Implant Component Types
1.2.840.10008.6.1.1041	CID 7310	Implant Fixation Method
1.2.840.10008.6.1.1042	CID 7445	Device Participating Roles
1.2.840.10008.6.1.1043	CID 8101	Container Types
1.2.840.10008.6.1.1044	CID 8102	Container Component Types
1.2.840.10008.6.1.1045	CID 8103	Anatomic Pathology Specimen Types
1.2.840.10008.6.1.1046	CID 8104	Breast Tissue Specimen Types
1.2.840.10008.6.1.1047	CID 8109	Specimen Collection Procedure
1.2.840.10008.6.1.1048	CID 8110	Specimen Sampling Procedure
1.2.840.10008.6.1.1049	CID 8111	Specimen Preparation Procedure
1.2.840.10008.6.1.1050	CID 8112	Specimen Stains
1.2.840.10008.6.1.1051	CID 8113	Specimen Preparation Steps
1.2.840.10008.6.1.1052	CID 8114	Specimen Fixatives
1.2.840.10008.6.1.1053	CID 8115	Specimen Embedding Media
1.2.840.10008.6.1.1054	CID 10020	Source of Projection X-Ray Dose Information
1.2.840.10008.6.1.1055	CID 10021	Source of CT Dose Information
1.2.840.10008.6.1.1056	CID 10025	Radiation Dose Reference Points
1.2.840.10008.6.1.1057	CID 501	Volumetric View Description
1.2.840.10008.6.1.1058	CID 502	Volumetric View Modifier
1.2.840.10008.6.1.1059	CID 7260	Diffusion Acquisition Value Types
1.2.840.10008.6.1.1060	CID 7261	Diffusion Model Value Types
1.2.840.10008.6.1.1061	CID 7262	Diffusion Tractography Algorithm Families
1.2.840.10008.6.1.1062	CID 7263	Diffusion Tractography Measurement Types
1.2.840.10008.6.​1.​1063	CID 7490	Research Animal Source Registries
1.2.840.10008.6.​1.​1064	CID 231	Yes-No Only
1.2.840.10008.6.​1.​1065	CID 601	Biosafety Levels
1.2.840.10008.6.​1.​1066	CID 602	Biosafety Control Reasons
1.2.840.10008.6.​1.​1067	CID 7457	Sex - Male Female or Both
1.2.840.10008.6.​1.​1068	CID 603	Animal Room Types
1.2.840.10008.6.​1.​1069	CID 604	Device Reuse
1.2.840.10008.6.​1.​1070	CID 605	Animal Bedding Material
1.2.840.10008.6.​1.​1071	CID 606	Animal Shelter Types
1.2.840.10008.6.​1.​1072	CID 607	Animal Feed Types
1.2.840.10008.6.​1.​1073	CID 608	Animal Feed Sources
1.2.840.10008.6.​1.​1074	CID 609	Animal Feeding Methods
1.2.840.10008.6.​1.​1075	CID 610	Water Types
1.2.840.10008.6.​1.​1076	CID 611	Anesthesia Category Code Type for Small Animal Anesthesia
1.2.840.10008.6.​1.​1077	CID 612	Anesthesia Category Code Type from Anesthesia Quality Initiative (AQI)
1.2.840.10008.6.​1.​1078	CID 613	Anesthesia Induction Code Type for Small Animal Anesthesia
1.2.840.10008.6.​1.​1079	CID 614	Anesthesia Induction Code Type from Anesthesia Quality Initiative (AQI)
1.2.840.10008.6.​1.​1080	CID 615	Anesthesia Maintenance Code Type for Small Animal Anesthesia
1.2.840.10008.6.​1.​1081	CID 616	Anesthesia Maintenance Code Type from Anesthesia Quality Initiative (AQI)
1.2.840.10008.6.​1.​1082	CID 617	Airway Management Method Code Type for Small Animal Anesthesia
1.2.840.10008.6.​1.​1083	CID 618	Airway Management Method Code Type from Anesthesia Quality Initiative (AQI)
1.2.840.10008.6.​1.​1084	CID 619	Airway Management Sub-Method Code Type for Small Animal Anesthesia
1.2.840.10008.6.​1.​1085	CID 620	Airway Management Sub-Method Code Type from Anesthesia Quality Initiative (AQI)
1.2.840.10008.6.​1.​1086	CID 621	Type of Medication for Small Animal Anesthesia
1.2.840.10008.6.​1.​1087	CID 622	Medication Type Code Type from Anesthesia Quality Initiative (AQI)
1.2.840.10008.6.​1.​1088	CID 623	Medication for Small Animal Anesthesia
1.2.840.10008.6.​1.​1089	CID 624	Inhalational Anesthesia Agents for Small Animal Anesthesia
1.2.840.10008.6.​1.​1090	CID 625	Injectable Anesthesia Agents for Small Animal Anesthesia
1.2.840.10008.6.​1.​1091	CID 626	Premedication Agents for Small Animal Anesthesia
1.2.840.10008.6.​1.​1092	CID 627	Neuromuscular Blocking Agents for Small Animal Anesthesia
1.2.840.10008.6.​1.​1093	CID 628	Ancillary Medications for Small Animal Anesthesia
1.2.840.10008.6.​1.​1094	CID 629	Carrier Gases for Small Animal Anesthesia
1.2.840.10008.6.​1.​1095	CID 630	Local Anesthetics for Small Animal Anesthesia
1.2.840.10008.6.​1.​1096	CID 631	Phase of Procedure Requiring Anesthesia
1.2.840.10008.6.​1.​1097	CID 632	Phase of Surgical Procedure Requiring Anesthesia
1.2.840.10008.6.​1.​1098	CID 633	Phase of Imaging Procedure Requiring Anesthesia
1.2.840.10008.6.​1.​1099	CID 634	Phase of Animal Handling
1.2.840.10008.6.​1.​1100	CID 635	Heating Method
1.2.840.10008.6.​1.​1101	CID 636	Temperature Sensor Device Component Type for Small Animal Procedures
1.2.840.10008.6.​1.​1102	CID 637	Exogenous Substance Types
1.2.840.10008.6.​1.​1103	CID 638	Exogenous Substance
1.2.840.10008.6.​1.​1104	CID 639	Tumor Graft Histologic Type
1.2.840.10008.6.​1.​1105	CID 640	Fibrils
1.2.840.10008.6.​1.​1106	CID 641	Viruses
1.2.840.10008.6.​1.​1107	CID 642	Cytokines
1.2.840.10008.6.​1.​1108	CID 643	Toxins
1.2.840.10008.6.​1.​1109	CID 644	Exogenous Substance Administration Sites
1.2.840.10008.6.​1.​1110	CID 645	Exogenous Substance Tissue of Origin
1.2.840.10008.6.​1.​1111	CID 646	Preclinical Small Animal Imaging Procedures
1.2.840.10008.6.​1.​1112	CID 647	Position Reference Indicator for Frame of Reference
1.2.840.10008.6.​1.​1113	CID 241	Present-Absent Only
1.2.840.10008.6.​1.​1114	CID 10024	Water Equivalent Diameter Method
1.2.840.10008.6.​1.​1115	CID 7022	Radiotherapy Purposes of Reference
1.2.840.10008.6.​1.​1116	CID 701	Content Assessment Types
1.2.840.10008.6.​1.​1117	CID 702	RT Content Assessment Types
1.2.840.10008.6.​1.​1118	CID 703	Basis of Assessment
1.2.840.10008.6.​1.​1119	CID 7449	Reader Specialty
1.2.840.10008.6.​1.​1120	CID 9233	Requested Report Types
1.2.840.10008.6.1.1121	CID 1000	CT Transverse Plane Reference Basis
1.2.840.10008.6.1.1122	CID 1001	Anatomical Reference Basis
1.2.840.10008.6.1.1123	CID 1002	Anatomical Reference Basis - Head
1.2.840.10008.6.1.1124	CID 1003	Anatomical Reference Basis - Spine
1.2.840.10008.6.1.1125	CID 1004	Anatomical Reference Basis - Chest
1.2.840.10008.6.1.1126	CID 1005	Anatomical Reference Basis - Abdomen/Pelvis
1.2.840.10008.6.1.1127	CID 1006	Anatomical Reference Basis - Extremities
1.2.840.10008.6.1.1128	CID 1010	Reference Geometry - Planes
1.2.840.10008.6.1.1129	CID 1011	Reference Geometry - Points
1.2.840.10008.6.1.1130	CID 1015	Patient Alignment Methods
1.2.840.10008.6.1.1131	CID 1200	Contraindications For CT Imaging
1.2.840.10008.6.1.1132	CID 7110	Fiducials Categories
1.2.840.10008.6.1.1133	CID 7111	Fiducials
1.2.840.10008.6.1.1134	CID 7013	Non-Image Source Instance Purposes of Reference
1.2.840.10008.6.1.1135	CID 7023	RT Process Output
1.2.840.10008.6.1.1136	CID 7024	RT Process Input
1.2.840.10008.6.1.1137	CID 7025	RT Process Input Used
1.2.840.10008.6.1.1138	CID 6300	Prostate Sector Anatomy
1.2.840.10008.6.1.1139	CID 6301	Prostate Sector Anatomy from PI-RADS v2
1.2.840.10008.6.1.1140	CID 6302	Prostate Sector Anatomy from European Concensus 16 Sector (Minimal) Model
1.2.840.10008.6.1.1141	CID 6303	Prostate Sector Anatomy from European Concensus 27 Sector (Optimal) Model
1.2.840.10008.6.1.1142	CID 12301	Measurement Selection Reasons
1.2.840.10008.6.1.1143	CID 12302	Echo Finding Observation Types
1.2.840.10008.6.1.1144	CID 12303	Echo Measurement Types
1.2.840.10008.6.1.1145	CID 12304	Echo Measured Properties
1.2.840.10008.6.1.1146	CID 12305	Basic Echo Anatomic Sites
1.2.840.10008.6.1.1147	CID 12306	Echo Flow Directions
1.2.840.10008.6.1.1148	CID 12307	Cardiac Phases and Time Points
1.2.840.10008.6.1.1149	CID 12300	Core Echo Measurements
1.2.840.10008.6.1.1150	CID 4270	OCT-A Processing Algorithm Families
1.2.840.10008.6.1.1151	CID 4271	En Face Image Types
1.2.840.10008.6.1.1152	CID 4272	Opt Scan Pattern Types
1.2.840.10008.6.1.1153	CID 4273	Retinal Segmentation Surfaces
1.2.840.10008.6.1.1154	CID 10060	Organs for Radiation Dose Estimates
1.2.840.10008.6.1.1155	CID 10061	Absorbed Radiation Dose Types
1.2.840.10008.6.1.1156	CID 10062	Equivalent Radiation Dose Types
1.2.840.10008.6.1.1157	CID 10063	Radiation Dose Estimate Distribution Representation
1.2.840.10008.6.1.1158	CID 10064	Patient Model Type
1.2.840.10008.6.1.1159	CID 10065	Radiation Transport Model Type
1.2.840.10008.6.1.1160	CID 10066	Attenuator Category
1.2.840.10008.6.1.1161	CID 10067	Radiation Attenuator Materials
1.2.840.10008.6.1.1162	CID 10068	Estimate Method Types
1.2.840.10008.6.1.1163	CID 10069	Radiation Dose Estimation Parameter
1.2.840.10008.6.1.1164	CID 10070	Radiation Dose Types
1.2.840.10008.6.1.1165	CID 7270	MR Diffusion Component Semantics
1.2.840.10008.6.1.1166	CID 7271	MR Diffusion Anisotropy Indices
1.2.840.10008.6.1.1167	CID 7272	MR Diffusion Model Parameters
1.2.840.10008.6.1.1168	CID 7273	MR Diffusion Models
1.2.840.10008.6.1.1169	CID 7274	MR Diffusion Model Fitting Methods
1.2.840.10008.6.1.1170	CID 7275	MR Diffusion Model Specific Methods
1.2.840.10008.6.1.1171	CID 7276	MR Diffusion Model Inputs
1.2.840.10008.6.1.1172	CID 7277	Units of Diffusion Rate Area Over Time
1.2.840.10008.6.1.1173	CID 7039	Pediatric Size Categories
1.2.840.10008.6.1.1174	CID 7041	Calcium Scoring Patient Size Categories
1.2.840.10008.6.1.1175	CID 10034	Reason for Repeating Acquisition
1.2.840.10008.6.1.1176	CID 800	Protocol Assertion Codes
1.2.840.10008.6.1.1177	CID 7026	Radiotherapeutic Dose Measurement Devices
1.2.840.10008.6.1.1178	CID 7014	Export Additional Information Document Titles
1.2.840.10008.6.1.1179	CID 7015	Export Delay Reasons
1.2.840.10008.6.1.1180	CID 7016	Level of Difficulty
1.2.840.10008.6.1.1181	CID 7017	Category of Teaching Material - Imaging
1.2.840.10008.6.1.1182	CID 7018	Miscellaneous Document Titles
1.2.840.10008.6.1.1183	CID 7019	Segmentation Non-Image Source Purposes of Reference
1.2.840.10008.6.1.1184	CID 280	Longitudinal Temporal Event Types
1.2.840.10008.6.1.1185	CID 6401	Non-lesion Object Type - Physical Objects
1.2.840.10008.6.1.1186	CID 6402	Non-lesion Object Type - Substances
1.2.840.10008.6.1.1187	CID 6403	Non-lesion Object Type - Tissues
1.2.840.10008.6.1.1188	CID 6404	Chest Non-lesion Object Type - Physical Objects
1.2.840.10008.6.1.1189	CID 6405	Chest Non-lesion Object Type - Tissues
1.2.840.10008.6.1.1190	CID 7191	Tissue Segmentation Property Types
1.2.840.10008.6.1.1191	CID 7192	Anatomical Structure Segmentation Property Types
1.2.840.10008.6.1.1192	CID 7193	Physical Object Segmentation Property Types
1.2.840.10008.6.1.1193	CID 7194	Morphologically Abnormal Structure Segmentation Property Types
1.2.840.10008.6.1.1194	CID 7195	Function Segmentation Property Types
1.2.840.10008.6.1.1195	CID 7196	Spatial and Relational Concept Segmentation Property Types
1.2.840.10008.6.1.1196	CID 7197	Body Substance Segmentation Property Types
1.2.840.10008.6.1.1197	CID 7198	Substance Segmentation Property Types
1.2.840.10008.6.1.1198	CID 9303	Interpretation Request Discontinuation Reasons
1.2.840.10008.6.1.1199	CID 7475	Gray Level Run Length Based Features
1.2.840.10008.6.1.1200	CID 7476	Gray Level Size Zone Based Features
1.2.840.10008.6.1.1201	CID 7060	Encapsulated Document Source Purposes of Reference
1.2.840.10008.6.1.1202	CID 7061	Model Document Titles
1.2.840.10008.6.1.1203	CID 7062	Purpose of Reference to Predecessor 3D Model
1.2.840.10008.6.1.1204	CID 7063	Model Scale Units
1.2.840.10008.6.1.1205	CID 7064	Model Usage
1.2.840.10008.6.1.1206	CID 10071	Radiation Dose Units
1.2.840.10008.6.1.1207	CID 7112	Radiotherapy Fiducials
1.2.840.10008.6.1.1208	CID 300	Multi-energy Relevant Materials
1.2.840.10008.6.1.1209	CID 301	Multi-energy Material Units
1.2.840.10008.6.1.1210	CID 9500	Dosimetric Objective Types
1.2.840.10008.6.1.1211	CID 9501	Prescription Anatomy Categories
1.2.840.10008.6.1.1212	CID 9502	RT Segment Annotation Categories
1.2.840.10008.6.1.1213	CID 9503	Radiotherapy Therapeutic Role Categories
1.2.840.10008.6.1.1214	CID 9504	RT Geometric Information
1.2.840.10008.6.1.1215	CID 9505	Fixation or Positioning Devices
1.2.840.10008.6.1.1216	CID 9506	Brachytherapy Devices
1.2.840.10008.6.1.1217	CID 9507	External Body Models
1.2.840.10008.6.1.1218	CID 9508	Non-specific Volumes
1.2.840.10008.6.1.1219	CID 9509	Purpose of Reference For RT Physician Intent Input
1.2.840.10008.6.1.1220	CID 9510	Purpose of Reference For RT Treatment Planning Input
1.2.840.10008.6.1.1221	CID 9511	General External Radiotherapy Procedure Techniques
1.2.840.10008.6.1.1222	CID 9512	Tomotherapeutic Radiotherapy Procedure Techniques
1.2.840.10008.6.1.1223	CID 9513	Fixation Devices
1.2.840.10008.6.1.1224	CID 9514	Anatomical Structures For Radiotherapy
1.2.840.10008.6.1.1225	CID 9515	RT Patient Support Devices
1.2.840.10008.6.1.1226	CID 9516	Radiotherapy Bolus Device Types
1.2.840.10008.6.1.1227	CID 9517	Radiotherapy Block Device Types
1.2.840.10008.6.1.1228	CID 9518	Radiotherapy Accessory No-slot Holder Device Types
1.2.840.10008.6.1.1229	CID 9519	Radiotherapy Accessory Slot Holder Device Types
1.2.840.10008.6.1.1230	CID 9520	Segmented RT Accessory Devices
1.2.840.10008.6.1.1231	CID 9521	Radiotherapy Treatment Energy Unit
1.2.840.10008.6.1.1232	CID 9522	Multi-source Radiotherapy Procedure Techniques
1.2.840.10008.6.1.1233	CID 9523	Robotic Radiotherapy Procedure Techniques
1.2.840.10008.6.1.1234	CID 9524	Radiotherapy Procedure Techniques
1.2.840.10008.6.1.1235	CID 9525	Radiation Therapy Particle
1.2.840.10008.6.1.1236	CID 9526	Ion Therapy Particle
1.2.840.10008.6.1.1237	CID 9527	Teletherapy Isotope
1.2.840.10008.6.1.1238	CID 9528	Brachytherapy Isotope
1.2.840.10008.6.1.1239	CID 9529	Single Dose Dosimetric Objectives
1.2.840.10008.6.1.1240	CID 9530	Percentage and Dose Dosimetric Objectives
1.2.840.10008.6.1.1241	CID 9531	Volume and Dose Dosimetric Objectives
1.2.840.10008.6.1.1242	CID 9532	No-Parameter Dosimetric Objectives
1.2.840.10008.6.1.1243	CID 9533	Delivery Time Structure
1.2.840.10008.6.1.1244	CID 9534	Radiotherapy Targets
1.2.840.10008.6.1.1245	CID 9535	Radiotherapy Dose Calculation Roles
1.2.840.10008.6.1.1246	CID 9536	Radiotherapy Prescribing and Segmenting Person Roles
1.2.840.10008.6.1.1247	CID 9537	Effective Dose Calculation Method Categories
1.2.840.10008.6.1.1248	CID 9538	Radiation Transport-based Effective Dose Method Modifiers
1.2.840.10008.6.1.1249	CID 9539	Fractionation-based Effective Dose Method Modifiers
1.2.840.10008.6.1.1250	CID 60	Imaging Agent Administration Adverse Events
1.2.840.10008.6.1.1251	CID 61	Time Relative to Procedure
1.2.840.10008.6.1.1252	CID 62	Imaging Agent Administration Phase Type
1.2.840.10008.6.1.1253	CID 63	Imaging Agent Administration Mode
1.2.840.10008.6.1.1254	CID 64	Imaging Agent Administration Patient State
1.2.840.10008.6.1.1255	CID 65	Pre-medication For Imaging Agent Administration
1.2.840.10008.6.1.1256	CID 66	Medication For Imaging Agent Administration
1.2.840.10008.6.1.1257	CID 67	Imaging Agent Administration Completion Status
1.2.840.10008.6.1.1258	CID 68	Imaging Agent Administration Pharmaceutical Unit of Presentation
1.2.840.10008.6.1.1259	CID 69	Imaging Agent Administration Consumables
1.2.840.10008.6.1.1260	CID 70	Flush
1.2.840.10008.6.1.1261	CID 71	Imaging Agent Administration Injector Event Type
1.2.840.10008.6.1.1262	CID 72	Imaging Agent Administration Step Type
1.2.840.10008.6.1.1263	CID 73	Bolus Shaping Curves
1.2.840.10008.6.1.1264	CID 74	Imaging Agent Administration Consumable Catheter Type
1.2.840.10008.6.1.1265	CID 75	Low-high-equal
1.2.840.10008.6.1.1266	CID 76	Type of Pre-medication
1.2.840.10008.6.1.1267	CID 245	Laterality with Median
1.2.840.10008.6.1.1268	CID 4029	Dermatology Anatomic Sites
1.2.840.10008.6.1.1269	CID 218	Quantitative Image Features
1.2.840.10008.6.1.1270	CID 7477	Global Shape Descriptors
1.2.840.10008.6.1.1271	CID 7478	Intensity Histogram Features
1.2.840.10008.6.1.1272	CID 7479	Grey Level Distance Zone Based Features
1.2.840.10008.6.1.1273	CID 7500	Neighbourhood Grey Tone Difference Based Features
1.2.840.10008.6.1.1274	CID 7501	Neighbouring Grey Level Dependence Based Features
1.2.840.10008.6.1.1275	CID 4242	Cornea Measurement Method Descriptors
1.2.840.10008.6.1.1276	CID 7027	Segmented Radiotherapeutic Dose Measurement Devices
1.2.840.10008.6.1.1277	CID 6098	Clinical Course of Disease
1.2.840.10008.6.1.1278	CID 6099	Racial Group

Note
For some Context Group UIDs, no Context Group Name or Identifier is specified; these are "placeholders" that are not assigned but will not be reused.
Table A-4. Template UID Values
UID Value	UID Name	UID Type	Part
1.2.840.10008.9.1	Imaging Report	Document TemplateID	PS3.20
1.2.840.10008.9.2	Clinical Information	Section TemplateID	PS3.20
1.2.840.10008.9.3	Imaging Procedure Description	Section TemplateID	PS3.20
1.2.840.10008.9.4	Comparison Study	Section TemplateID	PS3.20
1.2.840.10008.9.5	Impression	Section TemplateID	PS3.20
1.2.840.10008.9.6	Addendum	Section TemplateID	PS3.20
1.2.840.10008.9.7	Request	Section TemplateID	PS3.20
1.2.840.10008.9.8	Radiation Exposure and Protection Information	Section TemplateID	PS3.20
1.2.840.10008.9.9	Fetus Findings	Section TemplateID	PS3.20
1.2.840.10008.9.10	Labeled Subsection	Section TemplateID	PS3.20
1.2.840.10008.9.11	Communication of Actionable Findings	Section TemplateID	PS3.20
1.2.840.10008.9.12	Recommendation	Section TemplateID	PS3.20
1.2.840.10008.9.13	Procedural Medication	Entry TemplateID	PS3.20
1.2.840.10008.9.14	Procedure Technique	Entry TemplateID	PS3.20
1.2.840.10008.9.15	Image Quality	Entry TemplateID	PS3.20
1.2.840.10008.9.16	Study Act	Entry TemplateID	PS3.20
1.2.840.10008.9.17	Series Act	Entry TemplateID	PS3.20
1.2.840.10008.9.18	SOP Instance Observation	Entry TemplateID	PS3.20
1.2.840.10008.9.19	Section Text	Element Set TemplateID	PS3.20
1.2.840.10008.9.20	General Header	Element Set TemplateID	PS3.20
1.2.840.10008.9.21	Imaging Header	Element Set TemplateID	PS3.20
1.2.840.10008.9.22	Parent Document	Element Set TemplateID	PS3.20
1.2.840.10008.9.23	General Section Entries	Element Set TemplateID	PS3.20
1.2.840.10008.9.24	Imaging Addendum Report	Document TemplateID	PS3.20

B Well-Known Color Palettes (Normative)
B.1 Standard Color Palettes
Table B.1-1 lists the color palettes that are defined by the DICOM Standard.
Table B.1-1. Standard Color Palettes
Well-known SOP Instance UID	Content Label (0070,0080)	Content Description (0070,0081)	Section	URL of Reference Encoded Instance
1.2.840.10008.1.​5.​1	HOT_IRON	Hot Iron	B.1.1	ftp://​medical.nema.org/​Medical/​Dicom/​Palettes/​hotiron.dcm
1.2.840.10008.1.​5.​2	PET	PET	B.1.2	ftp://​medical.nema.org/​Medical/​Dicom/​Palettes/​pet.dcm
1.2.840.10008.1.​5.​3	HOT_METAL_BLUE	Hot Metal Blue	B.1.3	ftp://​medical.nema.org/​Medical/​Dicom/​Palettes/​hotmetalblue.dcm
1.2.840.10008.1.​5.​4	PET_20_STEP	PET 20 Step	B.1.4	ftp://​medical.nema.org/​Medical/​Dicom/​Palettes/​pet20step.dcm
1.2.840.10008.1.​5.​5	SPRING	Spring	B.1.5	ftp://​medical.nema.org/​Medical/​Dicom/​Palettes/​spring.dcm
1.2.840.10008.1.​5.​6	SUMMER	Summer	B.1.6	ftp://​medical.nema.org/​Medical/​Dicom/​Palettes/​summer.dcm
1.2.840.10008.1.​5.​7	FALL	Fall	B.1.7	ftp://​medical.nema.org/​Medical/​Dicom/​Palettes/​fall.dcm
1.2.840.10008.1.​5.​8	WINTER	Winter	B.1.8	ftp://​medical.nema.org/​Medical/​Dicom/​Palettes/​winter.dcm

B.1.1 Hot Iron Color Palette
B.1.1.1 Hot Iron Color Palette Description (Informative)
The Hot Iron color palette is often used in nuclear medicine applications to make differences in signal intensity (counts) more apparent to the human observer. A typical example is illustrated in Figure B.1.1.1-1.

Figure B.1.1.1-1. Nuclear Medicine image with and without Hot Iron Palette applied.
B.1.1.2 Hot Iron Color Palette Definition
The ICC Profile shall define the sRGB space.
The value of Content Label (0070,0080) shall be "HOT_IRON".
This color palette is defined to contain the values for Red Palette Color Lookup Table Descriptor (0028,1101), Green Palette Color Lookup Table Descriptor (0028,1102) and Blue Palette Color Lookup Table Descriptor (0028,1103) defined in Table B.1.1.2-1.
Table B.1.1.2-1. Hot Iron Color Palette Descriptor
Value 1 (Number of entries)	Value 2 (First value mapped)	Value 3 (Number of bits)
256	0	8

This color palette is defined to contain the values in Table B.1.1.2-2, where the values in the columns Red, Green and Blue are the values of the Red Palette Color Lookup Table Data (0028,1201), Green Palette Color Lookup Table Data (0028,1202) and Blue Palette Color Lookup Table Data (0028,1203), respectively.
Table B.1.1.2-2. Hot Iron Color Palette Data
Red	Green	Blue
0	0	0
2	0	0
4	0	0
6	0	0
8	0	0
10	0	0
12	0	0
14	0	0
16	0	0
18	0	0
20	0	0
22	0	0
24	0	0
26	0	0
28	0	0
30	0	0
32	0	0
34	0	0
36	0	0
38	0	0
40	0	0
42	0	0
44	0	0
46	0	0
48	0	0
50	0	0
52	0	0
54	0	0
56	0	0
58	0	0
60	0	0
62	0	0
64	0	0
66	0	0
68	0	0
70	0	0
72	0	0
74	0	0
76	0	0
78	0	0
80	0	0
82	0	0
84	0	0
86	0	0
88	0	0
90	0	0
92	0	0
94	0	0
96	0	0
98	0	0
100	0	0
102	0	0
104	0	0
106	0	0
108	0	0
110	0	0
112	0	0
114	0	0
116	0	0
118	0	0
120	0	0
122	0	0
124	0	0
126	0	0
128	0	0
130	0	0
132	0	0
134	0	0
136	0	0
138	0	0
140	0	0
142	0	0
144	0	0
146	0	0
148	0	0
150	0	0
152	0	0
154	0	0
156	0	0
158	0	0
160	0	0
162	0	0
164	0	0
166	0	0
168	0	0
170	0	0
172	0	0
174	0	0
176	0	0
178	0	0
180	0	0
182	0	0
184	0	0
186	0	0
188	0	0
190	0	0
192	0	0
194	0	0
196	0	0
198	0	0
200	0	0
202	0	0
204	0	0
206	0	0
208	0	0
210	0	0
212	0	0
214	0	0
216	0	0
218	0	0
220	0	0
222	0	0
224	0	0
226	0	0
228	0	0
230	0	0
232	0	0
234	0	0
236	0	0
238	0	0
240	0	0
242	0	0
244	0	0
246	0	0
248	0	0
250	0	0
252	0	0
254	0	0
255	0	0
255	2	0
255	4	0
255	6	0
255	8	0
255	10	0
255	12	0
255	14	0
255	16	0
255	18	0
255	20	0
255	22	0
255	24	0
255	26	0
255	28	0
255	30	0
255	32	0
255	34	0
255	36	0
255	38	0
255	40	0
255	42	0
255	44	0
255	46	0
255	48	0
255	50	0
255	52	0
255	54	0
255	56	0
255	58	0
255	60	0
255	62	0
255	64	0
255	66	0
255	68	0
255	70	0
255	72	0
255	74	0
255	76	0
255	78	0
255	80	0
255	82	0
255	84	0
255	86	0
255	88	0
255	90	0
255	92	0
255	94	0
255	96	0
255	98	0
255	100	0
255	102	0
255	104	0
255	106	0
255	108	0
255	110	0
255	112	0
255	114	0
255	116	0
255	118	0
255	120	0
255	122	0
255	124	0
255	126	0
255	128	4
255	130	8
255	132	12
255	134	16
255	136	20
255	138	24
255	140	28
255	142	32
255	144	36
255	146	40
255	148	44
255	150	48
255	152	52
255	154	56
255	156	60
255	158	64
255	160	68
255	162	72
255	164	76
255	166	80
255	168	84
255	170	88
255	172	92
255	174	96
255	176	100
255	178	104
255	180	108
255	182	112
255	184	116
255	186	120
255	188	124
255	190	128
255	192	132
255	194	136
255	196	140
255	198	144
255	200	148
255	202	152
255	204	156
255	206	160
255	208	164
255	210	168
255	212	172
255	214	176
255	216	180
255	218	184
255	220	188
255	222	192
255	224	196
255	226	200
255	228	204
255	230	208
255	232	212
255	234	216
255	236	220
255	238	224
255	240	228
255	242	232
255	244	236
255	246	240
255	248	244
255	250	248
255	252	252
255	255	255

B.1.2 PET Color Palette
B.1.2.1 PET Color Palette Description (Informative)
The PET color palette is often used in PET applications to pseudo-color the superimposed PET images when displayed fused with underlying CT images. A typical example is illustrated in Figure B.1.2.1-1.

Figure B.1.2.1-1. PET image with PET Palette superimposed over grayscale CT image.
B.1.2.2 PET Color Palette Definition
The ICC Profile shall define the sRGB space.
The value of Content Label (0070,0080) shall be "PET".
This color palette is defined to contain the values for Red Palette Color Lookup Table Descriptor (0028,1101), Green Palette Color Lookup Table Descriptor (0028,1102) and Blue Palette Color Lookup Table Descriptor (0028,1103) defined in Table B.1.2.2-1.
Table B.1.2.2-1. PET Color Palette Descriptor
Value 1 (Number of entries)	Value 2 (First value mapped)	Value 3 (Number of bits)
256	0	8

This color palette is defined to contain the values in Table B.1.2.2-2, where the values in the columns Red,Green and Blue are the values of the Red Palette Color Lookup Table Data (0028,1201), Green Palette Color Lookup Table Data (0028,1202) and Blue Palette Color Lookup Table Data (0028,1203), respectively.
Table B.1.2.2-2. PET Color Palette Data
Red	Green	Blue
0	0	0
0	2	1
0	4	3
0	6	5
0	8	7
0	10	9
0	12	11
0	14	13
0	16	15
0	18	17
0	20	19
0	22	21
0	24	23
0	26	25
0	28	27
0	30	29
0	32	31
0	34	33
0	36	35
0	38	37
0	40	39
0	42	41
0	44	43
0	46	45
0	48	47
0	50	49
0	52	51
0	54	53
0	56	55
0	58	57
0	60	59
0	62	61
0	65	63
0	67	65
0	69	67
0	71	69
0	73	71
0	75	73
0	77	75
0	79	77
0	81	79
0	83	81
0	85	83
0	87	85
0	89	87
0	91	89
0	93	91
0	95	93
0	97	95
0	99	97
0	101	99
0	103	101
0	105	103
0	107	105
0	109	107
0	111	109
0	113	111
0	115	113
0	117	115
0	119	117
0	121	119
0	123	121
0	125	123
0	128	125
1	126	127
3	124	129
5	122	131
7	120	133
9	118	135
11	116	137
13	114	139
15	112	141
17	110	143
19	108	145
21	106	147
23	104	149
25	102	151
27	100	153
29	98	155
31	96	157
33	94	159
35	92	161
37	90	163
39	88	165
41	86	167
43	84	169
45	82	171
47	80	173
49	78	175
51	76	177
53	74	179
55	72	181
57	70	183
59	68	185
61	66	187
63	64	189
65	63	191
67	61	193
69	59	195
71	57	197
73	55	199
75	53	201
77	51	203
79	49	205
81	47	207
83	45	209
85	43	211
86	41	213
88	39	215
90	37	217
92	35	219
94	33	221
96	31	223
98	29	225
100	27	227
102	25	229
104	23	231
106	21	233
108	19	235
110	17	237
112	15	239
114	13	241
116	11	243
118	9	245
120	7	247
122	5	249
124	3	251
126	1	253
128	0	255
130	2	252
132	4	248
134	6	244
136	8	240
138	10	236
140	12	232
142	14	228
144	16	224
146	18	220
148	20	216
150	22	212
152	24	208
154	26	204
156	28	200
158	30	196
160	32	192
162	34	188
164	36	184
166	38	180
168	40	176
170	42	172
171	44	168
173	46	164
175	48	160
177	50	156
179	52	152
181	54	148
183	56	144
185	58	140
187	60	136
189	62	132
191	64	128
193	66	124
195	68	120
197	70	116
199	72	112
201	74	108
203	76	104
205	78	100
207	80	96
209	82	92
211	84	88
213	86	84
215	88	80
217	90	76
219	92	72
221	94	68
223	96	64
225	98	60
227	100	56
229	102	52
231	104	48
233	106	44
235	108	40
237	110	36
239	112	32
241	114	28
243	116	24
245	118	20
247	120	16
249	122	12
251	124	8
253	126	4
255	128	0
255	130	4
255	132	8
255	134	12
255	136	16
255	138	20
255	140	24
255	142	28
255	144	32
255	146	36
255	148	40
255	150	44
255	152	48
255	154	52
255	156	56
255	158	60
255	160	64
255	162	68
255	164	72
255	166	76
255	168	80
255	170	85
255	172	89
255	174	93
255	176	97
255	178	101
255	180	105
255	182	109
255	184	113
255	186	117
255	188	121
255	190	125
255	192	129
255	194	133
255	196	137
255	198	141
255	200	145
255	202	149
255	204	153
255	206	157
255	208	161
255	210	165
255	212	170
255	214	174
255	216	178
255	218	182
255	220	186
255	222	190
255	224	194
255	226	198
255	228	202
255	230	206
255	232	210
255	234	214
255	236	218
255	238	222
255	240	226
255	242	230
255	244	234
255	246	238
255	248	242
255	250	246
255	252	250
255	255	255

B.1.3 Hot Metal Blue Color Palette
B.1.3.1 Hot Metal Blue Color Palette Description (Informative)
The Hot Metal Blue color palette is often used in nuclear medicine or PET applications to make differences in signal intensity (counts) more apparent to the human observer. A typical example is illustrated in Figure B.1.3.1-1.

Figure B.1.3.1-1. PET image with Hot Metal Blue Palette applied.
B.1.3.2 Hot Metal Blue Color Palette Definition
The ICC Profile shall define the sRGB space.
The value of Content Label (0070,0080) shall be "HOT_METAL_BLUE".
This color palette is defined to contain the values for Red Palette Color Lookup Table Descriptor (0028,1101), Green Palette Color Lookup Table Descriptor (0028,1102) and Blue Palette Color Lookup Table Descriptor (0028,1103) defined in Table B.1.3.2-1.
Table B.1.3.2-1. Hot Metal Blue Color Palette Descriptor
Value 1 (Number of entries)	Value 2 (First value mapped)	Value 3 (Number of bits)
256	0	8

This color palette is defined to contain the values in Table B.1.3.2-2, where the values in the columns Red,Green and Blue are the values of the Red Palette Color Lookup Table Data (0028,1201), Green Palette Color Lookup Table Data (0028,1202) and Blue Palette Color Lookup Table Data (0028,1203), respectively.
Table B.1.3.2-2. Hot Metal Blue Color Palette Data
Red	Green	Blue
0	0	0
0	0	2
0	0	4
0	0	6
0	0	8
0	0	10
0	0	12
0	0	14
0	0	16
0	0	17
0	0	19
0	0	21
0	0	23
0	0	25
0	0	27
0	0	29
0	0	31
0	0	33
0	0	35
0	0	37
0	0	39
0	0	41
0	0	43
0	0	45
0	0	47
0	0	49
0	0	51
0	0	53
0	0	55
0	0	57
0	0	59
0	0	61
0	0	63
0	0	65
0	0	67
0	0	69
0	0	71
0	0	73
0	0	75
0	0	77
0	0	79
0	0	81
0	0	83
0	0	84
0	0	86
0	0	88
0	0	90
0	0	92
0	0	94
0	0	96
0	0	98
0	0	100
0	0	102
0	0	104
0	0	106
0	0	108
0	0	110
0	0	112
0	0	114
0	0	116
0	0	117
0	0	119
0	0	121
0	0	123
0	0	125
0	0	127
0	0	129
0	0	131
0	0	133
0	0	135
0	0	137
0	0	139
0	0	141
0	0	143
0	0	145
0	0	147
0	0	149
0	0	151
0	0	153
0	0	155
0	0	157
0	0	159
0	0	161
0	0	163
0	0	165
0	0	167
3	0	169
6	0	171
9	0	173
12	0	175
15	0	177
18	0	179
21	0	181
24	0	183
26	0	184
29	0	186
32	0	188
35	0	190
38	0	192
41	0	194
44	0	196
47	0	198
50	0	200
52	0	197
55	0	194
57	0	191
59	0	188
62	0	185
64	0	182
66	0	179
69	0	176
71	0	174
74	0	171
76	0	168
78	0	165
81	0	162
83	0	159
85	0	156
88	0	153
90	0	150
93	2	144
96	4	138
99	6	132
102	8	126
105	9	121
108	11	115
111	13	109
114	15	103
116	17	97
119	19	91
122	21	85
125	23	79
128	24	74
131	26	68
134	28	62
137	30	56
140	32	50
143	34	47
146	36	44
149	38	41
152	40	38
155	41	35
158	43	32
161	45	29
164	47	26
166	49	24
169	51	21
172	53	18
175	55	15
178	56	12
181	58	9
184	60	6
187	62	3
190	64	0
194	66	0
198	68	0
201	70	0
205	72	0
209	73	0
213	75	0
217	77	0
221	79	0
224	81	0
228	83	0
232	85	0
236	87	0
240	88	0
244	90	0
247	92	0
251	94	0
255	96	0
255	98	3
255	100	6
255	102	9
255	104	12
255	105	15
255	107	18
255	109	21
255	111	24
255	113	26
255	115	29
255	117	32
255	119	35
255	120	38
255	122	41
255	124	44
255	126	47
255	128	50
255	130	53
255	132	56
255	134	59
255	136	62
255	137	65
255	139	68
255	141	71
255	143	74
255	145	76
255	147	79
255	149	82
255	151	85
255	152	88
255	154	91
255	156	94
255	158	97
255	160	100
255	162	103
255	164	106
255	166	109
255	168	112
255	169	115
255	171	118
255	173	121
255	175	124
255	177	126
255	179	129
255	181	132
255	183	135
255	184	138
255	186	141
255	188	144
255	190	147
255	192	150
255	194	153
255	196	156
255	198	159
255	200	162
255	201	165
255	203	168
255	205	171
255	207	174
255	209	176
255	211	179
255	213	182
255	215	185
255	216	188
255	218	191
255	220	194
255	222	197
255	224	200
255	226	203
255	228	206
255	229	210
255	231	213
255	233	216
255	235	219
255	237	223
255	239	226
255	240	229
255	242	232
255	244	236
255	246	239
255	248	242
255	250	245
255	251	249
255	253	252
255	255	255

B.1.4 PET 20 Step Color Palette
B.1.4.1 PET 20 Step Color Palette Description (Informative)
The PET 20 Step color palette is often used in PET applications to make differences in signal intensity (counts) more apparent to the human observer. A typical example is illustrated in Figure B.1.4.1-1.

Figure B.1.4.1-1. PET image with PET 20 Step Palette applied.
B.1.4.2 PET 20 Step Color Palette Definition
The ICC Profile shall define the sRGB space.
The value of Content Label (0070,0080) shall be "PET_20_STEP".
This color palette is defined to contain the values for Red Palette Color Lookup Table Descriptor (0028,1101), Green Palette Color Lookup Table Descriptor (0028,1102) and Blue Palette Color Lookup Table Descriptor (0028,1103) defined in Table B.1.4.2-1.
Table B.1.4.2-1. PET 20 Step Color Palette Descriptor
Value 1 (Number of entries)	Value 2 (First value mapped)	Value 3 (Number of bits)
256	0	8

This color palette is defined to contain the values in Table B.1.4.2-2, where the values in the columns Red,Green and Blue are the values of the Red Palette Color Lookup Table Data (0028,1201), Green Palette Color Lookup Table Data (0028,1202) and Blue Palette Color Lookup Table Data (0028,1203), respectively.
Table B.1.4.2-2. PET 20 Step Color Palette Data
Red	Green	Blue
0	0	0
0	0	0
0	0	0
0	0	0
0	0	0
0	0	0
0	0	0
0	0	0
0	0	0
0	0	0
0	0	0
0	0	0
0	0	0
96	0	80
96	0	80
96	0	80
96	0	80
96	0	80
96	0	80
96	0	80
96	0	80
96	0	80
96	0	80
96	0	80
96	0	80
96	0	80
48	48	80
48	48	80
48	48	80
48	48	80
48	48	80
48	48	80
48	48	80
48	48	80
48	48	80
48	48	80
48	48	80
48	48	80
48	48	80
48	48	112
48	48	112
48	48	112
48	48	112
48	48	112
48	48	112
48	48	112
48	48	112
48	48	112
48	48	112
48	48	112
48	48	112
80	80	128
80	80	128
80	80	128
80	80	128
80	80	128
80	80	128
80	80	128
80	80	128
80	80	128
80	80	128
80	80	128
80	80	128
80	80	128
96	96	176
96	96	176
96	96	176
96	96	176
96	96	176
96	96	176
96	96	176
96	96	176
96	96	176
96	96	176
96	96	176
96	96	176
96	96	176
112	112	192
112	112	192
112	112	192
112	112	192
112	112	192
112	112	192
112	112	192
112	112	192
112	112	192
112	112	192
112	112	192
112	112	192
112	112	192
128	128	224
128	128	224
128	128	224
128	128	224
128	128	224
128	128	224
128	128	224
128	128	224
128	128	224
128	128	224
128	128	224
128	128	224
48	96	48
48	96	48
48	96	48
48	96	48
48	96	48
48	96	48
48	96	48
48	96	48
48	96	48
48	96	48
48	96	48
48	96	48
48	96	48
48	144	48
48	144	48
48	144	48
48	144	48
48	144	48
48	144	48
48	144	48
48	144	48
48	144	48
48	144	48
48	144	48
48	144	48
48	144	48
80	192	80
80	192	80
80	192	80
80	192	80
80	192	80
80	192	80
80	192	80
80	192	80
80	192	80
80	192	80
80	192	80
80	192	80
80	192	80
64	224	64
64	224	64
64	224	64
64	224	64
64	224	64
64	224	64
64	224	64
64	224	64
64	224	64
64	224	64
64	224	64
64	224	64
224	224	80
224	224	80
224	224	80
224	224	80
224	224	80
224	224	80
224	224	80
224	224	80
224	224	80
224	224	80
224	224	80
224	224	80
224	224	80
208	208	96
208	208	96
208	208	96
208	208	96
208	208	96
208	208	96
208	208	96
208	208	96
208	208	96
208	208	96
208	208	96
208	208	96
208	208	96
208	176	64
208	176	64
208	176	64
208	176	64
208	176	64
208	176	64
208	176	64
208	176	64
208	176	64
208	176	64
208	176	64
208	176	64
208	176	64
208	144	0
208	144	0
208	144	0
208	144	0
208	144	0
208	144	0
208	144	0
208	144	0
208	144	0
208	144	0
208	144	0
208	144	0
192	96	0
192	96	0
192	96	0
192	96	0
192	96	0
192	96	0
192	96	0
192	96	0
192	96	0
192	96	0
192	96	0
192	96	0
192	96	0
176	48	0
176	48	0
176	48	0
176	48	0
176	48	0
176	48	0
176	48	0
176	48	0
176	48	0
176	48	0
176	48	0
176	48	0
176	48	0
255	0	0
255	0	0
255	0	0
255	0	0
255	0	0
255	0	0
255	0	0
255	0	0
255	0	0
255	0	0
255	0	0
255	0	0
255	0	0
255	255	255
255	255	255
255	255	255
255	255	255
255	255	255
255	255	255
255	255	255
255	255	255
255	255	255
255	255	255
255	255	255
255	255	255
255	255	255

B.1.5 Spring Color Palette
B.1.5.1 Spring Color Palette Description (Informative)
The Spring Color Palette is suggested for use in color fMRI activation maps. It shades from one pastel color to another which is distinctly different, making it suitable for illustrating either unipolar or bipolar activation. As part of a complementary set of color palettes (Spring, Summer, Fall, Winter), it conveys activation strength within one statistical parametric map, while making it possible for the human observer to distinguish between different fMRI activation maps in the same blended display. A typical example is illustrated in Figure B.1.5.1-1.

Figure B.1.5.1-1. MR image with Spring LUT Palette applied.
B.1.5.2 Spring Color Palette Definition
The ICC Profile shall define the sRGB space.
The value of Content Label (0070,0080) shall be "SPRING".
This color palette is defined to contain the values for Red Palette Color Lookup Table Descriptor (0028,1101), Green Palette Color Lookup Table Descriptor (0028,1102) and Blue Palette Color Lookup Table Descriptor (0028,1103) defined in Table B.1.5.2-1.
Table B.1.5.2-1. Spring Color Palette Descriptor
Value 1 (Number of entries)	Value 2 (First value mapped)	Value 3 (Number of bits)
256	0	8

This color palette is defined using the segmented lookup table data specified in Table B.1.5.2-2, where the values in the columns Red, Green and Blue are the values of the Segmented Red Palette Color Lookup Table Data (0028,1221), Segmented Green Palette Color Lookup Table Data (0028,1222) and Segmented Blue Palette Color Lookup Table Data (0028,1223), respectively
Table B.1.5.2-2. Spring Segmented Color Palette Data
Red	Green	Blue
0	0	0
1	1	1
255	0	255
1	1	1
255	255	255
255	255	0

B.1.6 Summer Color Palette
B.1.6.1 Summer Color Palette Description (Informative)
The Summer Color Palette is suggested for use in color fMRI activation maps. It shades from one pastel color to another which is distinctly different, making it suitable for illustrating either unipolar or bipolar activation. As part of a complementary set of color palettes (Spring, Summer, Fall, Winter), it conveys activation strength within one statistical parametric map, while making it possible for the human observer to distinguish between different fMRI activation maps in the same blended display. A typical example is illustrated in Figure B.1.6.1-1.

Figure B.1.6.1-1. MR image with Summer LUT Palette applied.
B.1.6.2 Summer Color Palette Definition
The ICC Profile shall define the sRGB space.
The value of Content Label (0070,0080) shall be "SUMMER".
This color palette is defined to contain the values for Red Palette Color Lookup Table Descriptor (0028,1101), Green Palette Color Lookup Table Descriptor (0028,1102) and Blue Palette Color Lookup Table Descriptor (0028,1103) defined in Table B.1.6.2-1.
Table B.1.6.2-1. Summer Color Palette Descriptor
Value 1 (Number of entries)	Value 2 (First value mapped)	Value 3 (Number of bits)
256	0	8

This color palette is defined using the segmented lookup table data specified in Table B.1.6.2-2, where the values in the columns Red, Green and Blue are the values of the Segmented Red Palette Color Lookup Table Data (0028,1221), Segmented Green Palette Color Lookup Table Data (0028,1222) and Segmented Blue Palette Color Lookup Table Data (0028,1223), respectively
Table B.1.6.2-2. Summer Segmented Color Palette Data
Red	Green	Blue
0	0	0
1	1	1
0	255	0
1	1	1
255	255	127
0	128	0
		1
		128
		254

B.1.7 Fall Color Palette
B.1.7.1 Fall Color Palette Description (Informative)
The Fall Color Palette is suggested for use in color fMRI activation maps. It shades from one pastel color to another which is distinctly different, making it suitable for illustrating either unipolar or bipolar activation. As part of a complementary set of color palettes (Spring, Summer, Fall, Winter), it conveys activation strength within one statistical parametric map, while making it possible for the human observer to distinguish between different fMRI activation maps in the same blended display. A typical example is illustrated in Figure B.1.7.1-1.

Figure B.1.7.1-1. MR image with Fall LUT Palette applied.
B.1.7.2 Fall Color Palette Definition
The ICC Profile shall define the sRGB space.
The value of Content Label (0070,0080) shall be "FALL".
This color palette is defined to contain the values for Red Palette Color Lookup Table Descriptor (0028,1101), Green Palette Color Lookup Table Descriptor (0028,1102) and Blue Palette Color Lookup Table Descriptor (0028,1103) defined in Table B.1.7.2-1.
Table B.1.7.2-1. Fall Color Palette Descriptor
Value 1 (Number of entries)	Value 2 (First value mapped)	Value 3 (Number of bits)
256	0	8

This color palette is defined using the segmented lookup table data specified in Table B.1.7.2-2, where the values in the columns Red, Green and Blue are the values of the Segmented Red Palette Color Lookup Table Data (0028,1221), Segmented Green Palette Color Lookup Table Data (0028,1222) and Segmented Blue Palette Color Lookup Table Data (0028,1223), respectively
Table B.1.7.2-2. Fall Segmented Color Palette Data
Red	Green	Blue
0	0	0
1	1	1
255	255	0
1	1	1
255	255	255
255	0	0

B.1.8 Winter Color Palette
B.1.8.1 Winter Color Palette Description (Informative)
The Winter Color Palette is suggested for use in color fMRI activation maps. It shades from one pastel color to another which is distinctly different, making it suitable for illustrating either unipolar or bipolar activation. As part of a complementary set of color palettes (Spring, Summer, Fall, Winter), it conveys activation strength within one statistical parametric map, while making it possible for the human observer to distinguish between different fMRI activation maps in the same blended display. A typical example is illustrated in Figure B.1.8.1-1.

Figure B.1.8.1-1. MR image with Winter LUT Palette applied.
B.1.8.2 Winter Color Palette Definition
The ICC Profile shall define the sRGB space.
The value of Content Label (0070,0080) shall be "WINTER".
This color palette is defined to contain the values for Red Palette Color Lookup Table Descriptor (0028,1101), Green Palette Color Lookup Table Descriptor (0028,1102) and Blue Palette Color Lookup Table Descriptor (0028,1103) defined in Table B.1.8.2-1.
Table B.1.8.2-1. Winter Color Palette Descriptor
Value 1 (Number of entries)	Value 2 (First value mapped)	Value 3 (Number of bits)
256	0	8

This color palette is defined using the segmented lookup table data specified in Table B.1.8.2-2, where the values in the columns Red, Green and Blue are the values of the Segmented Red Palette Color Lookup Table Data (0028,1221), Segmented Green Palette Color Lookup Table Data (0028,1222) and Segmented Blue Palette Color Lookup Table Data (0028,1223), respectively
Table B.1.8.2-2. Winter Segmented Color Palette Data
Red	Green	Blue
0	0	0
1	1	1
0	0	255
1	1	1
127	255	255
0	255	128
1
128
127

B.2 Localized Standard Color Palette Description Values
B.2.1 French
Table B.2.1-1. French Standard Color Palette Description Values
Content Label (0070,0080)	English Value of Content Description (0070,0081)	French Value of Content Description (0070,0081)
HOT_IRON	Hot Iron	Hot Iron
PET	PET	TEP
HOT_METAL_BLUE	Hot Metal Blue	Hot Metal Blue
PET_20_STEP	PET 20 Step	TEP Vingt étapes
SPRING	Spring	Printemps
SUMMER	Summer	Été
FALL	Fall	Automne
WINTER	Winter	Hiver

Note
In France, the English terms for "Hot Iron" and "Hot Metal Blue" are used.
B.2.2 German
Table B.2.2-1. German Standard Color Palette Description Values
Content Label (0070,0080)	English Value of Content Description (0070,0081)	German Value of Content Description (0070,0081)
HOT_IRON	Hot Iron	Heißes Eisen
PET	PET	PET
HOT_METAL_BLUE	Hot Metal Blue	Heißes Metallblau
PET_20_STEP	PET 20 Step	PET 20 Schritte
SPRING	Spring	Frühling
SUMMER	Summer	Sommer
FALL	Fall	Herbst
WINTER	Winter	Winter